# APPENDIX 1: Technical Spécification & Supplier Response Document

# Tender Reference: NCTR2687

# Request for Tenders

for the Supply, Delivery, Installation of a precision fermentation platform for TU Dublin

# Location

Central Quad, Grangegorman

| Issue Date: | Friday 26th of June 2026 |
| --- | --- |
| Contact Details for Queries: | E-tenders Messaging facility |
| Closing Date for Receipt of Queries: | 2pm Friday the 10th of July 2026 |
| Deadline for Receipt of Tenders:. | 2pm Friday the 17th of July 2026 |

Introduction

TU Dublin, Ireland’s first Technological University, is where career-focused students, dedicated staff and academic excellence in science, the arts, business, engineering and technology converge to create the leaders of tomorrow. We offer an inclusive and open learning experience with pathways to graduation, from Apprenticeship to PhD. Our 28,500 students learn in a practice-based environment informed by the latest research and enabled by technological advances.

TU Dublin hosts a thriving research community engaged in applying innovation and technology to solve the world’s most pressing challenges. We are deeply committed to collaborating with our national and international academic partners and our many networks in industry and civic society to create new learning experiences and develop impactful research.

TU Dublin is an exciting new milestone in Irish Higher Education.  With campuses in Dublin City, Tallaght and Blanchardstown, it spans the largest population centres of Ireland’s capital city.  Building on the rich heritage of its founding organisations - DIT, IT Blanchardstown and IT Tallaght – TU Dublin is inclusive and adaptable, creating educational opportunities for students at all stages.  TU Dublin students are socially responsible, open-minded global thinkers who are ambitious to change the world for the better. As graduates, they will be enterprising and daring in all their endeavours, ready to play their part in transforming the future.

Summary of requirements

Tenders are sought for establishing a precision fermentation platform that will build on the national landscape capabilities in the field of precision fermentation as it will allow a detailed monitoring of the fermentation process and facilitate an inline optimization of the processes. The platform requires (i) three research scale benchtop bioreactors 500 ml to 1000 ml (Lot 1) and (ii) One Light module to transform a 2 Litres bioreactor into a photobioreactor (Lot 2)

The fermentation platform must be capable of running parallel experiments in all the benchtop bioreactors with independent process control.

This competition will be split into two lots. Suppliers can tender for one or both lots.

Funding Acknowledgment

TU Dublin Research Equipment Grant Programme (2025) (TREG 2025)

SECTION A: REQUIREMENTS

| REF | PASS/FAIL CRITERIA | PASS REQUIREMENT |
| --- | --- | --- |
| A1 | Applicant Summary | Applicants must complete this section. If the Applicant is a grouping, then a separate questionnaire must be completed for each group member. |
| A2 | Financial | Applicants are required to self-declare compliance with the requirements. Applicants should note that if deemed the most economically advantageous tender, they will be required to provide the evidence self-declared prior to contract award. |
| A3 | Tax Compliance |  |
| A4 | Insurance |  |
| A5 | Previous Experience | Applicants must supply contact details and information on services provided for three referees to whom you have provided similar services to the requirements of this RFT. |
| A6 | Health & Safety | Applicants must demonstrate that they comply with the relevant legislation with reference to Environmental, Health and Safety. |
| A7 | Data Protection Compliance | The Tenderer must ensure the security of Personal Data at all times is in line with the Data Protection Acts 1998 and 2003 and guidance issued by the Data Protection Commissioner of Ireland. Tenderers are required to be compliant under the EU General Data Protection Regulation (GDPR) which came into effect on 25/05/2018. Please confirm compliance. Compliant Yes/No:. Signature:. |

SECTION A: PASS/FAIL CRITERIA

| A1. APPLICANT SUMMARY Weighting: Pass/Fail only Pass requirement: Applicants must complete this section |  |
| --- | --- |
| Organisation Name | Insert Response |
| Contact Name | Insert Response |
| Address | Insert Response |
| Telephone | Insert Response |
| Email | Insert Response |
| Legal Status (if any) (e.g., Company, Partnership, Sole Trader, etc.) | Insert Response |

| The below is a requirement of each Applicant |  |
| --- | --- |
| Requirement | Yes / No |
| The Applicant must submit a signed Tenderers statement.as per Appendix 3 of the RFT | Choose an item. |
| The Applicant must submit a Signed Declaration as to Personal Circumstances. as per Article 45 of Directive 2004/18/EC & Regulation 53 of SI329); Appendix 4 | Choose an item. |

A2/A3/A4. FINANCIAL/ TAX COMPLIANCE / INSURANCE Weighting: Pass/Fail only Pass requirement: Applicants must confirm that the meet the following minimum rules

| A2. Financial Summary |  |
| --- | --- |
| Requirement | Yes / No |
| The company is solvent and of good economic standing and can if requested supply an accountant’s/ auditor’s statement for the last three financial years. Availability of a signed statement from accountant/auditor confirming an annual company turnover of €250,000 (Lot 1) & €33,000 (Lot 2)for each of the last 3 years | Choose an item. |

| A3. Tax Compliance |  |
| --- | --- |
| Requirement | Yes / No |
| The Candidate holds a current and valid Tax Clearance Certificate(s) from the Revenue Commissioners or in the case of a Candidate outside of the Republic of Ireland hold or can provide an equivalent statement from the relevant taxation authority. | Choose an item. |

| A4. Insurances |  |
| --- | --- |
| Requirement | Yes / No |
| That the Candidate holds, or will hold should they be successful, the following insurances. TU Dublin reserves the right to amend the requirements before a Purchase Order is issued. |  |
| Employer’s Liability - €13 million limit for any one claim or series of claims arising out of a single occurrence | Choose an item. |
| Public Liability - €6.5 million limit for any one claim or series of claims arising out of a single occurrence | Choose an item. |
| Product Liability - €6.5 million limit for any one claim or series of claims arising out of a single occurrence | Choose an item. |

Documentary evidence of economic and financial standing is not required at this stage of the Competition. Instead, Candidates are asked to declare that they meet the minimum standards required by the Contracting Authority and will be in a position to produce the necessary documentation (e.g., bank statements, audited accounts, proof of insurance, etc.) if prior to any award of contract.

A5: PREVIOUS EXPERIENCE/REFERENCES

The tenderer MUST have a track record of successful supply and installation of similar systems to those being tendered for. The Tenderer MUST provide the relevant reference sites of where two (2) similar systems related to that being tendered  have been successfully installed. Please complete tables in appendix 1 prior to specification requests for the relevant lot s)

Lot 1 Supply, delivery and installation of Benchtop Bioreactors

| PREVIOUS EXPERIENCE (CONTRACT ONE)- |  |  |  |
| --- | --- | --- | --- |
| Title of Project |  |  |  |
| Client Name |  |  |  |
| Client Contact Details | Contact Name |  |  |
|  | Contact E-mail Address |  |  |
|  | Contact Telephone Number |  |  |
| Annual Value |  |  |  |
| Start Date |  | End Date |  |
| Please provide a summary of the service provided and indicate how it is similar to the goods being sought by TU Dublin. |  |  |  |
| Please insert details here |  |  |  |
| PREVIOUS EXPERIENCE (CONTRACT TWO)- |  |  |  |
| Title of Project |  |  |  |
| Client Name |  |  |  |
| Client Contact Details | Contact Name |  |  |
|  | Contact E-mail Address |  |  |
|  | Contact Telephone Number |  |  |
| Annual Value |  |  |  |
| Start Date |  | End Date |  |
| Please provide a summary of the service provided and indicate how it is similar to the goods being sought by TU Dublin. |  |  |  |
| Please insert details here |  |  |  |

Lot 2 Supply, delivery and installation of a light module for a Bioreactor

| PREVIOUS EXPERIENCE (CONTRACT ONE)- |  |  |  |
| --- | --- | --- | --- |
| Title of Project |  |  |  |
| Client Name |  |  |  |
| Client Contact Details | Contact Name |  |  |
|  | Contact E-mail Address |  |  |
|  | Contact Telephone Number |  |  |
| Annual Value |  |  |  |
| Start Date |  | End Date |  |
| Please provide a summary of the service provided and indicate how it is similar to the goods being sought by TU Dublin. |  |  |  |
| Please insert details here |  |  |  |
| PREVIOUS EXPERIENCE (CONTRACT TWO)- |  |  |  |
| Title of Project |  |  |  |
| Client Name |  |  |  |
| Client Contact Details | Contact Name |  |  |
|  | Contact E-mail Address |  |  |
|  | Contact Telephone Number |  |  |
| Annual Value |  |  |  |
| Start Date |  | End Date |  |
| Please provide a summary of the service provided and indicate how it is similar to the goods being sought by TU Dublin. |  |  |  |
| Please insert details here |  |  |  |

| A6. Health & Safety |  |
| --- | --- |
| Requirement | Yes / No |
| Candidates must demonstrate that they comply with the relevant legislation with reference to Environmental, Health and Safety. | Choose an item. |

A7: DATA PROTECTION COMPLIANCE

Tenderers are required to be compliant under the EU General Data Protection Regulation (GDPR) replacing the Data Protection Directive 95/46/EC which came into effect on 25/05/2018. Please confirm compliance.

| Compliant Yes/No: |  |
| --- | --- |
| Signature: |  |

PLEASE NOTE

In the first instance, all submissions will be checked to ensure they meet the mandatory pass/fail requirements in Section A in order to determine compliancy and responsiveness. All submissions that meet the mandatory pass/fail criteria will be deemed to be compliant and responsive and will be assessed under the weighted criteria set out in Section B. Any submission that does not meet the mandatory pass/fail requirements will be deemed to be non-compliant and non-responsive and will be eliminated.

In relation to minimum score required - Tenderers should note that they must achieve a minimum rating set out for each of the individual qualitative criteria to avoid elimination from the competition

Requirements & Specification

Suppliers MUST provide all information relating to the instrument and the associated consumables in the response boxes below. Responses will be scored accordingly by the Technical Evaluation team as set out in the Award Criteria in the SRFT document.

Please ensure all cost submissions are kept separate and included in the Excel file entitled: Appendix 2 – Pricing Schedule.

Where the bidder has more than one system matching the specifications, details for each system should be provided separately and should be limited to 2 per tenderer.

Technical Specification

This is the minimum specification required.  All tenderers must meet or exceed this minimum specification to progress through the award criteria to cost evaluation.

Additional functionality should be identified in the comments section to facilitate scoring under the award criteria.

Notes

- Where specifications include reference to manufacturer or brand names this is a result of difficulties in specifying the product required. In these instances the phrase “or equivalent” is inferred.

- Compliance with all the standards as regards health and safety, quality assurance and analytical protocols is expected.

It is acknowledged that each Tender’s product line may differ from these specifications.  As such, each seller is free to propose variances from these specifications. Alternative proposals favourable to the Contracting Authority may be suggested. It is required, however, that, whenever a variance from these specifications occurs, the proposed item must meet or exceed the specified characteristics or level of performance.

Programme Requirements

2.1	Tenderers MUST note that contractors do not have exclusive access and will be required to work around and accommodate other suppliers and contractors who will be active in the building and laboratories simultaneously.

2.2	TU Dublin working hours are Monday to Friday from 9.00am to 5.00 pm (Business Hours). Access to the site outside the Business Hours shall at all times be arranged through the appointed representative in TU Dublin.

Note: Working outside Business Hours may be required to comply with work programmes. Where this becomes necessary as a result of changes to the start dates set out above, all reasonable efforts will be made to accommodate same.

2.3	Contractors are responsible for security of their own equipment, materials, etc.

Prior to any deliveries, the successful supplier(s) (and/or  relevant sub-contractors, etc.) will be required to meet the representatives from TU Dublin to clarify the requirements for Health & Safety, site access, services etc. to ensure compliance with standard operating procedures, and other regulatory/ statutory requirements

Award Criteria

The Goods Contract will be awarded on the basis of the most economically advantageous tender(s) as identified in accordance with the following award criteria:

Lot 1 Award Criteria

| Award Criteria |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | Award Criteria: Please Note: Minimum Scores Requirement 60% for C &D For Section B, Tenderers must meet the minimum score of 60% in each sub-criteria. | Overall Weight | Total Maximum Score | Minimum Score (60%) Requirement |
| B | Quality & Technical Merit - General Functional Requirements - Hardware, Software & Interaction of Devices | 40 % 30 % 10% | 400 300 100 | 180 60 |
| C | - After Sales and Technical Assistance - Customer Support, Repair, Maintenance & Breakdown Requirements - Warranty & Preventative Maintenance - Training Proposal & Documentation | 20% 10 % 5 % 5 % | 200 | 120 |
| D | Delivery Date or Date of Completion - Lead Time & Delivery/Completion Plan - Installation, Commissioning, Calibration & Validation - Sustainability | 10% | 100 | 60 |
| E | Ultimate Cost (Total Cost of Ownership) Including - Capital Cost, Delivery costs, Installation Costs, Training Costs, Warranty, Maintenance & PM Costs, (IF NOT INCLUDED IN OVERALL COST). | 30% | 300 | n/a |
|  | Total Maximum Score Available 100% |  | 1000 |  |

Lot 2 Award Criteria

| Award Criteria |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | Award Criteria: Please Note: Minimum Scores Requirement 60% for B, C &D | Overall Weight | Total Maximum Score | Minimum Score (60%) Requirement |
| B | Quality & Technical Merit - General Functional Requirements | 40 % | 400 | 240 |
| C | After Sales and Technical Assistance - Customer Support, Repair, Maintenance & Breakdown Requirements - Warranty & Preventative Maintenance - Training Proposal & Documentation | 20% 10 % 5 % 5 % | 200 | 120 |
| D | Delivery Date or Date of Completion - Lead Time & Delivery/Completion Plan - Installation, Commissioning, Calibration & Validation - Sustainability | 10% | 100 | 60 |
| E | Ultimate Cost (Total Cost of Ownership) Including - Capital Cost, Delivery costs, Installation Costs, Training Costs, Warranty, Maintenance & PM Costs, (IF NOT INCLUDED IN OVERALL COST). | 30% | 300 | n/a |
|  | Total Maximum Score Available 100% |  | 1000 |  |

The tender award criteria for each will be applied to each tender who has successfully passed the Selection Criteria. Tenderers must achieve a minimum of 60% on all qualitative tender award criteria above, excluding Ultimate Cost, in order to avoid elimination from the competition. Only tenderers who have achieved the 60% will be evaluated against Ultimate Cost.

The lowest Ultimate Cost tender, will receive the maximum score achievable under this criterion.

| Pricing Score = | Lowest Tendered Price x Max Points |  |
| --- | --- | --- |
|  | Price of bid being evaluated |  |

Tenderers must use “Appendix 1: Requirements and Specifications to structure their response on how they meet the requirements.

No commitment of any kind, contractual or otherwise shall exist unless and until a formal written contract has been executed by or on behalf of the Contracting Authority. Any award of notification of preferred bidder status by the Contracting Authority shall not give rise to any enforceable rights by the Tenderer. The Contracting Authority may cancel this public procurement competition at any time prior to a formal written contract being executed by or on behalf of the Contracting Authority. The Contracting Authority does not bind itself to accept the lowest priced or any Tender.

Responses for Non-cost/qualitative Award Criteria items will be scored using the following matrix:

|  | Scoring Methodology |
| --- | --- |
| Score | Meaning |
| 5 | Excellent response that fully meets and exceeds the requirements, and provides very comprehensive, detailed, and convincing assurance that the Tenderer will deliver to an excellent standard. |
| 4 | A very good response that meets the requirements and exceeds them in some respects. Response provides detailed assurance that the Tenderer will deliver to a very good or high standard. |
| 3 | A satisfactory response which demonstrates a reasonable understanding of requirements and gives reasonable assurance of delivery to an adequate standard but does not provide sufficiently convincing assurance to award a higher mark. |
| 2 | A response where reservations exist. Lacks full credibility/convincing detail, and does not provide confidence to the Contracting Authority that the required Goods/Services will be successfully delivered. |
| 1 | A response where serious reservations exist. This may be because, for example, insufficient detail is provided, or the response has fundamental flaws, or is seriously inadequate or seriously lacks credibility with a high risk of non-delivery. |
| 0 | No response or partial response only and poor evidence provided in support of it: failure to meet the requirements. |
|  | Note: Scores will be awarded on the basis of whole numbers only. |

Suppliers please note:

- This specification document is designed to assist the contracting authority in its evaluation of your tender and to enable suppliers provide structured responses to key deliverables. It is structured to reflect the award criteria as set out in the SRFT.

- Please note responses to individual questions will NOT be scored separately but collectively they will provide a basis for assessing the merits of the tenderer’s overall response to each award criterion and sub-criterion.

A failure to answer question(s) may leave the contracting authority with insufficient information on which to base a thorough evaluation and in turn may result in loss of marks in respect of the relevant Award Criterion. Therefore, please ensure you complete ALL sections of this questionnaire.

Next Ranked Supplier

The award of any contract is subject to the successful tenderer meeting the required specification and standards as set out in this RFT and Appendix 1. Should the required specification and standards not be met, the TU Dublin reserves the right Authority, to award any contract to the next highest-ranked Tenderer.

Tie Break

If the evaluation outcome results in a tie between two or more Tenders, then the Tender with the highest score for the highest weighted criterion shall be deemed the most economically advantageous tender. If the score remains in a tie, the remaining criterion will be assessed by the highest weighted until there is no longer a tie.  If the score remains a tie after all criterion are assessed, the tenderers will be asked to submit a Best and Final Offer (BAFO).

- Specification Required & Supplier Response

Suppliers MUST provide all information relating to the instrument and the associated consumables in the blue response boxes below. Responses will be scored accordingly by the Evaluation team as set out in the Award Criteria in the RFT document.

Lot 1 – Three benchtop Bioreactors

| (Suppliers to state model type here) |  |  |  |  |
| --- | --- | --- | --- | --- |
| REF. |  | REQUIREMENT | CONFIRM (YES/NO) | COMMENT |
| B. Technical Merit |  |  |  |  |
| Functional Requirements |  |  |  |  |
| B1 | The required three bioreactors MUST all be of the same total volume which may range from 500 ml to 1000 ml with independent process control required. The working capacity MUST be at least 60% of total volume for each bioreactor. |  | Compliant (Yes / No) | Insert Response |
| B2 | The three bioreactors MUST each have capability for batch, fed- batch microbial (bacterial and fungal) culturing and fermentation for aerobic and anaerobic processes. This should include high-cell-density cultures, biotransformation, and mixed-culture consortia to support process development and scale-up studies. Desirably each of the three bioreactors should have capability for continuous microbial (bacterial and fungal) culturing and fermentation for aerobic and anaerobic processes. |  | Compliant (Yes / No) | Insert Response |
| B3 | The three bioreactors MUST each, along with any other necessary operational parts, include sufficient gas sparger, a baffle(s), Rushton impeller(s), a sampling mechanism and a condensing mechanism. Additional marine impellers for the 3 bioreactors will desirably be included. |  | Compliant (Yes / No) | Insert Response |
| B4 | The three bioreactors MUST each have at least 4 pumps. The pumps would desirably be capable of adjustable speed. Desirably all pumps should be able to operate with more than one tubbing diameter conducive to different flow rates for the same operation speed of the pump. Desirably there would be a sufficient number of pumps to enable working in continuous mode Desirably there would be the capability to integrate more pumps manageable by the controlling unit. |  | Compliant (Yes / No) | Insert Response |
| B5 | The three bioreactors MUST each have the capability to operate in aerobic and anaerobic conditions depending on user requirements. |  | Compliant (Yes / No) | Insert Response |
| B6 | The three bioreactors MUST have the capability to operate and be controllable both individually and/or in an integrated mode across multiple bioreactors. |  | Compliant (Yes / No) | Insert Response |
| B7 | As well as the required ports for feeding, sparger, thermal sensor etc the bioreactors MUST each have at least 3 x 12 mm ports for probes to operate simultaneously. Desirably there should be capability for addition of further probe ports to each bioreactor. |  | Compliant (Yes / No) | Insert Response |
| B8 | The three bioreactors MUST be capable of programming specific temperature, agitation, pH, feeding rate, and at least two sparged gases using a PC or equivalent. |  | Compliant (Yes / No) | Insert Response |
| B9 | For each of the three bioreactors, online measurement, real-time display, and data logging and recording of data collected by the installed probes and impeller (stirrer) speed, with direct measurement of the impeller’s rotational speed MUST be possible and the capability included. |  | Compliant (Yes / No) | Insert Response |
| B10 | The three bioreactors MUST be capable of monitoring and controlling the level of foam in the bioreactor vessels. |  | Compliant (Yes / No) | Insert Response |
| B11 | For each of the three bioreactors, pH measurements MUST be possible within a range of pH 2 to 12 with an accuracy of ± 0.1 through a provided probe. |  | Compliant (Yes / No) | Insert Response |
| B12 | The three bioreactors MUST each be capable of measurement of DO2 (range 0 to 100%, accuracy ±0.1%) through a provided probe. |  | Compliant (Yes / No) | Insert Response |
| B13 | At least one bioreactor MUST be able to measure the Oxidation-Reduction Potential (ORP) (range between –1500 mV and 1500 mV) through a provided probe. |  | Compliant (Yes / No) | Insert Response |
| B14 | The three bioreactors MUST each be capable of a measurement range of Temperature (0 to 50°C, accuracy ±0.1°C) through a provided probe. |  | Compliant (Yes / No) | Insert Response |
| B15 | For the three bioreactors it MUST be possible to calibrate all the probes with any necessary equipment provided by the supplier. Desirably, at least two-point calibration of pH and DO2, Optic density and ORP sensors in parallel across all the bioreactors. |  | Compliant (Yes / No) | Insert Response |
| B16 | A fully assembled bioreactor with all probes and sensors MUST be able to withstand a full autoclaving cycle (121 oC for 30 mins). Where not possible, an alternative sterilisation strategy and details MUST be provided. Additional marks will be awarded for probes with longer lifespans without loss of accuracy, calibration stability, or functional integrity e.g. measured by withstanding a larger number of autoclaving cycles |  | Compliant (Yes / No) | Insert Response |
| B17 | The three-bioreactor platform MUST provide effective control of ORP (if in use) and DO2 values to set-point using a combination of sparged gases and agitation. Please confirm and provide details of control strategy. |  | Compliant (Yes / No) | Insert Response |
| B18 | Auto PID control and counters for quantifying consumable usage (acid/alkali) is highly DESIRABLE for each of the three bioreactors. |  | Compliant (Yes / No) | Insert Response |
| B19 | For the three-bioreactor platform it is DESIRABLE to have the capability to achieve effective kLa (volumetric mass transfer coefficient) and oxygen mass transfer rates during operations. |  | Compliant (Yes / No) | Insert Response |
| B20 | All 3 bioreactors MUST be pre-prepared for off-gas sensor connectivity. At least one of the bioreactors MUST be equipped with an off-gas detection capability, able todetect CO2, O2. Desirably the bioreactor with this capability should be flexible for detection of pressure/humidity and of other off-gases such as CH4 or EtOH. Any required additional software for control of and data collection from this capability MUST be provided. |  | Compliant (Yes / No) | Insert Response |
| B21 | Tenderers MUST provide details of the physical dimensions and weight of the equipment. The minimum/maximum distance between the bioreactor vessel and controller, and between two bioreactors SHOULD be recommended. |  | Compliant (Yes / No) | Insert Response |
| B22 | Gas and water connections, feed line and sampling connections (e.g. connectors/fittings/tubing, male-female end connections, etc.) MUST be included in the Bioreactor installation set. Please confirm and provide details with end connections, flow-rate/pressure requirements, etc. |  | Compliant (Yes / No) | Insert Response |
| B23 | A comprehensive solution is DESIRABLE for the installation of the complete bioreactor platform (i.e. all three bioreactors, control units, computers, etc). Please confirm and provide optimum layout for configuration of all bioreactors on workbenches. |  | Compliant (Yes / No) | Insert Response |
| B24 | The platform MUST be provided with detailed SOPs for cleaning and sterilizing the bioreactors and any other part in direct contact with the fermentation. |  | Compliant (Yes / No) | Insert Response |
| System Software Requirements |  |  |  |  |
| B25 | It MUST be possible to control multiple bioreactors through a single interface which should be supplied with the required software installed. The software MUST be easy to use, accessible, and allow authorised users to securely access the system remotely while it is running. It is desirable that authorised users can also securely access the system remotely. Remote access to the controller unit of the bioreactor(s) should be possible. A user friendly interface MUST be provided and installed that allows the simultaneous control of, and data collection from, multiple bioreactors. Administrator access privileges MUST be provided for any hardware supplied, where applicable. |  | Compliant (Yes / No) | Insert Response |
| B26 | If a dedicated PC or workstation is required for the operation, monitoring, control, data acquisition, or configuration of the bioreactors, it shall be supplied by the vendor as part of the complete system. The supplied PC shall be fully configured, tested, and compatible with all required hardware and software components |  | Compliant (Yes / No) | Insert Response |
| B27 | The Software must allow control, recording and analysis of fermenter parameters including pH, ORP, DO2, temperature, gas readouts, etc. The associated licenses MUST be provided and remain valid for the lifetime of the precision fermentation platform. Please confirm and detail. |  | Compliant (Yes / No) | Insert Response |
| B28 | Software MUST allow users to export data and results to word processor or spreadsheets. Please confirm and provide details. |  | Compliant (Yes / No) | Insert Response |
| B29 | The system MUST include licenses for any proprietary software required for system operation that permits continued use without recurring subscription fees. Please confirm and detail. |  | Compliant (Yes / No) | Insert Response |
| B30 | Capability to provide configurational freedom — connect any actuator to any sensor input and create a new control loop when required is DESIRABLE. A provision to connect and integrate any third-party sensors in the bioreactor, subject to capability to connect, is DESIRABLE. Please confirm and provide details. |  | Compliant (Yes / No) | Insert Response |
| B31 | All software updates MUST be provided by the Tenderer throughout warranty. All software updates, which become available during the lifetime of the system, MUST be provided free of charge. Please confirm and detail. |  | Compliant (Yes / No) | Insert Response |

| - After Sales and Technical Assistance |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Customer Support, Repair, Maintenance & Breakdown Requirements |  |  |  |  |  |
| C1 | The tenderer MUST have dedicated specialist at their disposal and fully qualified service personnel who can respond to service calls. | Compliant (Yes / No) |  | Insert Response |  |
| C2 | The Tenderer MUST provide details of guaranteed response times in event of equipment issues or breakdown both in relation to phone support, online/remote support and arrival on site by a fully trained service engineer during the warranty period (s) and post-warranty. Please detail. | Compliant (Yes / No) |  | Insert Response |  |
| C3 | Please outline the company customer service plan and your escalation policy including what process is in place for escalating a call-out, if a breakdown cannot be resolved on the first visit. | Compliant (Yes / No) |  | Insert Response |  |
| C4 | The Tenderer SHOULD provide details of guaranteed response times in event of equipment breakdown both in relation to phone support, online/remote support and arrival on site by a fully trained service engineer during the warranty period (s) and post-warranty. Please detail. | Compliant (Yes / No) |  | Insert Response |  |
| Warranty & Preventative Maintenance |  |  |  |  |  |
| C5 | At least one year’s warranty MUST be included as part of the tendered price, to include all hardware, software, call outs, Travel, and Parts. Detail specific inclusions and exclusions of warranty. Please confirm and detail any additional pricing in Appendix 2, Pricing Schedule. |  | Compliant (Yes / No) |  | Insert Response |
| C6 | The Tenderer SHOULD state the details of any extended Warranty. Detail specific inclusions and exclusions of extended warranty for up to 3 years optional to purchase warranties on a yearly basis. Please include price per annum, details of coverage and maximum term of coverage. Please confirm and detail and any additional cost in the Appendix 2 – Pricing Schedule. |  | Compliant (Yes / No) |  | Insert Response |
| C7 | The Tenderer must provide details of the servicing and maintenance arrangements available for the proposed system. Responses should clearly identify the scope of the service, including all inclusions and exclusions. Please provide details of: - The recommended servicing and maintenance schedule. - Preventative maintenance activities included. - Calibration, performance verification, and system health checks, where applicable. - Software and firmware updates included as part of the service. - Coverage of replacement parts, consumables, labour, and travel costs. - Technical support arrangements and response times. - Any exclusions, limitations, or additional chargeable services. Please also provide details of the available servicing and maintenance packages, including associated costs and contract options. |  | Compliant (Yes / No) |  | Insert Response |
| C8 | Please describe the preventative maintenance activities that can be undertaken by the end user to maximise system availability, reliability, and uptime. Your response should include: - Routine maintenance tasks that can be performed by trained users. - Recommended maintenance frequency and procedures. - Components that are user-serviceable or user-replaceable. - Diagnostic tools, self-monitoring features, or system health checks available. - Cleaning, calibration, and performance verification procedures. - Training and documentation provided to support preventative maintenance activities. - Any maintenance activities that must be performed exclusively by the supplier or an authorised service engineer. Please explain how these preventative maintenance measures contribute to minimising downtime and maintaining optimal system performance over the lifetime of the equipment. |  | Compliant (Yes / No) |  | Insert Response |
| C10 | The tender MUST specify the lead time for repairs during the warranty period(s) and post-warranty, with and without service contract options. |  | Compliant (Yes / No) |  | Insert Response |

| Training Proposal |  |  |  |
| --- | --- | --- | --- |
| C11 | Training for at least 6 personnel MUST be included after commissioning to cover set up / system familiarization, operation of the tendered equipment, troubleshooting, maintenance, and associated software once the system is in place. (Date(s) and timing of training to be confirmed with TUDublin and the winning tenderer). Please confirm and detail. | Compliant (Yes / No) | Insert Response |
| C12 | Please describe the technical learning and user support resources available for the proposed system. Your response should include: • Initial user training provided at installation and commissioning. • Availability of refresher training or advanced user training. • Access to user manuals, application guides, tutorials, webinars, and other learning resources. • Availability of online knowledge bases, training portals, or e-learning platforms. • Application support for experimental design, image acquisition, and data analysis. • Access to technical specialists or application scientists for user queries. • Details of any ongoing user support programmes, workshops, or user communities. Please explain how these resources will support users in developing proficiency with the system and maximising the value and utilisation of the equipment over its operational lifetime. | Compliant (Yes / No) | Insert Response |
| - Delivery Date or Date of Completion |  |  |  |
| Lead Time & Delivery/Completion Plan |  |  |  |
| D1 | Delivery MUST include all Shipping Charges, insurance, customs duties and all aspects as outlined in DDP (Delivery Duty Paid) incoterms. | Compliant (Yes / No) | Insert Response |
| D2 | Tenderers MUST give details of lead time for delivery, installation, commissioning (if required), and training from the date the Purchase Order is received (In weeks). Please confirm and detail | Compliant (Yes / No) | Insert Response |
| Installation, Commissioning & Validation |  |  |  |
| D3 | The instrument MUST be installed (if required) and commissioned in a location defined by the Contracting Authority, by an appropriately qualified engineer. Please confirm, providing details of experience/qualifications of installation team. | Compliant (Yes / No) | Insert Response |
| D4 | The installation of all electrical items MUST be in accordance with but not limited to European standards/ regulations and all relevant amendments. | Compliant (Yes / No) | Insert Response |
| D5 | All instrument documentation relating to installation of hardware and software, hardware/software interfacing, service and maintenance, parts list, and instrument operation MUST be supplied in English language. | Compliant (Yes / No) | Insert Response |

Sustainability

| D6 | Tenderers SHOULD make all reasonable efforts to minimise adverse environmental impact in the methods of services delivery and in materials used. What actions have you taken to reduce your impact on the environment? Please confirm and detail. | Compliant (Yes / No) | Insert Response |
| --- | --- | --- | --- |
| D7 | It is DESIRABLE that all Tenderers dispose of all Packaging & Waste, using environmentally friendly sources, after installation and commission has been completed. Please confirm and detail. | Compliant (Yes / No) | Insert Response |
| D8 | Tenderers MUST state the energy rating of the proposed system; the heat level generated and the range of environmental conditions in which it will operate and the energy consumption per hour. The system SHOULD ideally have low running costs allowing the equipment to be as energy and economically efficient as possible. Please confirm and detail. | Compliant (Yes / No) | Insert Response |

| - Ultimate Cost |  |
| --- | --- |
| E1 | Please see Appendix 2 – Pricing Schedule for details. |

This form must be completed and signed by an authorised officer of the tenderers organisation.

- This offer will remain open for acceptance by you for a period of 12 months from the date of deadline for submission of tenders.

- We acknowledge that you are not obliged to accept the lowest or any offer.

- I/We undertake to deliver the products/services in accordance with the terms and conditions of the tender specification.

- I/We undertake to maintain full confidentiality with regard to all aspects of this tender process

| Signed: |  | Position: |  |
| --- | --- | --- | --- |
| Print Name: |  | Phone No: |  |  |
| Company Name: |  | Date: |  |  |
| Address: |  |  |  |  |
| Email: |  |  |  |  |

Lot 2 - Light module for Bioreactor

| (Suppliers to state model type here) |  |  |  |  |
| --- | --- | --- | --- | --- |
| REF. |  | REQUIREMENT | CONFIRM (YES/NO) | COMMENT |
| B.Technical Merit |  |  |  |  |
| Functional Requirements |  |  |  |  |
| B1 | The proposed instrument MUST be suitable for usage as a light module for the bioreactor Solaris Jupiter 2.0 Pressure (serial number JP-2.0-R02-22-39) |  | Compliant (Yes / No) | Insert Response |
| B2 | The proposed instrument MUST be able to operate in an independently and controlled manner. |  | Compliant (Yes / No) | Insert Response |
| B3 | The proposed instrument MUST have Photon flux density and other light parameters adequate for growth of unicellular algae. Light intensity Must be dimmable from 0–100%. Circadian cycle simulation MUST be possible. |  | Compliant (Yes / No) | Insert Response |
| B4 | It is DESIRABLE that light intensity is able to reach up to 3000 μmol (photon)/m2. |  | Compliant (Yes / No) | Insert Response |
| B5 | It is DESIRABLE that the proposed instrument have at least 3 vertical bars with white dynamic LEDs each. Each LED should consist of at least four chips: two dedicated to warm white light (2700 K) and two for cool white light (6500 K) - LED total power: 180 W - LED brightness: 6120 lm |  | Compliant (Yes / No) | Insert Response |

| - After Sales and Technical Assistance |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Customer Support, Repair, Maintenance & Breakdown Requirements |  |  |  |  |  |
| C1 | The tenderer MUST have dedicated specialist at their disposal and fully qualified service personnel who can respond to service calls. Please confirm and detail. | Compliant (Yes / No) |  | Insert Response |  |
| C2 | The Tenderer MUST provide details of guaranteed response times in event of equipment issues or breakdown both in relation to phone support, online/remote support and arrival on site by a fully trained service engineer during the warranty period (s) and post-warranty. Please detail. | Compliant (Yes / No) |  | Insert Response |  |
| C3 | Please outline the company customer service plan and your escalation policy including what process is in place for escalating a call-out, if a breakdown cannot be resolved on the first visit. | Compliant (Yes / No) |  | Insert Response |  |
| C4 | The Tenderer SHOULD provide details of guaranteed response times in event of equipment breakdown both in relation to phone support, online/remote support and arrival on site by a fully trained service engineer during the warranty period (s) and post-warranty. Please detail. | Compliant (Yes / No) |  | Insert Response |  |
| Warranty & Preventative Maintenance |  |  |  |  |  |
| C5 | At least One year warranty MUST be included as part of the tendered price, to include all hardware, software, call outs, Travel, and Parts. Detail specific inclusions and exclusions of warranty. Please confirm and detail any additional pricing in Appendix 2, Pricing Schedule. |  | Compliant (Yes / No) |  | Insert Response |
| C6 | The Tenderer MUST state details of yearly servicing included. Detail specific inclusions and exclusions of the yearly servicing included. Please confirm pricing of yearly servicing. |  | Compliant (Yes / No) |  | Insert Response |
| C7 | The Tenderer must provide details of the servicing and maintenance arrangements available for the proposed system. Responses should clearly identify the scope of the service, including all inclusions and exclusions. Please provide details of: • The recommended servicing and maintenance schedule. • Preventative maintenance activities included. • Calibration, performance verification, and system health checks, where applicable. • Software and firmware updates included as part of the service. • Coverage of replacement parts, consumables, labour, and travel costs. • Technical support arrangements and response times. • Any exclusions, limitations, or additional chargeable services. Please also provide details of the available servicing and maintenance packages, including associated costs and contract options. |  | Compliant (Yes / No) |  | Insert Response |
| C8 | Please describe the preventative maintenance activities that can be undertaken by the end user to maximise system availability, reliability, and uptime. Your response should include: • Routine maintenance tasks that can be performed by trained users. • Recommended maintenance frequency and procedures. • Components that are user-serviceable or user-replaceable. • Diagnostic tools, self-monitoring features, or system health checks available. • Cleaning, calibration, and performance verification procedures. • Training and documentation provided to support preventative maintenance activities. • Any maintenance activities that must be performed exclusively by the supplier or an authorised service engineer. Please explain how these preventative maintenance measures contribute to minimising downtime and maintaining optimal system performance over the lifetime of the equipment. |  | Compliant (Yes / No) |  | Insert Response |
| C9 | The tender MUST specify the lead time for repairs during the warranty period(s) and post-warranty, with and without service contract options. |  | Compliant (Yes / No) |  | Insert Response |

| Training Proposal |  |  |  |
| --- | --- | --- | --- |
| C10 | Training for at least 6 personnel MUST be included after commissioning to cover set up / system familiarization, operation of the tendered equipment, troubleshooting, maintenance, and associated software once the system is in place. (Date(s) and timing of training to be confirmed with TU Dublin and the winning tenderer). Please confirm and detail. | Compliant (Yes / No) | Insert Response |
| C11 | Please describe the technical learning and user support resources available for the proposed system. Your response should include: • Initial user training provided at installation and commissioning. • Availability of refresher training or advanced user training. • Access to user manuals, application guides, tutorials, webinars, and other learning resources. • Availability of online knowledge bases, training portals, or e-learning platforms. • Application support for experimental design, image acquisition, and data analysis. • Access to technical specialists or application scientists for user queries. • Details of any ongoing user support programmes, workshops, or user communities. Please explain how these resources will support users in developing proficiency with the system and maximising the value and utilisation of the equipment over its operational lifetime. | Compliant (Yes / No) | Insert Response |
| - Delivery Date or Date of Completion |  |  |  |
| Lead Time & Delivery/Completion Plan |  |  |  |
| D1 | Delivery MUST include all Shipping Charges, insurance, customs duties and all aspects as outlined in DDP (Delivery Duty Paid) incoterms. | Compliant (Yes / No) | Insert Response |
| D2 | Tenderers MUST give details of lead time for delivery, installation, commissioning (if required), and training from the date the Purchase Order is received (In weeks). Please confirm and detail | Compliant (Yes / No) | Insert Response |
| Installation, Commissioning & Validation |  |  |  |
| D3 | The instrument MUST be installed (if required) and commissioned in a location defined by the Contracting Authority, by an appropriately qualified engineer. Please confirm, providing details of experience/qualifications of installation team. | Compliant (Yes / No) | Insert Response |
| D4 | The installation of all electrical items MUST be in accordance with but not limited to European standards/ regulations and all relevant amendments. | Compliant (Yes / No) | Insert Response |
| D5 | All instrument documentation relating to installation of hardware and software, hardware/software interfacing, service and maintenance, parts list, and instrument operation MUST be supplied in English language. | Compliant (Yes / No) | Insert Response |

Sustainability

| D6 | Tenderers SHOULD make all reasonable efforts to minimise adverse environmental impact in the methods of services delivery and in materials used. What actions have you taken to reduce your impact on the environment? Please confirm and detail. | Compliant (Yes / No) | Insert Response |
| --- | --- | --- | --- |
| D7 | It is DESIRABLE that all Tenderers dispose of all Packaging & Waste, using environmentally friendly sources, after installation and commission has been completed. Please confirm and detail. | Compliant (Yes / No) | Insert Response |
| D8 | Tenderers MUST state the energy rating of the proposed system; the heat level generated and the range of environmental conditions in which it will operate and the energy consumption per hour. The system SHOULD ideally have low running costs allowing the equipment to be as energy and economically efficient as possible. Please confirm and detail. | Compliant (Yes / No) | Insert Response |

| - Ultimate Cost |  |
| --- | --- |
| E1 | Please see Appendix 2 – Pricing Schedule for details. |

This form must be completed and signed by an authorised officer of the tenderers organisation.

- This offer will remain open for acceptance by you for a period of 12 months from the date of deadline for submission of tenders.

- We acknowledge that you are not obliged to accept the lowest or any offer.

- I/We undertake to deliver the products/services in accordance with the terms and conditions of the tender specification.

- I/We undertake to maintain full confidentiality with regard to all aspects of this tender process

| Signed: |  | Position: |  |
| --- | --- | --- | --- |
| Print Name: |  | Phone No: |  |  |
| Company Name: |  | Date: |  |  |
| Address: |  |  |  |  |
| Email: |  |  |  |  |