Single Party Framework Agreement for Waterways Ireland’s Hydrometric Measurement Network & associated services

Scope of Service

| # Document Control Sheet |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| WI Section: |  | Technical Services |  |  |  |  |  |
| Project Title: |  | Single Party Framework Agreement for Waterways Ireland’s Hydrometric Measurement Network & associated services |  |  |  |  |  |
| Document Title: |  | Scope of Service |  |  |  |  |  |
| Project No: |  |  |  |  |  |  |  |
| Text Pages: |  |  |  | Appendices: |  |  |  |
| Rev | Status | Date | Author(s) |  | Reviewed By |  | Approved By |

Waterways Ireland

2 Sligo Road

Enniskillen

Co Fermanagh

BT74 7JY

Tel: +44 (0) 28 6632 3004

Fax: +44 (0) 28 6634 6257

E: info@waterwaysireland.org

Contents

1.0	Introduction	1

2.0	Background	2

3.0	Scope of requirements under the framework	4

3.1	Summary specification of the initial Orders	4

3.2	Details of Contracts Arising Over Life of Framework Agreement	4

3.3	Parties’ Representative	5

3.3.1	WATERWAYS IRELAND’S Representative/ project manager	5

3.3.2	The Service Provider’s Representative	5

3.4	On-going Communications	5

3.4.1	Liaison between Employer and the Service Provider	5

3.4.2	Liaison between Waterways Ireland Operations and the Service Provider	5

3.4.3	Liaison with the General Public	6

3.5	Progress meetings	6

3.6	Health and Safety	6

3.6.1	General	6

3.6.2	Accidents and Dangerous Occurrences	7

3.6.3	Health and Safety Audits	7

3.7	Quality Management	7

3.8	Innovation, Environmental and Sustainability Management	9

4.0	Detailed scope of services under the initial Orders	9

4.1	Order 1 - Annual Servicing Requirements	9

4.1.1	General	10

4.1.2	Existing Equipment	11

4.1.3	Site Access	11

4.1.4	Servicing checklist	12

4.1.5	Call outs for onsite diagnostic testing and device review	13

4.1.6	Spares	14

4.1.7	Stock Storage	14

4.1.8	KPI’s	14

4.2	Order 2 - Web hosting platform	15

4.2.1	Overview of requirements	15

4.2.2	Specific Requirements	15

4.2.2.1	Information management	15

4.2.2.2	Data Hosting	15

4.2.2.3	Data transmission	15

4.2.2.4	Ease of use	16

4.2.2.5	Domain name	16

4.2.2.6	Maintaining the platform	16

4.2.2.6.1	Retrieving lost data	16

4.2.2.6.2	Platform interuptions	16

4.2.2.7	Security	17

4.2.3	Device Compatiblity	17

4.2.4	Data Transfer	17

4.2.5	KPI’s	17

5.0	Replacement device and equipment specification	18

5.1.1	general specficiation	18

5.1.1.1	Design Life	18

5.1.1.2	Device Longivity and Warranty	18

5.1.1.3	Ingress protection	18

5.1.1.4	Materials	18

5.1.1.5	CE mark	18

5.1.1.6	Power source and Energy Efficiency	19

5.1.1.7	End of life service	19

5.1.2	Water level recorder	19

5.1.2.1	Water level device Type 1	20

5.1.2.2	Water level device type 2	21

5.1.3	Fixed flow velocity and water level recorder	21

5.1.4	Weather station	22

5.1.5	Data logger	23

5.1.6	Rittal Cabinets	23

5.1.7	Security Cage	23

6.0	Scope of services for future Orders	24

6.1	Relocation or new device installation	24

6.2	Provision of training	24

6.3	Sustainability improvements	24

6.4	Maintenance, data collection, procurement and /or installation of smart controllers, smart pumps and associated water level and flow devices and ancillaries	24

6.5	Project specific monitoring, testing, data collection and analysis	24

6.6	Spot Flow Measurement	25

6.7	Topographical datum’s collection of the network	25

Introduction

Waterways Ireland is one of six North-South Implementation Bodies, established under the British-Irish Agreement Act of 1999. We are the cross-border navigational authority responsible for the management, maintenance, development, and restoration of 1,100 km of inland navigable waterways principally for recreational purposes. Our Headquarters are in Enniskillen Co. Fermanagh, with regional offices in Dublin, Carrick-on-Shannon, Co. Leitrim, and Scariff, Co. Clare. We employ 350 permanent staff located at offices and sites close to the inland waterways who are assisted by a team of seasonally recruited staff reflecting the seasonality of the organisation’s remit. Waterways Ireland creates social, economic, and environmental well-being to the value of €600m annually.

Waterways Ireland is the navigation authority for 1,200 km of navigable inland waterways, comprising:

- Lower Bann Navigation

- Erne System

- Shannon-Erne Waterway (SEW)

- Shannon Navigation

- Royal Canal

- Grand Canal

- Barrow Navigation

- Ulster Canal (Upper Lough Erne to Clones)

In July 2007, it was agreed by the North/South Ministerial Council (NSMC) to include responsibility for the reconstruction of the Ulster Canal from Upper Lough Erne to Clones, and following restoration, its management, maintenance, and development, principally for recreational purposes to Waterways Ireland’s remit.

At an operational level, Waterways Ireland is responsible for a vast range of infrastructure assets including navigation channels, embankments, towpaths, adjoining lands, harbours, jetties, fishing stands, bridges, culverts, aqueducts, overflows, locks, sluices, and lock houses along with buildings and archives. The current valuation of the rebuild costs of this infrastructure is estimated at €809 million. Waterways Ireland’s work programs are critical to providing a safe and high-quality recreational environment for customers, whilst preserving the industrial and environmental heritage of the waterways for future generations. Our Purpose is:

To be the custodian of the inland navigations and collaborate to reimagine, maintain, develop and promote them to sustain communities, environment and heritage.

Our Vision is to bring our waterways to life – full of recreational and commercial activities, havens rich in wildlife and biodiversity, and positively transformational both in terms of enriching people's lives and regenerating the many rural and urban communities and regions through which the waterways pass. Our Vision is:

Creating inspirational inland navigations and waterways experiences through conservation and sustainable development for the benefit of all.

To realise our Vision, Waterways Irelands focus remains on achievement of six strategic priorities based on extensive user and stakeholder engagement following the completion of our Long-term Plan:

- Organisation Development & Governance

- Sustainable Funding Model

- Asset Portfolio Management

- Participation & Reputation

- Sustainable Development

- Climate Action, Environment & Heritage

Our Values, as behavioural guidance, are critically important to the organisation as they are essential to the successful implementation of the Long-Term Plan as well as driving cultural change and development in the organisation

Further information is available at our corporate website www.waterwaysireland.org

Background

Waterways Ireland’s current hydrometric network comprises of 133 (with four additional units expected to be installed prior to award of this contract) water level, weather and flow monitoring devices installed across seven inland navigable waterways. Most of the instrumentation forming the existing network has been supplied by one manufacturer OTT HydroMet (OTT). A number of Side Looking Doppler flow monitors manufacture by Xylem have also been procured. Each device remotely transmits data to the existing web-hosting platform Hydromet Cloud, which is also currently supplied by OTT and is stored for a minimum of three years

The geographical spread of the devices are indicated in Figure 1 below, each green marker representing an individual water monitoring instrument. The network comprises of instruments which are in both NI and in the RoI.

Fig.1

The existing instrumentation network has been installed since 2009 and the current asset value of the network is in the region of €1,000,000 (one million Euros).

Scope of requirements under the framework

The scope of the Framework is for the servicing, maintenance and web-hosting of the Waterways Ireland existing hydrometric network and is further set out below. It is intended to establish the Framework Agreement on foot of this competition. At the point of establishing the Framework Agreement two (2) initial Orders shall also be accepted. Details of the requirements for these Orders are detailed within this document along with an overview of potential future Orders.

Summary specification of the initial Orders

The two (2) initial Orders shall be as follows: -

- Annual Servicing -Service and Supply of hydrometric devices and their ancillaries to ensure the existing hydrometric network is maintained. This initial Order and thus the associated Contract there after shall be for a 4 year duration and shall include;

- Calibration, repair and maintenance of the existing network of OTT instruments;

- Provide and manage a system that allows each device (existing and new) to collect and store the associated data as well as transmit it remotely at varying intervals to a web-hosting platform;

- Supply, install, calibrate and maintain replacement water level devices, water temperature recorders, flow devices and weather stations;

- Provide a call out service, which includes out of hours in order to rectify a network failure or device failures.

- Liaise with WI in regards to the surrounding maintenance and access to device.

- Web-hosting platform - Provide and manage a web-hosting platform for hydrometric data. This initial Order and thus the associated Contract there after shall be for a 4 year duration, exact start date and as such the duration of this Order is subject to who the supplier of web-hosting system is due to a potential cross over of existing arrangements. Under this Order the Service Provider shall:

- Provide and manage a web-hosting platform. This platform will be required to, as a minimum, store and monitor hydrometric data and to raise alarms when abnormally high or low water level or flow events occur;

- Migrate the existing historical data so it can be stored in the new web-hosting platform.

In relation to the above, the Service Provider shall carry out any work (including any construction work) whether of a permanent or temporary nature and provide everything necessary and required for performance of the Order(s). Furthermore, due to the nature of the hydrometric network the Service Provider may be required to supply their own boat to facilitate the installation, calibration and maintenance of the devices.

Details of Contracts Arising Over Life of Framework Agreement

In addition to the initial Order(s), the Framework Agreement may also be used, in particular, for the following types of Orders:

- Supply/ relocate, install, calibrate and maintain new water level devices, water temperature recorders, flow devices and weather stations;

- Supply, install, calibrate and maintain new environmental/hydrometric devices to help meet WI changing requirements;

- Provide training to Waterways Ireland employees regarding any device existing or new and/or the web-hosting platform;

- Sustainability improvements to the existing network;

- Maintenance, data collection, procurement and /or installation of smart controllers, smart pumps and associated water level and flow devices and ancillaries, likely with different requirements to those detailed in the Scope of Services.

- Project specific monitoring, testing and data collection: structural, environmental and/or hydrological (e.g, spot flow measurement);

- Analysis of data for water, structural or environmental management and/or research purposes;

- Check existing topographical survey device datum’s of the network; and

- Any other related device installation, monitoring or reporting that is required to meet WI changing requirements.

There is no guarantee of Orders, but where they arise, the Contracting Authority will consult with the Service Provider under and in accordance with the Framework Agreement.

Parties’ Representative

WATERWAYS IRELAND’S Representative/ project manager

The Waterways Ireland Representative/ Project Manager for the project is Kevin Macken, Waterways Ireland, Ashtown Gate, Navan Road, Dublin 15

The Service Provider’s Representative

The Service Provider will appoint a Framework Manager to be a primary point of contact between Waterways Ireland and the Service Provider.

On-going Communications

Liaison between Employer and the Service Provider

The Service Provider shall make themselves available for regular meetings and discussions throughout the Framework Period. They will be required to attend formal progress meetings as identified in Section 3.5.

The Service Provider should allow for one project meeting every 6 months  in person in Waterways Ireland HQ, Enniskillen on in Waterways Ireland office in Ashtown, Dublin 15. Any additional meetings/discussions will be held via telephone conference between each party, as and when required by Waterways Ireland.

Liaison between Waterways Ireland Operations and the Service Provider

Under each Order the Service Provider is required to liaise directly with Waterways Ireland’s operational staff in each region as necessary in order to gain access to specific site locations. On award of an Order a list of relevant WI personnel will be provided to the Service Provider and shall be updated as and when staff changes are made within Waterways Ireland.

Liaison with the General Public

It is not anticipated that the Service Provider will have to liaise with the general public. If the Service Provider or any of its Personnel are approached on site they must conduct themselves in a professional manor as they will be representing Waterway Ireland. Should any issues arise on-site they must inform the Project Manager as soon as it is safe to do so.

Progress meetings

The Service Provider will be required to attend biannual progress meeting with Waterways Ireland. It is anticipated that these meetings will be held in person at the Waterways Ireland Headquarters in Enniskillen or in Waterways Ireland offices in Ashtown, Dublin 15. The Service Provider shall be required to prepare and submit a progress report not later than 3 days before the scheduled meeting detailing what works have been complete to date under each Order, what issues are still outstanding and what works are planned for the following 6 months. The typical agenda for each meeting is as follows:

- Previous Minutes

- Progress/Programme (Servicing and other works undertaken during the framework)

- Health and Safety

- Communications

- Quality

- KPI review and reporting

- Commercial

- Any other business

These meetings are intended to ensure that the Framework Agreement is working for both parties and the Service Provider is encouraged to report any issues or difficulties they have encountered to efficiently find a solution.

Health and Safety

General

The Service Provider must adhere to all relevant Health and Safety legislation and regulations for the area they are working in including but limited to:

- the Safety, Health & Welfare at Work Act, 2005 and Safety, Health and Welfare at Work (Construction) Regulations 2013 for the Republic of Ireland; and

- Health and Safety at Work (Northern Ireland) Order 1978 and The Construction (Design and Management) Regulations (Northern Ireland) 2016 for Northern Ireland.

The Service Provider must regularly update and supply Waterways Ireland with copies of:

- their corporate health and safety policy

- method statements and risk assessments for annual servicing including any lone working policies; and

- any specific method statements and risk assessments for any future Orders entered into under the Framework Agreement.

Accidents and Dangerous Occurrences

The Service Provider shall immediately advise Waterways Ireland verbally, and afterward in writing of any accident or dangerous occurrence, loss or damage to any personnel, third party or property arising during the execution of their duties in relation to the Framework Agreement and/or any Order. Such accident reporting to Waterways Ireland shall not relive the Service Provider of their statutory responsibilities for accident reporting to the relevant statutory authorities and/or insurers.

Health and Safety Audits

The Service Provider shall permit Waterways Ireland’s Health and Safety staff access to their and their Personnel’s offices, sites etc. at any reasonable time; permit these Health and Safety staff to make any reasonable enquiry; and provide an appropriate information/data requested relevant thus enabling Waterways Ireland to undertake any health and safety audit, inspection, or investigation in assessing health and safety compliance.

Quality Management

The Service Provider shall institute and operate a quality management system complying with IS EN ISO 9001 or be able to produce an internally approved quality management plan of equivalent and suitable quality.

The Quality Plan shall cover the following items:

- Service Provider’s Organisation and Management.

- Service Provider’s Method Statements and Servicing Procedures.

- Service Provider’s Quality Control.

- Supplier’s Quality Plans

(for each of the quality management schemes listed below).

Quality Plans shall conform with the requirements as follows:

- Service Provider’s Organisation and Management.

This section of the Quality Plan shall include:

- Definition of the Framework Agreement and its documentation.

- The organisation of the Service Provider, including the line of command and communication links between parties involved in the framework.

- Names, roles, responsibilities and authority of the Service Provider’s principals and key personnel.

- Control of liaison and meetings with third parties, if any.

- Identification of the Service Provider’s staff responsible for overseeing each activity.

- The Service Provider’s control of sub-contractors, if applicable.

- Document control.

- Programme for submission of method statements.

The Quality Plan shall identify procedures (which may be a part of the Service Provider’s general procedures) that cover the topics listed below. Copies of these procedures shall be made available to the Waterways Ireland Representative on request.

- The quality plans for sub-contractors and suppliers of services, work, goods and materials which are the subject of quality management schemes.

- Procedure for the preparation, review and adjustment of programmes for the effective progression of Orders and the recording of this.

- Control and approval of purchases of materials.

- Control of off-site activities (where appropriate).

- Procedures for the regular review and recording by the Service Provider of the quality of the Works and/or Services under the Framework and each Order.

- Control of personnel selection, based on their care, skill and experience.

- Management review/audits to monitor and exercise adequate control over the implementation of the Quality Plan.

- Any other relevant item.

- Service Provider’s Method Statements and Works Procedures

This section of the Quality Plan shall include:

        - Detailed method statements for each major activity whether directly controlled or subcontracted. The method statements shall identify hold points and invoke:

- work instructions;

- quality control procedures;

- compliance testing/inspection arrangements;

- and work acceptance procedures.

for all activities that might affect the quality of the Works and / or Services.

        - Identify the relevant procedures in the Service Provider’s own Quality Management System (and provide copies on request).

- Service Provider’s Quality Control.

This section of the Quality Plan shall include:

- Statement of the Service Provider’s organisation for quality control.

The quality plan shall identify procedures (which may be a part of the Service Provider’s general procedures) that cover the topics listed below. Copies of these procedures shall be made available to the Employer’s Representative on request.

- Document control.

- Procedure for monitoring and recording the inspection, test and approval status of Work.

- Procedure for the collation of quality records as identified in BS EN ISO 9001, and provision of copies when requested by the Waterways Ireland Representative.

- Supplier’s Quality Plans.

The Quality Plan shall include:

- Definition of the elements of the Works and/or Services to be provided by a sub-contractor or supplier.

- The organisation of the sub-contractor or supplier describing the line of command and stating the name of the senior manager responsible for the relevant Works and/or Services. Contact addresses, telephone numbers etc shall be provided.

- Identification of the relevant parts of the sub-contractor or supplier’s quality system relevant to the relevant Works and/or Services being provided. (Copies to be provided to the Project Manager on request).

- The control of personnel selection (at works and on site), including special requirements for skilled personnel, e.g. training of operatives, experience requirements etc.

Specific procedures for the following:

- Receipt and examination of certificates of conformity and tests results for purchased products.

- Product identification and traceability.

- Handling, storage, packaging and delivery to Site and storage and handling on site.

- Quality records.

- KPI’s

Items (i) and (iii) of the Quality Plan shall be submitted to the Project Manager for its acceptance not later than 21 days after Framework Agreement and shall be updated within 21 days following entry into each Order.

The Service Provider shall submit other parts of the Quality Plan prior to commencement of any related Works and/or Services or appointment of a sub-contractor or supplier.

Method statements will be required for all Orders of Works and/or Services, as reasonably determined by Waterways Ireland.

Innovation, Environmental and Sustainability Management

In line with our values, Waterways Ireland strives to continuously improve in the delivery of quality services and infrastructure. We are passionate about what we do and promote collaboration, innovation, and accountability were possible.

In line with the Environmental Protection Agency’s Green Procurement Guidance and to help achieve Waterways Ireland’s long term net zero goal the Service Provider shall as a minimum comply with both the Directive 2011/65/EU on the restriction of certain hazardous substances in electrical and electronic equipment (RoHs) and the Directive 2012/19/EU on waste electrical and electronic equipment (WEEE Directive).

Waterways Ireland encourages Service Providers to, at its own cost, seek to go above these requirements were possible considering the environmental impact that their organisation and in particular the Works and/or Services they supply will have in the long term.

Detailed scope of services under the initial Orders

Order 1 - Annual Servicing Requirements

General

The Service Provider shall undertake annual servicing of all 137 (133 existing and 4 sites pending installation) number of devices installed in Waterways Ireland’s Navigations. Servicing of devices covers the calibration, repair and/or replacement of each device and must conform to a specific checklist for each device as provided to the Service Provider by Waterways Ireland in Section 4.1.1 and Appendix A. Any new devices installed by the Service Provider during the Framework Period shall also require annual servicing.

Servicing must be complete within a 12-month timeframe between January and the end of December each calendar year. Each device must be serviced within 12months plus/minus 2 weeks of its previous annual service; exceptions may be giving to the second annual servicing allowing it to be undertaken in advance of a full 12 months from the first due to the time restrictions associated with the first annual service. Waterways Ireland’s hydrometric network is divided between the following regions and navigation areas: -

Northern Region – 38 devices

- LB - Lower Bann

- Erne – Erne

- SE - Shannon Erne

- UC – Ulster Canal

Western Region – 29 devices

- NS - North Shannon

- SS -South Shannon

Eastern Region – 70 devices

- BAR – Barrow

- RCR - Royal Canal Rural

- RC -Royal Canal Metro

- GCR - Grand Canal Rural

- GC - Grand Canal Metro

Timeframes and preliminary servicing to be agreed January of each year and a schedule must be supplied by the Service Provider to Waterways Ireland for review and acceptance before servicing commences for the year.

It should be noted that where reference is made to existing devices the Service Provider can propose to replace all devices should they be unable to service the existing branded devices. The cost of any full replacement must be allowed for in their prices if they wish to take this approach and the proposed devices must meet the minimum requirements outlined within this document, a list of device ages, types and locations has been provided in Appendix B for reference.

In addition to the annual servicing, the devices at the following locations require a weekly check via the web-hosting platform to ensure they are active, should they be inactive a call out will be required within 24 hours to physically inspect the relevant device(s) and ensure they are functioning correctly and noted as ‘active’ on the web-hosting platform. Additional devices may be added to this list throughout the duration of the Framework Agreement and/or relevant Order.

- Royal Canal Metro, Spencer Dock – OTT Ecolog

- Royal Canal Metro, Lock 12 U/S – OTT Ecolog

- Royal Canal Metro, Lock 13 –  OTT Orpheus Mini and ITC

- Grand Canal Metro, Grand Canal Dock – OTT Orpheus Mini and ITC

- Grand Canal Metro, Lock 6 -  OTT Orpheus Mini and ITC

- Grand Canal Metro, Lock 18 U/S – OTT Orpheus Mini and ITC

- Grand Canal Metro, Lock C4 – OTT Orpheus Mini and ITC

- South Shannon, Banagher, Lufft Weather Station and NetDL logger

- South Shannon, Meelick Weir, OTT Orpheus Mini and ITC

- South Shannon, River Suck Air Draft, OTT Ecolog

- Ulster Canal, Clones Marina, OTT Ecolog

Existing Equipment

Waterways Ireland network consists of the following types of hydrometric devices:

- Water level and temperature sensors installed at various locations,

- Flow monitoring stations with level and temperature sensors installed at various locations,

- Weather station which includes rainfall, wind speed, wind direction, water level and temperature sensors installed at one location, and

- Data loggers

Waterways Ireland currently has a combination of the following devices installed throughout the 10 navigation areas:

- 108 No. OTT Orpheus mini and ITC level recorders with temperature sensors,

- 22 No. OTT Ecolog 500/1000 water level recorder and temperature sensor,

- One  Side Looking Doppler (SLD) measuring water flow and temperature with OTT netDL logger powered by solar panel,

- One OTT Surface Velocity Radar (SVR) measuring water flow and temperature with OTT netDL logger powered by solar panel,

- One OTT Lufft weather station (wind speed and direction, precipitation rate and rain fall level, pressure sensor probe for water level and temperature recording) with netDL logger powered by solar panel and battery

- Prior to award of this contract, it is expected that 2 No. further OTT Ecolog 500/1000 will be installed and 2 SLD flow recording devices (either SonTek-SL 1500 or an OTT SLD device)

Site Access

The Service Provider shall include in their servicing prices all costs associated with land travel to each of the devices. Several of the devices in the South Shannon and the Erne navigation can only be accessed by power boat due to their location being on a pile in open water, this access can be facilitated by Waterways Ireland provided at least 2 weeks’ notice is given by the Service Provider.

Waterways Ireland will supply one staff member with a RYA level 2 powerboat certificate of compliance. It is the Service Providers responsibility to provide a 2nd crew member with RYA level 2 powerboat certification to ensure Waterways Ireland’s health and safety polices are met. The Service Provider will be required to provide evidence that the required RYA level 2 powerboat training is up to date on a yearly basis.

The Service Provider shall include in their tender price all costs associated with providing a small watercraft capable of acting as a working platform to conduct non intensive tasks associated with servicing and/or repairs as and when required at other locations including but not limited to visual inspections, bracket repair and pipe realignment. It is expected that this would only be a requirement should significant repairs need to be undertaken to ensure pipes or brackets are secure and not expected to be a requirement as part of the standard annual service. It shall also be dependent on the location as Waterways Ireland may be able to facilitate this type of access if a work vessel is in the local area.

Waterways Ireland shall provide the Service Provider with a list of regional contacts upon execution of the Framework Agreement and it shall be the responsibility of the Service Provider to liaise with the local operational contacts about the planned servicing schedule. This is to ensure that the relevant Waterways Ireland operational team has adequate time to ensure a member of staff is available to provide access to sites that are restricted to the general public.

If access to a device is unavailable due to maintenance issues associated with infrastructure outside the remit of the Service Provider, the Service Provider shall provide photographic evidence of the access issue and submit them to Waterways Ireland through the standard checklist.

Servicing checklist

Any issues encountered during the annual servicing of devices must be flagged on the checklist as seen in Appendix A. If a device cannot be serviced on the day a checklist must still be completed and the reason why the service could not take place shall be clearly stated by the Service Provider. This checklist must be submitted within 7 days of the service to Waterways Ireland to ensure all issues are actioned in a timely manner and as required by Waterways Ireland.

As a minimum the following is expected to be undertaken during the servicing of devices:

| Device component | Service task | NOTE |
| --- | --- | --- |
| Water level sensor (accurate to <1cm ref. local RTK benchmark). | Maintain in line with manufacturer’s instructions and calibrate against local datum POINT, Ordnance datum Malin (ODM) | Verify against benchmark |
| Doppler checking to ensure beams are with specified parameters | Maintain in line with manufacturer’s instructions, check all ancillaries and cabling | Check operation |
| Solar Panel | Check integrity and fitness for purpose, check if battery has been charged. | Check operation |
| Data logger | Maintain in line with manufacturer’s instructions, check all ancillaries and cabling | Check operation and save recorded data |
| Mobile network modem and antennae | Maintain in line with manufacturer’s instructions, check all ancillaries and cabling | Check operation and current signal strength |
| Cabling | Check integrity and fitness for purpose, any problems developing | Check operation |
| Power supplies | Check integrity and fitness for purpose, check remaining battery life, any problems developing / replace batteries as required | Check operation |
| Enclosures and penetrations | Check integrity and fitness for purpose, any problems developing / faults | Check operation |
| Pipe work, brackets and security cages | Check integrity and fitness for purpose, any problems developing / faults | Check operation |
| Desiccants | Check integrity and fitness for purpose, replace if required | Verify operation |

NB: not all elements will necessarily be present at all stations. This list is not an exhaustive list and other elements may need to be serviced or reported on depending on the device.

The Service Provider shall ensure that:

- A maintenance / calibration check procedure is provided for each device.

- Detailed records of all activities and observations are recorded through the checklist and provided to Waterways Irelands Project Manager within 7 days of completion.

- A detailed account of the maintenance procedure and what tasks will be carried out as part of the annual service is given for each device type.

- Batteries are only replaced after a battery life check is undertaken and it is proven that there is not sufficient battery remaining to power the device and transmit until the next service.

- All data is downloaded from the devices/loggers, saved in excel or CSV format and supplied to Waterways Ireland following the servicing programme for each navigation. Data must clearly state the device name and time frame for data downloaded.

Details of any added value are supplied to Waterways Ireland, e.g.  Software upgrades, to minimise downtime.

Call outs for onsite diagnostic testing and device review

Waterways Ireland currently has 11 priority stations that are flagged as essential to maintaining the navigations (such stations being those identified in Section 4.1.1 above) and as part of the initial Order it is a requirement that these stations are remotely checked weekly by the Service Provider to ensure they are active; this is currently achieved through the use of an online web hosting platform. Should, upon carrying out a remote weekly check, a priority station become inactive the Service Provider shall ensure that an onsite callout is undertaken within 24hrs of such check in order to identify the cause of the station becoming inactive and carrying out any necessary Works and/or Services to get it functioning correctly once again.

Should a non-priority station become inactive a call out to the location of the relevant device is required by the Service Provider within 7 days of identifying a device as inactive.

In the event that any device becomes inactive within 30 days of being serviced, the Service Provider shall revisit the physical device at their own cost. The Service Provider shall undertake all repairs and or replacement of devices within its Warranty Period at their own cost.

Once it is known that a device is inactive, priority or otherwise the Service Provider must inform Waterways Ireland as soon as possible and before a site visit is undertaken. Once agreed and a site visit is undertaken the Service Provider must provide an update summary to the Project Manager regarding the works undertaken and any further works required within 2 days of the site visit.

The Service Provider shall be responsible for regularly monitoring WI hydrometric network remotely to ensure devices are not inactive for extended periods of time, the minimum monitoring periods are listed below:

- Every Monday for the 11 priority stations

- Every 7 days for the first 30 days following servicing of a device

- Every 3 months for the entire network

Spares

The Service Provider must take spare devices with them during servicing and any call outs to ensure any inactive devices that can’t be repaired onsite are replaced with a new device to prevent the need for additional call outs. It is expected that a minimum of 4 spare devices of each type shall be held at any one time to ensure there is no delay in devices being operational. It will be the Service Providers responsibility to ensure approval for replacement stock is obtained from Waterways Ireland with sufficient lead in times so as to ensure there is no delay in performance of the Service Provider’s obligations.

Stock Storage

The Service Provider should have a secure facility available capable of storing approximately 20 devices at any time on behalf of Waterways Ireland. Upon award of the Framework Agreement Waterways Ireland will transfer all currently held hydrometric stock to the Service Provider with a completed list of stock. It shall be the responsibility of the Service Provider to update the stock list as new devices are purchased under the Framework Agreement and/or a relevant Order and devices are swapped out during servicing and call-outs. An updated stock list shall be issued by the Service Provider to Waterways Ireland at a minimum of 3 month intervals or if the number of devices in stock becomes low and new devices are required to ensure a minimum of 4 of each type of device are held by the Service Provider at any one time.

As a minimum Waterways Ireland require each stock list provided by the Service Provider to contain the following information:

- Device type of each device

- Serial number of each device

- Date of delivery into the Service Providers storage

For the avoidance of doubt, delivery or storing of Goods to the storage facility maintained by the Supplier under this section shall not be delivery to the Contracting Authority for the purpose of Section 3 of the Framework Agreement.

KPIs

The following KPIs are to be reported on in advance of each progress meeting:

- The servicing programme shall not slip by more than 1 week and servicing completion % to be above 97% by the end of the servicing period each year

- All devices should be active at any given time however it is understood this may not be possible and as such at least 90% of devices/ stations must be active at all times

- All call out shall be undertaken within the specified timeframes

Order 2 - Web hosting platform

Overview of requirements

Waterways Ireland requires a secure real-time data access platform that will receive, ingest, decode, process, display and store measurement data from each remote hydrometric monitoring device installed on Waterways Ireland property. Access to the data must be available from both a PC and mobile device. It must be able to be viewed graphically on a map or in tabular, and report forms. The data must also be able to be downloaded (in XLS & CSV format) immediately for internal processing by Waterways Ireland staff.

The proposed web hosting platform must meet the satisfaction of Waterways Ireland (and be based on currently supported technology).

Specific Requirements

Information management

The Service Provider should ensure that the platform retains the ability to plot single or multiple sensor information in graph format at all times. The platform must be capable of providing specific graphs between selected time and date ranges and must be capable of converting these graphs to selected formats i.e. PDF, JPEG for use with associated administrative functions.

The platform must also continue to provide single or multiple sensor data in a report formation. Similar to plotting the graphs the platform must be capable of providing specific reports between selected time and date ranges and must be capable of downloading these reports for use with Microsoft Excel applications.

Data Hosting

The Service Provider must ensure that the platform remains password protected, and the monitoring stations and sensors are only accessible to those that enter the designated Waterways Ireland login credentials. The web hosting platform must be available for access on both desktop and mobile app versions at all times and data updated and synced so that desktop and mobile app versions contain the same data. The platform must be capable of hosting the following sensor information:

    - Water level to Ordnance Datum Malin

    - Flow in m3/s

    - Water temperature,

    - Rainfall in mm.

      - Data transmission

Data from the monitoring stations shall be collected at the following intervals:

- Temperature every 60 Minutes.

- Water Level/ Rainfall/ Flow every 15minutes.

and shall be transmitted via SMS or Data dependent on the most appropriate method at the specific site and updated on the platform at the following intervals:

twice daily (every 12 hours) @ 07:00 GMT and 19:00 GMT for all stations; and

- every hour, on the hour for

- WR-SS-Banagher (Solar)

- WR-SS-Meelick Weir

- ER-GCM-Grand Canal Dock

- ER-RCM-Spencer Dock

- WR-SS-River Suck Air Draft

- WR-NS-Jamestown

Ease of use

The platform must be user friendly and contain the following elements:

- Username and password protected access for multiple accounts including a specific account for admin with permissions to edit device names, station information and access to all transmitted data.

- Waterways Ireland logo on the main platform page.

- Map based background with location, name of device and last transmitted data available.

- Selection of devices by navigation.

- Ability to set up individual dashboards for navigations for ease of data reading by operational staff.

- Graphical representation of data within the platform and ability to export the same in PDF and JPEG format for a specified time period.

- Tabular representation of data within the platform and the ability to export in CSV or Excel format for a specified time period.

- Ability to set alerts and alarms for low and/or high; water levels, flow, rain, and temperature recordings.

- Ability for Waterways Ireland to amend device names as and when required by business need.

- Station list stating which are active/inactive and the last date and time of transmission.

Domain name

The domain name must include Waterways Ireland.

At present it is not a requirement that the platform is publicly available however this option should be available to Waterways Ireland should the strategic decision be made for the information to be shared publicly.

Maintaining the platform

        - Retrieving lost data

In the event of a communications failure between any monitoring station and the platform the Service Provider should be capable of retrieving any dropped data when the commutations link has been re-established. This data should be integrated with existing hydrometric records to provide an uninterrupted account of sensor information.

        - Platform interuptions

If a fault occurs with the platform the Service Provider must contact the Web Hosting Provider to alert them of the fault and act as an intermediary between Waterways Ireland and the web Hosting Provider until the fault is resolved. Waterways Ireland’s Project Manager should be immediately notified when a fault occurs and when the fault has been rectified and the system is functioning normally. Similarly, if routine maintenance causes an interruption to the platform the Service Provider should issue a notification to Waterways Ireland of proposed start and finish times with a minimum of 24hrs notice.

Security

The Service Provider will be custodians of the data, and should be responsible for verifying data integrity

Cyber essential at minimum scoring points and ISO 27001-2 higher points

Base MFA (Multi Factor Authentication) to access the service

Resilience and high availability MUST be without our intervention

Backups

Scheduled Backups to facilitate PITR (Point in Time Recovery)

Backup while operational / in use (Must be backed up while online)

Device Compatiblity

The web hosting platform must be compatible with the existing and any new monitoring stations installed throughout Waterways Ireland’s hydrometric network and must allow for the necessary transmission intervals in addition to holding the previous historic data for the hydrometric network.

Data Transfer

Waterways Ireland currently utilises Hydromet Cloud by OTT for collating and displaying the information transmitted from the various hydrometric devices installed across the inland navigations managed by Waterways Ireland.

The Service Provider will be required to ensure that historical information for a minimum of 3 years prior to entry into the Framework Agreement is transferred onto the new platform without loss of data or continuity of data.

KPI’s

The following KPI shall be reported on in advance of each progress meeting:

- Data shall be available through the web hosting platform at least 98% of the time with downtime time not being for more than 24hrs at any one time

Replacement device and equipment specification

general specficiation

The following requirements are common to all devices:

Design Life

The minimum design life of all instruments shall be 10-years (at the rated conditions).

Those items requiring regular replacement must be identified.

Device Longivity and Warranty

In relation to product longevity and warranty, the Service Provider must ensure that each device component is covered by the warranty terms for a minimum of two years from the later of delivery into stock or installation as part of providing the Services. The Service Provider must ensure that genuine or equivalent spare parts will be available for at least ten years from the date of purchase. The Service Provider must provide a copy of the warranty terms which includes the above requirements.

All devices and their components must be designed for repairability with repair being the preferred approach rather than replace were suitable.

Ingress protection

Ingress protection for instrument casings and enclosures must be:

- appropriate for the environment in which the instrument is to be installed;

- have due regard to operator safety.

The minimum ingress protection rating as defined in BS EN 60529 (IP Code) must be as listed in associated tables for the devices.

Materials

All equipment, cable terminations, electrical power supplies and associated apparatus that is likely to be exposed to weather, vandalism, explosive atmospheres, corrosive atmospheres or other adverse conditions must be constructed or protected as necessary to prevent damage or danger arising from such exposure.

- Plastics - Any plastics used in the supplied devices must be both PUR – Certified for drinking water purposes and be sufficiently UV resistant so that any exposed material to UV radiation doesn’t deteriorate and allow micro plastics to enter the water course.

- Metal – All metal used must be either Aluminium / POM or stainless steel to a rating of 1.4539 (904L) regardless of the distance they may be from saltwater to ensure that the devices have an acceptable life cycle.

- Ceramics – All Ceramics used in the pressure sensors must be suitably temperature compensated for purpose.

CE mark

All products supplied by the Service Provider must have and display the European CE certification mark.

Power source and Energy Efficiency

Thought should be given to the sustainability of the power sources used for each device. If possible, Waterways Ireland would like the Service Provider to consider the use of rechargeable batteries to reduce waste. Regardless of the type of battery used the Service Provider must ensure battery life is tested before replacing them to maximise their efficiency. These details must be provided in the checklist during servicing/replacing any devices.

In some situations, it may be suitable for solar panels to be considered and are the preferred power source over mains supply in all cases.

The devices must be able to transmit data via SMS and via data to the web-hosting platform. They must be compatible with a mobile SIM card or equivalent to facilitate data transmission.

End of life service

When a device has reached the end of its life it shall be the responsibility of the Service Provider to remove the product and take the required steps needed to ensure that the device follows the waste hierarchy as illustrated below. A record of each devices’ end of life must be submitted to Waterways Ireland at the progress meetings and as a minimum the Service Provider is required to take back the devices for functional testing to determine if it can be repaired or if it can be dismantled for re-use, recycling and /or disposal.

Water level recorder

Waterways Ireland require the use of two different water level devices hereafter named as Type 1 and Type 2. Overtime there will be the possibility to phase out Type 1 and should the service provider be considering a full replacement of devices Type 2 devices are to be installed in this scenario.

Level measurement shall comply with the requirements of BS EN ISO 4373.

The below table outlines the minimum requirements that are common between the two types of devices with any differences highlighted in the later sections.

| Water Level | Measuring range | 0…4m water column / 0 … 0.4 bar |  |
| --- | --- | --- | --- |
|  |  | 0 … 10m water column / 0 … 1 bar |  |
|  |  | 0 … 20 m water column / 0 … 2 bar |  |
|  |  | 0 … 40m water column / 0 … 4 bar |  |
|  |  | 0 … 100m water column / 0 … 10 bar |  |
|  | Resolution | 0.001 m / 0.1 cm / 0.01 ft / 0.1 inch / 0.0001 bar / 0.001 psi |  |
|  | Accuracy pressure | ± 0.05 % full scale |  |
|  | Long term stability (Linearity + Hysteresis) | ± 0.1 %/a full scale |  |
|  | Units | Water level m/cm/mbar |  |
|  | Pressure Sensor | ceramic / temperature-compensated |  |
|  | Temperature-compensated operating range | (compensated) -5 °C (ice-free) ... +45 °C |  |
| Measurement | Measurement values | Water level / water pressure |  |
|  |  | Temperature |  |
|  |  | Supply Voltage |  |
|  |  | RSSI / Signal strength |  |
|  |  | PBAT/ Power consumption battery |  |
| Data transmission | Transmission interval | 1 min ... 1/week |  |
|  | SMS | SMS data transmission / commands |  |
|  | Digital | Compatible with web hosting platform |  |
| Environmental | Temperature range, operating | -30 °C ... +85 °C | -22 °F ... +185 °F |
|  | Temperature range, storage | -40 °C ... +85 °C | -40 °F ... +185 °F |
|  | Humidity | 5% ... 95 % (non-condensing) |  |
|  | IP rating logger unit | IP67 |  |
|  | IP rating senor | IP68 |  |

Water level device Type 1

Type 1 shall be a reliable water level and temperature monitoring and data storage device with an external data logger with a 4MB memory. Current devices of this nature on the network have a data logger that is no longer in manufacture and as such should the data logger need replacing a new water level device type 2 shall be installed rather than a replacement type 1.

The specific minimum requirements for the type 1 device are:

| Data memory | Measurement memory | Approx. 500,000 values |  |
| --- | --- | --- | --- |
| Measurement | Sample/ Storage interval | 1s…24h |  |
| Power | Battery Life | > 5 years @ average temperature of 20 °C, 1 hour sampling and 1 transmission per day |  |
| Temperature | Measurement range | -25 °C ... +70 °C |
|  | Resolution temperature | 0.1°C |
|  | Accuracy temperature | ± 0.5 °C; optional ± 0.1 °C |
|  | Units | °C |

Water level device type 2

Type 2 shall be a self-contained ground water level logger for measuring water level and temperature, providing reliable and accurate data with integrated mobile communication. The device shall allow wireless local communication via Bluetooth Low Energy (BLE) to help facilitate servicing and data download in difficult to reach areas of the network. It shall be a self-contained unit with integrated data logger reducing the number of components on site that require maintenance and replacement. The integrated logger shall support mobile devices, smart phones and tablets operating with iOS, Android, or Windows 10.

Other minimum requirements include:

| Data memory | Measurement memory | 28MB (approx. 1,000,000 values) |  |
| --- | --- | --- | --- |
| Measurement | Sample/ Storage interval | 5s…24h |  |
| Power | Battery Life | > 10 years @ average temperature of 20 °C, 1 hour sampling and 1 transmission per day |  |
| RTC Clock | Accuracy | ± 26s/month (at 25°C) /< 3 s using SNTP |  |
| Interface | Cellular networks | LTE Cat-1 (4G)B3 (1800MHz), B8 (900MHz), B20 (800MHz) |  |
|  |  | EDGE, GSM, GPRS (2G): 900/1800 MHz |  |
|  | Local communication | BLE 5.0 – up to 10m (free line of sight) |  |
| Data Transmission | IP COM | FTP |  |
|  |  | HTTP |  |
|  |  | HTTPS (TLS1.2) |  |
| Temperature | Measurement range | -25 °C ... +70 °C |
|  | Resolution temperature | 0.01 °C |
|  | Accuracy temperature | ± 0.1 °C |
|  | Units | °C |

Fixed flow velocity and water level recorder

The fixed flow velocity and water level recorder is to be a stationary unit for continuous measuring in river, streams, canals and open channels. It must be an energy efficient system that provides reliable accurate data even at high water levels and high concentrations of suspended material. Although its primary function will be to collect flow velocity it is required that the device also has the capability of collecting water levels at the same location.

It is envisaged that such a device is likely to require an external data logger however may also include the use of an internal logger for some measurements.

The minimum requirements for the device are outlined in the below table:

| Flow Velocity | Measuring range | -10m/s … +10m/s |
| --- | --- | --- |
|  | Accuracy | 1% of meas. value ±5 mm/s |
|  | Resolution | 1mm/s |
|  | Measurement averaging time | 1s … 3600s |
| Water Level | Measuring range | 0.15 … 10m |
|  | Accuracy | ±3mm |
|  | Resolution | 1mm |
|  | Measurement Averaging Time | 1s … 3600s |
|  | Min. water Depth above instrument | 0.15m (water level option) |
| Internal Memory | Capacity | 9 MB (non-volatile) |
| Cable | Protection Class | IP68 |
| Discharge Calculation | Within the unit or externally on a data logger |  |

Weather station

Waterway Ireland current has 1 weather station that is used for national recording of weather information. As Waterways continues to adapt to the challenges created by climate change the gathering of this raw data is likely to be required and for this purpose the Service provider must be able to provide a weather station that meets the minimum requirements outlined below.

| General | Protection level housing | IP66 |
| --- | --- | --- |
|  | Standards/Regulations | Compliant to IEC 61724-1:2017 Class C |
| Lighting detection | Lightning detection | Number of lightning events |
|  | Coverage area | 5 – 10km |
| Relative humidity | Principle | Capacitive |
|  | Measuring range | 0 … 100 % RH |
|  | Unit | % RH |
|  | Accuracy | ±2% RH |
| Air pressure | Principle | MEMS capacitive |
|  | Measuring range | 300 … 1200 hPa |
|  | Unit | hPa |
|  | Accuracy | ±0.5hPa (0…40°C) |
| Wind direction | Principle | Ultrasonic |
|  | Measuring range | 0 … 359.9° |
|  | Unit | ° |
|  | Accuracy | < 3° RMSE > 1.0m/s |
|  | Resolution | 0.1 |
| Wind speed | Principle | Ultrasonic |
|  | Measuring range | 0 … 75m/s |
|  | Unit | m/s |
|  | Accuracy | ±0.3m/s or ±3% (0 … 35m/s) ±5% (>35m/s) RMS |
|  | Resolution | 0.1 |
| Precipitation (liquid) | Droplet size | 0,3 … 5mm |
|  | Detection sensitivity | 0,01mm/h |
|  | Particle velocity | 0.9 … 15.5m/s |
|  | Precipitation types | Rain/snow |
|  | Solid precipitation | 5.1… approx.. 30mm |
|  | Intensity range | 0…200mm/h |
|  | Intensity resolution | 0.01mm/h |
|  | Amount resolution | 0.1mm |
|  | Accuracy | 20% under laboratory conditions |
|  | Reproducibility | Typical >90% under laboratory conditions |
| Radiation | Unit | W/m2 |
|  | Accuracy | 5% |
|  | Response time (95%) | <1s |
|  | Spectral range | 300 to 1100nm |
|  | Measuring range | 1400 W/m2 |

Data logger

A number of devices require the use of external data loggers to record and transmit the data to the dedicated web hosting platform and it is the Service Providers responsibility to ensure that the devices proposed are compatible with a data logger capable of meeting the minimum specification below:

| Communications interfaces | Ethernet RJ-45 10 Mbit |
| --- | --- |
|  | USB Host and USB Device |
|  | RS-232 |
|  | Compatible with web hosting platform |
| Sensor interfaces | SDI-12 |
|  | RS-485 )SDI-12) |
|  | RS-232 (optional) |
|  | Analogue in/Analogue out (configurable) |
|  | Status input and output/impulse |
|  | Switched output |
| IP communication | Integrated TCP/IP stack (HTTP, FTP, SMTP, SNTP…) |
|  | GPRS, Ethernet/DSL, PPP over landline |
|  | Integrated web server |
|  | Encrypted data transmission SSL/TLS (HTTPS) |
| Time synchronisation | NTP (network time protocol) |
| Protection class | IP 41 |

Rittal Cabinets

To ensure data loggers are protected from the elements a stainless steel rittal cabinet has been previously installed to house them. Should this be a requirement of the new devices and data loggers the Service Provider must provide suitable cabinets that are watertight, have a minimum rating of NEMA 4x when locked and hold at least 1 key for every installed cabinet. When providing the cabinet, the Service Provider must ensure that all the telemetry inside the cabinet is correctly assembled as per the manufacture’s instructions.

Security Cage

A number of monitoring stations are located in urban areas and as such hold the risk of vandalism. To minimise the impact to the devices Waterways Ireland has historically installed security cages around the devices. There may be a need going forward to replace a number of these cages. The typical security cage has been 1m x 370mm x 400mm in size but will be subject to the size of the device installed by the service provider.

Scope of services for future Orders

Relocation or new device installation

As Waterways Ireland develops its Climate Action Plan and Strategic Water Management Plan a need may arise for the relocation of or new environmental/ hydrometric devices to be installed under the Framework.

Any relocation or installation of a new device will be under an Order with relevant specific details be provided at the time.

For reference most of Waterways Ireland’s current devices are installed by fixing a 50mm OD stainless steel pole vertical pipe to a reinforced concrete structure, natural stone structure or RSJ driven into the river bed. The pipe is normally secured with at least 3 number saddle clamps which are held in place by 100mm M12 through bolts. A 14mm drill bit is used to create the colt hole and injected with an epoxy rated for metal/concrete to hold the bolts in place. On the canal sections some of the devices are located in the horizontal position under the fixed waiting jetty upstream or downstream of the lock chambers.

If the Service Provider wishes to use a different installation method they will need to provide a standard detail highlighting what fixtures will be utilised under each Order. It is expected that the Service Provider will have experienced staff members capable of securely fixing any future devices at the agreed upon locations of any future devices.

Provision of training

From time to time there may be a requirement that Waterways Ireland will require training on existing or new devices and/ or the web hosting platform. This will primarily be to facilitate the changing of staff and to expand understanding of the systems within Waterways Ireland as and when required.

Sustainability improvements

As Waterways Ireland strives to be a more sustainable organisation there may be the possibility to develop more sustainable practices and power sources for the monitoring equipment throughout Waterways Ireland’s remit. Waterways Ireland may seek support in this over the duration of the Framework Agreement.

Maintenance, data collection, procurement and /or installation of smart controllers, smart pumps and associated water level and flow devices and ancillaries

As Waterways Ireland’s Water Management Strategy develops there may be a need to include the maintenance, data collection, procurement and/or installation of a number of other devices linked with the supply of water through the use of pumps. The current water supply and associated monitoring devices have been supplied by varying companies and not OTT. Waterways Ireland may incorporate the requirements for water supply equipment and their associated monitoring devices within this Framework Agreement.

Project specific monitoring, testing, data collection and analysis

Depending on the programme of works each year there may be the potential need for monitoring, testing, data collection and analysis of environmental, structural and/or hydrological features. This will be subject to the need under specific projects.

Spot Flow Measurement

Following the introduction of Abstraction and Impoundment regulations in ROI, there is now a need for Spot Flow Measurement. This typically involves measurements of flows in the feeder channels that enter to the canals and flow measurement in the canals itself using handheld electromagnetic flow meters or floating acoustic doppler current profilers.

It is also expected that periodic flow measurements will be undertaken at installed flow meters to verify their accuracy.

Topographical datum’s collection of the network

Previously surveys have been undertaken of each device on the network to ensure datums are correct. New devices will require datums to be surveyed during installation however a full check of all devices may be required should the Geoid models be updated again during the Framework Period.

Appendix A

Appendix B

GPP-Guidance-for-the-Irish-Public-Sector.pdf (epa.ie)