# Appendix 8 - Technical Compliance Matrix and Tender Response Document – SVUH- 26017 Provision of Mini C-Arm Fluoroscopy System for Plastic, Reconstructive & Hand Surgery

To be used in conjunction with Appendix 1: Requirements and Specifications to the RFT.

- TECHNICAL COMPLIANCE MATRIX

Appendix 1: Requirements and Specifications

The Mini C-Arm Fluoroscopy System for Plastic, Reconstructive & Hand Surgery must meet the following technical specifications. Mandatory requirements must be fully met; failure to comply may result in exclusion.

Tenderers must explain how their proposal meets Mandatory and Desired requirements in below 2. TENDER RESPONSE DOCUMENT.

- Technical Requirements

| Requirement | Status | Checklist |  |
| --- | --- | --- | --- |
|  |  | Yes | No |
| General |  |  |  |
| All equipment supplied must be user friendly, with clear identification of all controls, which are to be marked with standard IEC Symbols | Mandatory |  |  |
| The tendered equipment (and all associated options), including accessories and software, must be CE marked and compliant with the European Medical Device Regulation (EU MDR) 2017/745 and must be fully compliant with the requirements of all relevant European and Irish radiation safety standards/legislation including SI 256 of 2018 | Mandatory |  |  |
| Compliant with relevant IEC Standards | Mandatory |  |  |
| Provide an EU Declaration of Conformity, this should include: CE Certification and Unique Device Identifier (UDI) | Mandatory |  |  |
| Designated EU Notified body and associated 4-digit identification number | Mandatory |  |  |
| Certified risk classification | Mandatory |  |  |
| Confirmation of EUDAMED registration | Mandatory |  |  |
| If the product is not yet certified under the EU MDR 2017/745 but is currently certified under the previous MDD/AIMDD Directives, the supplier must provide: Confirmation of a formal MDR application, Supporting documentation confirming valid transitional arrangements | Mandatory |  |  |
| Detector |  |  |  |
| Large flat-panel detector (≥ 20 × 20 cm) | Mandatory |  |  |
| High-resolution detector (≤ 100 µm pixel spacing) | Mandatory |  |  |
| CMOS technology | Mandatory |  |  |
| X-ray Generator, X-ray Tube and AssemblyPulsed generator with adjustable rates | Mandatory |  |  |
| Generator ripple must be less than 5% at full load | Mandatory |  |  |
| Radiographic mode | Mandatory |  |  |
| Tube cooling (state type) | Mandatory |  |  |
| The half-value layer of the beam should be equivalent to at least 2.9 mmAl for X-ray tube voltages of 80 kVp (mm of Al – give details) | Mandatory |  |  |
| Dose reduction features |  |  |  |
| Low-dose modes. (State pulse per second offered) | Mandatory |  |  |
| Preview collimator | Mandatory |  |  |
| Optimised filtration options e.g. Cu filters. State specific options | Mandatory |  |  |
| Automatic Exposure Control (AEC) | Mandatory |  |  |
| Provide dose tables for normal and low dose rate modes | Mandatory |  |  |
| Navigation aids |  |  |  |
| Laser Targeting | Mandatory |  |  |
| Live Zoom | Mandatory |  |  |
| Digital Pen | Mandatory |  |  |
| Measurements | Mandatory |  |  |
| Patient dose recording |  |  |  |
| Dose Area Product (DAP) either measured or calculated but must be validated and traceable | Mandatory |  |  |
| DAP to be reported separately as DAP fluoroscopy, DAP Digital Acquisition and total DAP. State units (dGycm2, uGym2, cGycm2 etc) | Mandatory |  |  |
| Dose rate | Mandatory |  |  |
| Total exposure time | Mandatory |  |  |
| Total images acquired | Mandatory |  |  |
| RDSR report | Mandatory |  |  |
| Features |  |  |  |
| Touchscreen interface | Mandatory |  |  |
| Motorised collimator to reduce field on view | Mandatory |  |  |
| Surgical LED lights on the C-arm (reduce shadows from theatre lights) | Mandatory |  |  |
| PACs and RIS |  |  |  |
| All devices requiring connectivity must be IPv4 compliant | Mandatory |  |  |
| The system must include all DICOM classes for complete and seamless DICOM integration with PACS/RiS | Mandatory |  |  |
| Conform to DICOM standard v3 | Mandatory |  |  |
| If CD/DVD burn is an option that their viewing is supported on standard PCs | Mandatory |  |  |
| Specify the network connectivity speed | Mandatory |  |  |
| State size of network card supplied with the modality | Mandatory |  |  |
| Specify how the modality communicated with any optional post processing/radiologist workstation | Mandatory |  |  |
| General Requirements | Mandatory |  |  |
| The unit must be straight forward to learn, operate and maintain | Mandatory |  |  |
| The unit must be easy to clean, disinfect and/or sterilise as required | Mandatory |  |  |
| The unit must be constructed with durable materials to withstand typical abuse and cleaning | Mandatory |  |  |
| The controls should be clearly marked and their functions must be self-evident | Mandatory |  |  |
| Controls must be protected against accidental changes | Mandatory |  |  |
| Controls must be sealed against penetration of liquids | Mandatory |  |  |
| Labels and markings must be durable to withstand routine cleaning and normal wear | Mandatory |  |  |
| If the equipment contains critical components which have a limited life, the duration must be specified in the quotation, and an indication of the cost of replacement given | Mandatory |  |  |
| Integrated image storage capacity to facilitate continued use during times of lost PACs connectivity. State available memory | Mandatory |  |  |
| Power supply requirements (voltage, frequency, and maximum power consumption) | Mandatory |  |  |
| Electrical protection features, including over-voltage and surge protection | Mandatory |  |  |
| Boot-up times and battery operating times | Mandatory |  |  |
| Expected system lifespan | Mandatory |  |  |
| The supplier must install the system on site, including completion of an electrical safety test | Mandatory |  |  |
| All packaging associated with the delivery and installation must be removed from site by the supplier | Mandatory |  |  |
| Equipment dimensions, weight, and manoeuvrability | Mandatory |  |  |
| Confirm the expected lead time for delivery and installation | Mandatory |  |  |
| ≥160°+ orbital rotation / over-rotation | Desired |  |  |
| Compact & lightweight design | Desired |  |  |
| Digital magnification & real-time enhancement | Desired |  |  |
| Environmental operating conditions (temperature and humidity) | Desired |  |  |
| Heat output and ventilation requirements | Desired |  |  |

- Clinical Quality & Compliance

| Requirement | Status | Checklist |  |
| --- | --- | --- | --- |
|  |  | Yes | No |
| Excellent high-contrast image quality with no distortion | Mandatory |  |  |
| Real-time image processing & enhancement algorithms optimised for hand/extremity incorporating noise reduction, edge enhancement options | Mandatory |  |  |
| Facilitate manual contrast and brightness adjustment | Mandatory |  |  |
| Anatomical programs/imaging modes (including ortho/extremity specific) | Mandatory |  |  |
| Last Image Hold / Cine loop / Fluoro Save options | Mandatory |  |  |
| Standards & Quality Assurance |  |  |  |
| A detailed list of all standards which the tendered equipment complied must be clearly identified and included with the tender | Mandatory |  |  |
| Tenders must indicate the quality assurance equipment that is supplied with the tendered equipment to allow later in-house quality assurance | Mandatory |  |  |
| Must provide support for physics staff in the performance of acceptance testing, commissioning, and routine QA | Mandatory |  |  |
| Raw or unprocessed data in DICOM format available for image quality analysis | Mandatory |  |  |
| Access to QA system and image data for physics staff including necessary dongles, passwords and access codes. Detail options | Mandatory |  |  |
| Infection control – cleaning disinfection instructions | Mandatory |  |  |
| Raw or unprocessed images should conform with NEMA XR-2013 (R2018) | Desired |  |  |

- After-sales

| Requirement | Status | Checklist |  |
| --- | --- | --- | --- |
|  |  | Yes | No |
| Local distributor network & manufacturer backed service | Mandatory |  |  |
| Availability of spare parts for the specific model line | Mandatory |  |  |
| Provide full details of guaranteed support/parts availability in terms of years including End-of-Life notification timelines | Mandatory |  |  |
| State the warranty period and provide full warranty details | Mandatory |  |  |
| Escalation process | Mandatory |  |  |
| Guaranteed uptime / repair time | Mandatory |  |  |
| Details of the process for notifying the hospital of future FSN’s software updates, recalls or safety alerts | Mandatory |  |  |
| Confirmation that all FSN’s will be implemented at no additional cost during the warranty period and under any maintenance agreement | Mandatory |  |  |
| Details of how urgent safety updates will be managed to minimise clinical downtime | Mandatory |  |  |
| Availability of spare parts and commitment to supply parts for a minimum of ten (10) years from date of installation | Mandatory |  |  |
| Proven system uptime / reliability record | Desired |  |  |

- Support

| Requirement | Status | Checklist |  |
| --- | --- | --- | --- |
|  |  | Yes | No |
| On-call technical and clinical application support | Mandatory |  |  |
| Remote diagnostics / support capability/ remote technical support | Mandatory |  |  |
| Dedicated service engineers for this equipment line | Mandatory |  |  |
| Preventative maintenance and service contracts | Mandatory |  |  |
| Provide details of preventive maintenance programme, including recommended service intervals/requirements | Mandatory |  |  |
| Details of the supplier's service and support organisation, including the location of service engineers supporting Ireland | Mandatory |  |  |
| Outline post-warranty service contract options for the extended lifespan of the system | Mandatory |  |  |
| Provide details of annual service costs | Mandatory |  |  |
| Outline all call-out charges, labour charges, and any other related service costs | Mandatory |  |  |
| Confirm expected turnaround times for servicing, maintenance, repairs, and technical support | Mandatory |  |  |
| State whether a loan unit, replacement system can be made available in the event of device failure | Mandatory |  |  |
| Response to site for emergency and non-emergency service calls | Mandatory |  |  |
| Software support, including security patches, software updates, and upgrade policy | Mandatory |  |  |
| Detail the process for managing Field Safety Notices (FSNs), Field Safety Corrective Actions (FSCAs), recalls, and software safety updates, including expected response times | Mandatory |  |  |
| Details of equipment obsolescence management and end-of-life support | Mandatory |  |  |

- Training

| Requirement | Status | Checklist |  |
| --- | --- | --- | --- |
|  |  | Yes | No |
| Comprehensive on-site user and technical training to include clinical applications training for hand surgery workflows | Mandatory |  |  |
| Train-the-trainer programme | Desired |  |  |
| Ongoing refresher training options | Desired |  |  |
| Outline any additional training options available for clinical and technical staff | Desired |  |  |
| Frontline training | Desired |  |  |

- Extended Warranty

| Requirement | Status | Checklist |  |
| --- | --- | --- | --- |
|  |  | Yes | No |
| Extended warranty options (2-5 years). Annual Maintenance pricing for each year (10 years in total incl. contract extensions) of the contract and including the extended warranty | Desired |  |  |
| Minimum 24 – 36 month comprehensive warranty including parts labour, travel, software updates and Preventative Maintenance visits | Desired |  |  |

- Green Procurement

| Requirement | Status | Checklist |  |
| --- | --- | --- | --- |
|  |  | Yes | No |
| Energy-efficient pulsed generator & low power consumption | Mandatory |  |  |
| Design features supporting dose reduction (lower overall energy use) | Mandatory |  |  |
| Reduced use of lead / environmentally friendly materials | Mandatory |  |  |

- Radiation Safety

| Requirement | Status | Checklist |  |
| --- | --- | --- | --- |
|  |  | Yes | No |
| A visible power-on indication must be provided on the unit | Mandatory |  |  |
| Warning light and audible indicator for radiation exposure | Mandatory |  |  |
| The exposure must terminate if the exposure control switch is released prematurely | Mandatory |  |  |
| Radiation leakage from the tube housing must not be greater than 1 mGy/hr at 1 metre in any direction from the tube focus at the maximum rating of the tube assembly average over an area not exceeding 100 cm2 | Mandatory |  |  |
| The system must allow manual control of movements in the event of power failure. List all the equipment features that are designed to protect the patient in this situation | Mandatory |  |  |
| Patient dose data must be transferrable to RIS and PACs systems | Mandatory |  |  |
| System must have a secure method to disable unit (either key or password or alternative) Give details | Mandatory |  |  |
| Foot pedal must be supplied | Mandatory |  |  |
| Standby mode must be available and/or foot pedal shield so that the foot pedal cannot be activated inadvertently | Mandatory |  |  |

- TENDER RESPONSE DOCUMENT

Tenderers are required to formulate their proposals when they are responding to the Qualitative Award Criteria under 1-8 of Section 3.3 of the RFT.

SVUH will use the responses below to determine if a proposal meets the minimum service requirements.

The standard expected to be met is elaborated in Appendix 1 to this RFT. In conjunction with Appendix 1, Tenderers are required to outline details of how their proposal responds to the criteria below as in 3.3 of RFT.

As reiterated in Appendix 1: Tenderers must address each of the issues and requirements in Appendix 1 of the RFT and submit a detailed description in each case which demonstrates how these issues and requirements will be dealt with / met and their approach to the proposed delivery of the Services.  A mere affirmative statement by the Tenderer that it can/will do so, or a reiteration of the tender requirements is NOT sufficient in this regard.

The page limits mentioned below do not include the questions / details on the criteria.

QUALITATIVE AWARD CRITERIA (1 to 8 below)

| 1. | Technical Requirements |
| --- | --- |

Tenderers are required to outline in the two (2) A4 pages (excess of this number will NOT be assessed) how their proposal meets technical requirements.

1.  – Response:

| 2. | Clinical Quality & Compliance |
| --- | --- |

Tenderers are required to outline in the two (2) A4 pages (excess of this number will NOT be assessed) how their proposed offer meets clinical quality and regulatory compliance requirements.

2.  – Response:

| 3. | After-sales |
| --- | --- |

Tenderers are required to outline in the two (2) A4 pages (excess of this number will NOT be assessed) how their proposal addresses SVUH outlined After-sales requirement.

3.  – Response:

| 4. | Support |
| --- | --- |

Tenderers are required to outline in the two (2) A4 pages (excess of this number will NOT be assessed) how their proposal aligns with SVUH specific Support requirements.

4.  – Response:

| 5. | Training |
| --- | --- |

Tenderers are required to outline in the two (2) A4 pages (excess of this number will NOT be assessed) their proposal to SVUH Training requirements.

5.  – Response:

| 6. | Extended Warranty |
| --- | --- |

Tenderers are required to outline in the two (2) A4 pages (excess of this number will NOT be assessed) how their warranty package meets SVUH Extended Warranty requirements.

6.  – Response:

| 7. | Green Procurement |
| --- | --- |

Tenderers are required to outline in the two (2) A4 pages (excess of this number will NOT be assessed) how their offer comply with SVUH requirements listed under Green Procurement.

7.  – Response:

| 8. | Radiation Safety |
| --- | --- |

Tenderers are required to outline in the two (2) A4 pages (excess of this number will NOT be assessed) how their proposal aligns with the Radiation Safety requirements outlined.

8.  – Response:

End of Document