Tender Response Document

Tender Reference: LA3215C

Request for Tenders under Open (OJEU) Procedure for the Supply, Delivery, Installation and Commissioning of a Suite of Plasma Etcher & Plasma Deposition Tools, [4 Lots], for Tyndall National Institute, University College Cork (UCC)

Location:

Tyndall National Institute – Cork, Lee Maltings Complex, Dyke Parade, Cork, T12 R5CP, Ireland.

| Issue Date: | Tuesday 16th June 2026 |
| --- | --- |
| Contact Person for Queries: | Gráinne Quinn |
| Contact Details for Queries: | E-Tenders Mailbox |
| Closing Date for Receipt of Queries: | 12 noon on Tuesday 07th July 2026 |
| Deadline For Receipt of Tenders: | 12 noon on Friday 24th July 2026 |

Suppliers, please refer to the “Clarifications” document within the CFT documents for further information on  UCC Responses to Clarifications requested on the Tender Document specifications outlined in the Appendix 1 below.

Contents

1.	Instructions for the Completion of the Tender Response Document	4

2.	Introduction	5

3.	Summary of Requirements	7

4.	Section A: Selection Criteria – RFT Section 3.2	10

A1: Insurance	10

A2: Tax Clearance	12

A3: Statement of Turnover	13

A4: Audited Accounts & Reports	13

A5: Data Protection	14

A6: Quality Assurance & Accreditation	14

A7: Organisational Summary	15

A8: Previous Experience	16

A9: Health & Safety	22

A10: Appendix 3 Tenderers’ Statement	23

A11: Appendix 4 Declaration as to Personal Circumstances of Tenderer	25

5.	Sections B: Tender Response	28

B1: Requirements & Specifications	28

B2: Award Criteria for ALL Lots	30

B3: Appendix 1-Technical Specification	33

Lot 1 Plasma Etcher (Silicon /Dielectrics)	33

Lot 2 Plasma Etcher (Metals / Piezoelectrics)	85

Lot 3 Plasma Etcher (Slow & Controllable Rate / Atomic Layer Etch of Silicon / Dielectrics)	137

Lot 4 Plasma Deposition (PECVD)	189

B4: Appendix 2- Pricing Schedule	241

B5-Checklist for Tender Submissions	242

# Instructions for the Completion of the Tender Response Document

Tenders MUST submit the Tender Response Document (TRD) duly completed.

When completing this TRD, please note the following:

- Please ensure you complete all of Section A before you progress to Section B.

- Tenderers should note they MUST meet and ‘Pass’ all the requirements contained in Section A, the Selection Criteria, to be considered eligible for assessment against the award criteria, Section B.

- You need to COMPLETE AND RETURN TWO (2) DOCUMENTS.

This TRD and the accompanying Appendix 2 Pricing Schedule.

- You DO NOT need to return the RFT Document.

The RFT document is for information purposes only.

TRD - Section A Selection Criteria

- A1-A4: Economic & Financial Standing

- A5-A9 Technical & Professional Ability

- A10: Appendix 3 Tenderer’s Statement

- A11: Appendix 4 Declaration as to Personal Circumstances of Tenderer

TRD - Section B: Tender Response

- B1: Requirements & Specifications

- B2: Award Criteria

- B3: Appendix 1-Technical Specification

- B4: Appendix 2- Pricing Schedule (separate Excel File)

- B5: Checklist for Tender Submissions

The completed TRD must be returned in accordance with Section 2.6 of the RFT.

Suppliers please note:

eESPD - When submitting tenders on the eTenders platform, please ensure to complete the online eESPD as part of the tender submission. Instructions are included in the tender documents.

# Introduction

University College Cork (UCC)

UCC is an award-winning institution with a history of independent thinking stretching back over 170 years. UCC is proud to be ranked in the top 2% of universities in the world. Their beautiful university opened its gates to just 115 students in 1849. They now have a student population of over 21,000.

Tyndall National Institute

Tyndall National Institute is a leading European research centre in integrated ICT (Information and Communications Technology) hardware and systems. Specialising in both electronics and photonics – materials, devices, circuits and systems – it is a global leader in the core research areas of:

•	Smart sensors and systems

•	Optical communication systems

•	Mixed signal and analog circuit design

•	Microelectronic and photonic integration

•	Semiconductor wafer fabrication

•	Nano materials and device processing

Central to our mission is delivering economic impact through research excellence. We work with industry and academia to transform research into products in our core market areas of electronics, communications, energy, health, agri-food and the environment. With a network of over 200 industry partners and customers worldwide, we are focused on delivering real impact from our excellent research. Our ambitious five-year strategic plan sets out a clear strategy to create employment and build critical mass within the Irish technology space.

Specialty Products & Services Centre, Tyndall National Institute

The Specialty Products & Services (SP&S) Centre comprises front-end activities encompassing cleanroom fabrication and training facilities along with back-end Device Forensics activities including analysis, test reverse engineering, packaging and reliability evaluation. The centre develops and fabricates a wide variety of semiconductor devices, ranging from traditional silicon microelectronics and silicon MEMs to compound semiconductor-based photonic devices. The centre’s extensive back-end expertise and processing capabilities enable the resulting devices, to be packaged and assembled into formats that allow the customer to use them in the real-world application areas.

Tyndall now wishes to expand its state-of-the-art etch and deposition capabilities through procurement of a suite of tools in 4 Lots for Plasma Etcher (Silicon / Dielectrics) + Plasma Etcher (Metals / Piezoelectrics) + Plasma Etcher (Slow & Controllable Rate / Atomic Layer Etch of Silicon / Dielectrics) + Plasma Deposition (PECVD) which will support its silicon microelectronics, silicon photonics and silicon MEMS research.

Funding Acknowledgement

This item is funded by Taighde Éireann (Research Ireland).

# Summary of Requirements

Tenders are sought for the supply delivery, installation and commissioning of a suite of tools in 4 Lots for silicon-based and dielectric materials for University College Cork.

The new plasma tools will enhance the capabilities at the Tyndall National Institute for the fabrication of semiconductor devices, especially silicon microelectronics and silicon MEMS, but also other materials such as germanium and silicon carbide. We invite proposals for the following lots;

| Lot No. | Title |
| --- | --- |
| 1 | Plasma Etcher (Silicon /Dielectrics) |
| 2 | Plasma Etcher (Metals /Piezoelectrics |
| 3 | Plasma Etcher (Slow & Controllable Rate /Atomic Layer Etch for Silicon / Dielectrics) |
| 4 | Plasma Deposition (PECVD) |

Lot 1- Plasma Etcher (Silicon /Dielectrics): advanced plasma etching system capable of etching for polysilicon (poly-Si) / silicon dioxide (SiO2) / silicon nitride (SiN) thin films using medium / high-density plasma processes.

Lot 2- Plasma Etcher (Metals /Piezoelectrics): advanced plasma etching system capable of etching for aluminium (Al) & alloys / molybdenum (Mo) / aluminium scandium nitride (AlScN) thin films using medium / high-density plasma processes.

Lot 3- Plasma Etcher (Slow & Controllable Rate /Atomic Layer Etch for Silicon / Dielectrics): advanced plasma etching system capable of both conventional high-rate etching and slow & controllable rate / atomic layer etching for silicon (Si) / silicon dioxide (SiO2) / silicon nitride (SiN) thin films using medium / high-density plasma processes.

Lot 4- Plasma Deposition (PECVD): advanced plasma enhanced chemical vapour deposition system capable of depositing dielectric (SiO2, SiN) and amorphous (a-Si) thin films on semiconductor substrates.

This equipment will be essential for etching and depositing a range of materials used in the fabrication of silicon microelectronics, silicon photonics and silicon MEMS devices. The ideal tools will offer precision, reliability, and efficiency to meet the requirements for etching and deposition on both 100 mm and 200 mm diameter wafers. Tenderers with cutting-edge solutions that can achieve these high standards are encouraged to submit their proposals for this vital component in our technological advancement.

The equipment will be supplied for and maintained at the following site:

- Silicon Fabrication (Room A1.1.04),

- Tyndall National Institute,

- University College Cork,

- Lee Maltings, Dyke Parade,

- Cork T12 R5CP,

- Ireland.

Additional site information

Appendix A- Cleanroom Footprint

The suite of plasma etch, plasma / atomic-layer etch / plasma deposition (PECVD) tools will be installed in the Silicon-Fab cleanroom as direct replacements for existing tools which will be decommissioned and moved-out to allow installation of the new tools. The available cleanroom space will be limited so tools and ancillary equipment must fit within a space approx. 5 m2 per tool (the total footprint for the suite of tools being tendered-for is approx. 20 m2).

Digital files will be made available upon request.

Appendix B – Sub-Fab Footprint:

Vacuum pumps / chillers etc. will be located remotely, positioned in a Sub-Fab with up-to a 10 m vertical drop and up-to a 10 m horizontal run to the point-of-use at the tool and as direct replacements for the ancillary equipment for existing tools which will be decommissioned and moved-out to allow installation of the new ancillary equipment. The available Sub-Fab space will be limited and must fit within a space approx. 4 m2 per tool (the total footprint for the ancillary equipment for the suite of tools being tendered-for is approx. 16 m2).

Digital files will be made available upon request.

UCC will consider tenders offering new instrumentation, or instrumentation which has been used for demonstration/ purposes by the manufacture/distributor/supplier- or a refurbished instrument. The Goods/Instrument(s) must be functional as new and supplied with a comprehensive mechanical and electrical warranty.

If the Instrument is ex demo it MUST be no more than three (3) years old, and all parts and components MUST be available for at least the next ten (10) years. It MUST include a full refurbishment history with appropriate calibration certificates and full warranty.

Budget Available

The Contracting Authority estimates that the expenditure on the Goods to be covered by the proposed Goods Contract (excluding optional items) may amount to €3,300,000 ex VAT (all 4 lots)

This is to include capital costs and associated costs.

Please note: This is a total budget for all lots, however the contracting authority estimates that approximately €1.2M of this budget will be assigned to Lot 4 (Plasma Deposition (PECVD) and approximately €0.8M assigned to Lot 3. The remaining budget will be split between the other two lots. The Contracting Authority aims to optimise functionality between the total suite of equipment and will assign budget based on maximising optional extras at award stage.

Note: The Contracting Authority reserves the right to exclude submissions from tenderers that exceed the maximum budget available as stated above.

UCC reserves the right to buy additional options available with this equipment at a later stage should the budget become available. This option will be available for 24 months.

Please refer to the RFT document for details regarding additional drawdown requirements.

# Section A: Selection Criteria – RFT Section 3.2

Tenderers will either Pass OR Fail each of the Selection Criteria.

In the event of one or more of the Selection Criteria achieving a fail, the Tenderer will be excluded from participating in this Competition.

Economic and Financial Standing A1 – A4

## A1: Insurance

| Note – UCC reserves the right to determine whether the insurance evidence provided is sufficient to demonstrate compliance. Where certificates or broker letters do not clearly set out all required information – including limits, basis of cover, excesses, business description or indemnity to principal wording – the tender may be deemed non-compliant. UCC also reserves the right to request copies of relevant policy endorsements (e.g., indemnity to principal, additional insured wording) to verify compliance. |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Employer Liability |  | Public Liability |  | Product Liability |  |  | Professional Indemnity |  |
| Minimum Limit of Indemnity |  | €13,000,000 |  | €6,500,000 |  | €6,500,000 |  |  | €6,500,000 |  |
| I confirm that we have the following insurances in place: |  |  |  |  |  |  |  |  |  |  |
| Insurance Type |  |  | Level in Place |  | Details of Any Excess |  |  | Expiry Date |  |  |
| Employers Liability Required Level: €13 Million (on an each and every occurrence basis) - Euro or foreign currency equivalent |  |  | € |  |  |  |  |  |  |  |
| Public Liability Required Level: €6.5Million (on an each and every occurrence basis) - Euro or foreign currency equivalent |  |  | € |  |  |  |  |  |  |  |
| Product Liability Required Level: €6.5Million (on an aggregate basis) – Euro or foreign currency equivalent |  |  | € |  |  |  |  |  |  |  |
| Professional Indemnity Required Level: €6.5Million (each and every claim or in the annual aggregate) - Euro or foreign currency |  |  | € |  |  |  |  |  |  |  |
| AND |  |  |  |  |  |  |  |  |  |  |
| I confirm that if successful, where the levels required under the contract are higher than those currently in our possession, I will be in a position to put the required forms and levels of insurance required in place. |  |  |  |  |  |  | Yes/No |  |  |  |
| AND |  |  |  |  |  |  |  |  |  |  |
| I confirm that I will provide the following promptly on request: - evidence of insurance in place or - letter from Insurance Broker confirming that the required levels could be put in place if successful |  |  |  |  |  |  | Yes/No |  |  |  |
| Requirements for ALL insurance certificates: - Full legal name of insured entity and any subsidiaries involved in the contract. - Business description, confirming the policy covers the activities associated with the UCC contract. - Reference to the tender (e.g., “Tender Ref #– University College Cork”). - Policy details, including: - Policy type and number - Insurer name - Start and expiry dates - Limits of indemnity, specifying whether limits apply per occurrence or in the aggregate - Basis of cover (e.g., occurrence or claims‑made; retroactive date if claims‑made) - Excess / Deductible / SIR – explicit euro or foreign currency equivalent amounts must be stated - Territorial limits and jurisdiction - Confirmation of any contract‑specific requirements, including: Indemnity to Principal / Specific Indemnity to UCC |  |  |  |  |  |  |  |  |  |  |
| Signed: |  |  |  |  |  |  |  |  |  |  |
| Position: |  |  |  |  |  |  |  |  |  |  |

## A2: Tax Clearance

- Please provide a copy of a valid Irish Tax Clearance Certificate and complete the information below:

| Valid Irish Tax Clearance Certificate: |  |
| --- | --- |
| Evidence Attached | Yes/No |
| Appendix Number |  |
| Expiry Date: |  |

(ii) Alternatively, please provide the following information to enable online verification.

| Tax Registration Number |  |
| --- | --- |
| Tax Clearance Access number |  |

(iii) Or - Please provide evidence that a Tax Clearance Certificate has been applied for from the Irish Revenue Commissioners

| Evidence Attached | Yes/No |
| --- | --- |
| Appendix Number |  |

Note: All non-resident companies should apply for a general TCC where tax clearance is required in connection with a Public-Sector Contract (formal or otherwise).

Both Resident and Non-Resident Tax Clearance Certificate information may be obtained from the Irish Revenue Commissioners http://www.revenue.ie/en/tax/it/forms/webtc1.pdf

## A3: Statement of Turnover

Tenderers must provide evidence, from your company’s Auditor/Accountant, that their turnover, relating solely to the Goods specified in Appendix 1, has exceeded the amount per Lot below for the years 2022, 2023, and 2024:

(Note: Where suppliers have submitted for multiple lots, turnover must reflect the combined turnover amounts)

| Lot | Yearly Turnover | 2022 | 2023 | 2024 |
| --- | --- | --- | --- | --- |
| 1 | €1,300,000 |  |  |  |
| 2 | €1,300,000 |  |  |  |
| 3 | €1,600,000 |  |  |  |
| 4 | €2,400,000 |  |  |  |

Please confirm below that this information has been submitted with your tender response and specify the Appendix number of the information provided.

| Confirm Yes/No | Yes/No |
| --- | --- |
| Appendix Number |  |

## A4: Audited Accounts & Reports

Three (3) years' audited accounts, from your company’s Auditor/Accountant, must be provided to support the turnover statement. In line with current legislation, where a company has an annual turnover of less than €12 million, then audited accounts are not necessary. In these instances, a Statement of Accounts from the company’s accountants will suffice.

Please note: a link to audited accounts will not be accepted.

Please confirm below that this information has been submitted with your tender response and specify the Appendix number of the information provided.

| Audited Accounts, Yes/No | Yes/No |
| --- | --- |
| Statement of Accounts from Accountant (if turnover < €12M), | Yes/No |
| Appendix Number |  |

Technical and Professional Ability A5-A9

## A5: Data Protection

The Tenderer must ensure that the security of Personal Data at all times complies with the Data Protection Acts 1998 and 2003 and guidance issued by the Data Protection Commissioner of Ireland.

Tenderers must comply with the EU General Data Protection Regulation (GDPR), which replaces the Data Protection Directive 95/46/EC and came into effect on 25/05/2018.

Please provide a copy of your Data Protection Policy.

Please confirm below that this information has been submitted with your tender response and specify the Appendix number of the information provided.

| Confirm Yes/No | Yes/No |
| --- | --- |
| Appendix Number |  |

## A6: Quality Assurance & Accreditation

A certificate confirming an accreditation (e.g. accredited partnership) exists between the Instrument source manufacturer and the Tenderer for the proposed solution, confirming the Tenderer’s ability as authorised provider of the solution to the Contracting Authority.

Please confirm below that this information has been submitted with your tender response and specify the Appendix number of the information provided.

| Confirm Yes/No | Yes/No |
| --- | --- |
| Appendix Number |  |
| Expiry Date of Certification |  |
| Source Manufacturer | Yes/No |
| Is it 3rd-party certified or in-house? | 3rd Party Certified/In-house |

Regulatory Compliance – tenderers must demonstrate adherence to all relevant Irish and European regulatory requirements, including but not limited to the following.

- Must be clearly CE marked.

If supplying through the UK, please specify the location of the awarding body of the CE mark, as since Brexit, all UK suppliers must provide evidence that the CE mark has been certified by an EU Notified Body. (if applicable).

| Confirm CE Compliance: Yes/No | Yes/No |
| --- | --- |

## A7: Organisational Summary

Please complete the following table(s) with your Company Summary.  If the Tenderer is a group, separate information must be provided for each member of the group.

| Tendering Party Name |  |
| --- | --- |
| Legal Structure – partnership, limited company, etc. |  |
| Please confirm if you are an SME (Small and Medium Enterprise) as defined in Commission Recommendation 2003/361/EC | YES / NO Definition as per 2003/361/EC The category of micro, small, and medium-sized enterprises (SMEs) comprises enterprises that employ fewer than 250 persons and have an annual turnover not exceeding EUR 50 million, and/or an annual balance sheet total not exceeding EUR 43 million. |
| Tendering Group Information – if relevant |  |
| Contact Name: |  |
| Address: |  |
| Landline telephone number: |  |
| Mobile number: |  |
| Out of hours contact telephone number: |  |
| Email address: |  |
| Please provide a brief company profile, including an organisation chart (Please expand this box as necessary) |  |

P

Technical Resources – Tenderers must demonstrate that they have effective and efficient manpower, skills, and technical resources available to support the supply, installation, and delivery of the goods.

| Technical Resources | Please outline effective and efficient manpower, skills, and technical resources available to support the supply, installation, and delivery of the goods. Please include details of the names of staff and third parties. |
| --- | --- |

## A8: Previous Experience

Details of at least two (2) relevant reference sites, per lot, where similar goods have been supplied within the last five (5) years, similar in nature to the supplies, installation, and commissioning and validation required in this contract, that have been successfully completed. These MUST be included in the tables below.

Note: Only one (1) reference may be submitted from the Contracting Authority specified in the title of this Tender Response Document. Please ensure permission has been obtained from the client contact person before submitting the reference information.

Please confirm that this information has been completed and submitted with your tender.

| Confirm Yes/No | Yes/No |
| --- | --- |

Lot 1- Plasma Etcher (Silicon / Dielectrics)

| Previous Experience | Reference 1 |  |  |
| --- | --- | --- | --- |
| Title of Project: |  |  |  |
| Client Name & address: (Note: where name is confidential, please indicate nature/type of client and location) |  |  |  |
| Client contact person: |  | Phone no: |  |
| Contact E-mail Address: |  |  |  |
| Annual Value: |  |  |  |
| Any associated subcontractors: |  |  |  |
| Start Date: |  | End Date: |  |
| Please describe this contract in full |  |  |  |
| Insert, expanding this table as necessary. |  |  |  |
| Please outline the extent to which this contract is similar in nature and scale to the Contracting Authority’s current requirements.. Insert, expanding this table as necessary. |

| Previous Experience | Reference 2 |  |  |
| --- | --- | --- | --- |
| Title of Project: |  |  |  |
| Client Name & address: (Note: where name is confidential, please indicate nature/type of client and location) |  |  |  |
| Client contact person: |  | Phone no: |  |
| Contact E-mail Address: |  |  |  |
| Annual Value: |  |  |  |
| Any associated subcontractors: |  |  |  |
| Start Date: |  | End Date: |  |
| Please describe this contract in full |  |  |  |
| Insert, expanding this table as necessary. |  |  |  |
| Please outline the extent to which this contract is similar in nature and scale to the Contracting Authority’s current requirements.. Insert, expanding this table as necessary. |

Lot 2- Plasma Etcher (Metals / Piezoelectrics)

| Previous Experience | Reference 1 |  |  |
| --- | --- | --- | --- |
| Title of Project: |  |  |  |
| Client Name & address: (Note: where name is confidential, please indicate nature/type of client and location) |  |  |  |
| Client contact person: |  | Phone no: |  |
| Contact E-mail Address: |  |  |  |
| Annual Value: |  |  |  |
| Any associated subcontractors: |  |  |  |
| Start Date: |  | End Date: |  |
| Please describe this contract in full |  |  |  |
| Insert, expanding this table as necessary. |  |  |  |
| Please outline the extent to which this contract is similar in nature and scale to the Contracting Authority’s current requirements.. Insert, expanding this table as necessary. |

| Previous Experience | Reference 2 |  |  |
| --- | --- | --- | --- |
| Title of Project: |  |  |  |
| Client Name & address: (Note: where name is confidential, please indicate nature/type of client and location) |  |  |  |
| Client contact person: |  | Phone no: |  |
| Contact E-mail Address: |  |  |  |
| Annual Value: |  |  |  |
| Any associated subcontractors: |  |  |  |
| Start Date: |  | End Date: |  |
| Please describe this contract in full |  |  |  |
| Insert, expanding this table as necessary. |  |  |  |
| Please outline the extent to which this contract is similar in nature and scale to the Contracting Authority’s current requirements.. Insert, expanding this table as necessary. |

Lot 3- Plasma Etcher (Slow & Controllable Rate / Atomic Layer Etch of Silicon / Dielectrics)

| Previous Experience | Reference 1 |  |  |
| --- | --- | --- | --- |
| Title of Project: |  |  |  |
| Client Name & address: (Note: where name is confidential, please indicate nature/type of client and location) |  |  |  |
| Client contact person: |  | Phone no: |  |
| Contact E-mail Address: |  |  |  |
| Annual Value: |  |  |  |
| Any associated subcontractors: |  |  |  |
| Start Date: |  | End Date: |  |
| Please describe this contract in full |  |  |  |
| Insert, expanding this table as necessary. |  |  |  |
| Please outline the extent to which this contract is similar in nature and scale to the Contracting Authority’s current requirements.. Insert, expanding this table as necessary. |

| Previous Experience | Reference 2 |  |  |
| --- | --- | --- | --- |
| Title of Project: |  |  |  |
| Client Name & address: (Note: where name is confidential, please indicate nature/type of client and location) |  |  |  |
| Client contact person: |  | Phone no: |  |
| Contact E-mail Address: |  |  |  |
| Annual Value: |  |  |  |
| Any associated subcontractors: |  |  |  |
| Start Date: |  | End Date: |  |
| Please describe this contract in full |  |  |  |
| Insert, expanding this table as necessary. |  |  |  |
| Please outline the extent to which this contract is similar in nature and scale to the Contracting Authority’s current requirements.. Insert, expanding this table as necessary. |

Lot 4- Plasma Deposition (PECVD)

| Previous Experience | Reference 1 |  |  |
| --- | --- | --- | --- |
| Title of Project: |  |  |  |
| Client Name & address: (Note: where name is confidential, please indicate nature/type of client and location) |  |  |  |
| Client contact person: |  | Phone no: |  |
| Contact E-mail Address: |  |  |  |
| Annual Value: |  |  |  |
| Any associated subcontractors: |  |  |  |
| Start Date: |  | End Date: |  |
| Please describe this contract in full |  |  |  |
| Insert, expanding this table as necessary. |  |  |  |
| Please outline the extent to which this contract is similar in nature and scale to the Contracting Authority’s current requirements.. Insert, expanding this table as necessary. |

| Previous Experience | Reference 2 |  |  |
| --- | --- | --- | --- |
| Title of Project: |  |  |  |
| Client Name & address: (Note: where name is confidential, please indicate nature/type of client and location) |  |  |  |
| Client contact person: |  | Phone no: |  |
| Contact E-mail Address: |  |  |  |
| Annual Value: |  |  |  |
| Any associated subcontractors: |  |  |  |
| Start Date: |  | End Date: |  |
| Please describe this contract in full |  |  |  |
| Insert, expanding this table as necessary. |  |  |  |
| Please outline the extent to which this contract is similar in nature and scale to the Contracting Authority’s current requirements.. Insert, expanding this table as necessary. |

## A9: Health & Safety

Confirmation of the Organisation’s Health and Safety statement must be provided.

All Tenderers should demonstrate any other relevant standards related to supplying the tendered equipment.

Please confirm below that this information has been submitted with your tender response and specify the Appendix number of the information provided.

| Confirm Yes/No | Yes/No |
| --- | --- |
| Appendix Number |  |

## A10: Appendix 3 Tenderers’ Statement

[Tenderers shall complete and return the following form of Tenderers’ Statement printed on the Tenderers’ headed notepaper and signed by the Tenderer.]

TENDERERS’ STATEMENT

TO:  University College Cork (the “Contracting Authority”)

RE: LA3215C – Request for Tenders under Open (OJEU) Procedure for the Supply, Delivery, Installation and Commissioning of suite of Plasma Etcher & Plasma Deposition Tools, [4 Lots], for Tyndall National Institute, University College Cork (UCC).

Having examined your Request for Tenders (the “RFT”), including the Instructions to Tenderers, the Selection and Award Criteria, the Requirements and Specifications, and the Terms and Conditions of the Framework Agreement and the Goods Contract, we hereby agree and declare the following:

| 1. | We understand the nature and extent of the Goods required to be delivered as described in Requirements and Specifications at Appendix 1 to the RFT. |
| --- | --- |
| 2. | We accept all of the Terms and Conditions of the RFT, the Goods Contract, and the Confidentiality Agreement and agree, if awarded a Goods Contract, to execute the Goods Contract at Appendix 5 to the RFT and the Confidentiality Agreement at Appendix 6 to the RFT. |
| 3. | We accept all the Selection and Award Criteria as set out in Part 3 of the RFT. |
| 4. | We agree to supply the Contracting Authority with the Goods in accordance with the RFT and our Tender. |
| 5. | We agree that, if awarded any Goods Contract, we shall, in the performance of such contract, comply with all applicable obligations in the field of environmental, social and labour law. |
| 6. | We confirm that we have complied with all requirements as set out at Part 2 of the RFT. |
| 7. | We confirm that all prices quoted in our Tender will remain valid for the period of time, commencing from the Tender Deadline, specified at paragraph 2.10.3 of the RFT. |
| 8. | We shall, if awarded any Goods Contract under the RFT, have in place on the Effective Date of the Goods Contract all insurances (if any) as required by paragraph 2.21.1 of the RFT. |
| 9. | We confirm that all Data Subjects whose Personal Data is provided in our Tender have consented to the processing of such Personal Data by us, the Contracting Authority, the Evaluation Team and the supplier of the etenders.gov.ie website, for the purposes of our participation in this Competition or that we otherwise have a legal basis for providing such Personal Data to the Contracting Authority for the purposes of our participation in this Competition and that we will provide evidence of such consent and/or legal basis to the Contracting Authority upon request. |
| 10. | We do not come within the category of prohibited economic operators identified in Regulation (EU) No 833/2014 of 31 July 2014 (as amended by EU Regulation 2022/576 or any subsequent amendments to same). |
| 11. | The origin of goods connected to our Tender, if any, are not subject to the prohibitions set out in Regulation (EU) No 833/2014 (as amended by EU Regulation 2022/576 or any subsequent amendments to same). |
| 12. | The subcontractor(s) on whose capacity we rely as part of our Tender (where the value of that subcontract exceeds 10% of the value of the Services Contract) does not come within the category of prohibited economic operators identified in Regulation (EU) No 833/2014 of 31 July 2014 (as amended by EU Regulation 2022/576 or any subsequent amendments to same). |

| SIGNED (Authorised Signatory) | Company |
| --- | --- |
| Print name | Address |
| Date |  |

## A11: Appendix 4 Declaration as to Personal Circumstances of Tenderer

RE: LA3215C – Request for Tenders under Open (OJEU) Procedure for the Supply, Delivery, Installation and Commissioning of suite of Plasma Etcher & Plasma Deposition Tools, [4 Lots], for Tyndall National Institute, University College Cork (UCC).

NAME:   	[Click here and insert name]

ADDRESS: 	[Click here and insert address]

I, [Click here and insert name of Declarant], of [Click here and insert name of entity]  do solemnly and sincerely declare that:

- I am a [insert role of Declarant] of [Click here and insert name of entity] and am authorized by [Click here and insert name of entity]  to make this declaration which relates to a tender (“the Tender”) submitted by [Click here and insert name of entity]  in response to an RFT dated titled [insert description of competition] published by [insert name of contracting authority] (“the Contracting Authority”).

- Neither [Click here and insert name of entity]  nor any person who is a member of the administrative, management or supervisory body of [Click here and insert name of entity]  nor any person who has powers of representation, decision or control in [Click here and insert name of entity]  has:

- ever been the subject of a conviction for participation in a criminal organisation, as defined in Article 2 of Council Framework Decision 2008/841/JHA.

- ever been the subject of a conviction for corruption, as defined in Article 3 of the Convention on the fight against corruption involving officials of the European Communities or officials of Member States of the European Union and Article 2(1) of Council Framework Decision 2003/568/JHA as well as corruption as defined in the national law of the Contracting Authority or the law of the state in [Click here and insert name of entity] is established.

- ever been the subject of a conviction for fraud within the meaning of Article 1 of the Convention on the protection of the European Communities’ financial interests.

- ever been the subject of a conviction for terrorist offences or offences linked to terrorist activities, as defined in Articles 1 and 3 of Council Framework Decision 2002/475/JHA respectively, or for inciting or aiding or abetting or attempting to commit an offence, as referred to in Article 4 of that Framework Decision.

- ever been the subject of a conviction for money laundering or terrorist financing, as defined in Article 1 of Directive 2005/60/EC of the European Parliament and of the Council.

- ever been the subject of a conviction for child labour and other forms of trafficking in human beings as defined in Article 2 of Directive 2011/36/EU of the European Parliament and of the Council.

- [Click here and insert name of entity]:

- is not in breach and has not breached its obligations relating to the payment of taxes or social security contributions.

- has carried out the preparation of the Tender independently.

- [Click here and insert name of entity]:

- has, in the performance of all public contracts, complied with applicable obligations in the field of environmental social and labour law that apply at the place where the works are carried out or the services provided, that have been established by EU law, national law, collective agreements or by international, environmental, social and labour law listed in Schedule 7 of the European Union (Award of Public Authority Contracts) Regulations 2016 (Statutory Instrument 284 of 2016).

- is not bankrupt or the subject of insolvency or winding-up proceedings, its assets are not being administered by a liquidator or by the court, it is not in an arrangement with creditors, its business activities are not suspended nor is it in any analogous situation arising from a similar procedure under national laws and regulations.

- is not guilty of grave professional misconduct.

- has not entered into agreements with other economic operators aimed at distorting competition.

- is not aware of any conflict of interest due to its participation in the Competition;

- has not had any prior involvement in the preparation of the Competition;

- has not shown significant or persistent deficiencies in the performance of a substantive requirement under a prior public contract, a prior contract with a contracting entity, or a prior concession contract, which led to early termination of that prior contract, damages or other comparable sanctions.

- is not guilty of serious misrepresentation in supplying the information required for the verification of the absence of grounds for exclusion or the fulfilment of the Selection Criteria for this Competition and did not withhold such information and did not fail or is not able to submit supporting documents in respect of this Competition as required under Regulation 59 of the European Union (Award of Public Authority Contracts) Regulations 2016 (Statutory Instrument 284 of 2016) .

- has not undertaken to unduly influence the decision-making process of the Contracting Authority in respect of the Competition or obtain confidential information that may confer upon it undue advantages in respect of the Competition; or negligently provided misleading information that may have a material influence on decisions concerning exclusion, selection or award.

- [Click here and insert name of entity] does not come within the category of prohibited economic operators identified in Regulation (EU) No 833/2014 of 31 July 2014 (as amended by EU Regulation 2022/576 or any subsequent amendments to same);

- The origin of goods connected to the Tender, if any, are not subject to the prohibitions set out in Regulation (EU) No 833/2014 (as amended by EU Regulation 2022/576 or any subsequent amendments to same);

- Any subcontractor, supplier or other entity on whose capacity [Click here and insert name of entity]  relies as part of the Tender does not come within the category of prohibited economic operators identified in Regulation (EU) No 833/2014 of 31 July 2014 (as amended by EU Regulation 2022/576 or any subsequent amendments to same).

I understand and acknowledge that the provision of inaccurate or misleading information in this declaration may lead to my business/firm/company/partnership being excluded from participation in this or future tenders, and I make this solemn declaration conscientiously believing the same to be true and by virtue of the Statutory Declarations Act, 1938. This declaration is made for the benefit of the Contracting Authority.

________________________

Signature of Declarant

Declared before me by ___________________________________ who is personally known to me

(or who is identified to me by ______________________________who is personally known to me) or*

at ____________________________  this ___________ day of _______________ 20__

________________________________

(signed)

Practicing Solicitor/Commissioner for Oaths

*Please include such other form of identification used to identify the Declarant as permitted by the Statutory Declarations Act, 1938 (as amended)

# Sections B: Tender Response

## B1: Requirements & Specifications

Suppliers MUST supply all details regarding the equipment and associated consumables in the response boxes below. Responses will be scored by the Technical Evaluation team in accordance with the Award Criteria detailed in the RFT document. Please ensure all cost submissions are kept separate and included in the Excel file titled: Appendix 2 – Pricing Schedule.

This is the minimum specification required. All tenderers must meet or exceed this minimum specification to progress through the award criteria to cost evaluation.  Additional functionality should be identified in the comments section to facilitate scoring under the award criteria.

Supplier Notes

- This specification document is designed to assist the contracting authority in its evaluation of your tender and to enable suppliers to provide structured responses to key deliverables.

It is structured to reflect the award criteria as set out in the RFT.

- Please note responses to individual questions will NOT be scored separately, but collectively they will provide a basis for assessing the merits of the tenderer’s overall response to each award criterion and sub-criterion.

- Where specifications include reference to manufacturer or brand names, this is a result of difficulties in specifying the product required. In these instances, the phrase “or equivalent” is inferred.

- It is acknowledged that each Tender’s product line may differ from these specifications.

As such, each seller is free to propose variances from these specifications.

Alternative proposals favourable to the Contracting Authority may be suggested.

It is required, however, that whenever a variance from these specifications occurs, the proposed item must meet or exceed the specified characteristics or level of performance.

- Where abnormally low tenders are received, the Evaluation team shall require the tenderer to explain the price proposed in Appendix 2. (Per Article 69 of the EU Procurement Directives 2014/24).

Where the tenderer cannot provide a sufficient explanation, the contracting authority is entitled to reject the tender.

- The award of any contract is subject to the successful tenderer meeting the required specification and standards as set out in this RFT and Tender Response Document (TRD).

Should the required specification and standards not be met, the EPS reserves the right, in consultation with the Contracting Authority, to award any contract to the next highest-ranked Tenderer.

- Should the required specification and standards not be met, the EPS reserves the right, in consultation with the Contracting Authority, to award any contract to the next highest-ranked Tenderer.

- If the evaluation outcome results in a tie between two or more Tenders, then the Tender with the highest score for the highest weighted criterion shall be deemed the most economically advantageous tender.

If the score remains in a tie, the remaining criterion will be assessed by the highest weighted until there is no longer a tie.

If the score remains a tie after all criterions are assessed, the tenderers will be asked to submit a Best and Final Offer (BAFO).

A failure to answer question(s) may leave the contracting authority with insufficient information on which to base a thorough evaluation and, in turn, may result in loss of marks in respect of the relevant Award Criterion. Therefore, please ensure you complete ALL sections of this questionnaire.

Only those Tenderers who have qualified in accordance with paragraph 3.2 of this RFT will proceed to be evaluated under this paragraph 3.3.1

## B2: Award Criteria for ALL Lots

The Goods Contract will be awarded on the basis of the most economically advantageous tender(s) as identified in accordance with the following award criteria:

| LA3215C – Request for Tenders under Open (OJEU) Procedure for the Supply, Delivery, Installation and Commissioning of suite of Plasma Etchers & Plasma Deposition Tools, [4 Lots], for Tyndall National Institute, University College Cork (UCC) |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | Award Criteria: Please Note: All sub criteria for sections B, C, D & E carry equal weight/marks. For section A some sub criteria carry its own weighting. | Weighting | Total Maximum Score | Minimum Score (60%) Requirement |
| - Technical Merit |  |  |  |  |
| A | - General/ Functional Requirements - Performance Requirements - Hardware/Software Requirements | 30% 5% 15% 10% | 300 50 150 100 | 30 90 60 |
| - Quality, Health & Safety |  |  |  |  |
| B | - Quality Requirements - Health & Safety Requirements | 5% | 50 | 30 |
| - Warranty/Spare Parts, After-Sales and Training Support |  |  |  |  |
| C | - Warranty & Customer Support, Repair/Breakdown Requirements - Replacement & Spare Parts & Equipment Consumables - Training & Technical Support Proposal | 10% | 100 | 60 |
| - Delivery Plan |  |  |  |  |
| D | - Lead Time & Delivery - Installation, Commissioning, Calibration & Validation Requirements - Documentation | 10% | 100 | 60 |
| - Energy, Environmental & Sustainable Considerations |  |  |  |  |
| E | - Environmental & Sustainable Considerations | 5% | 50 | 30 |
| - Ultimate Cost |  |  |  |  |
| F | Ultimate Cost Including: (this includes the Total Purchase Cost/Capital Cost (ex. VAT) and the After Sales Service & Warranty Costs (Service Cost), exc. VAT as per Appendix 2 – Pricing Schedule | 40% | 400 | No Minimum Score Applies to Ultimate Cost |
|  | - Total Maximum Score Available | 100% | 1000 |  |

Tenderers must achieve a minimum of 60% on all tender award criteria above, excluding Ultimate Cost, in order to avoid elimination from the competition. Only tenderers who have achieved the 60% will be evaluated against the Ultimate Cost.

No commitment of any kind, contractual or otherwise, shall exist unless and until a formal written contract has been executed by or on behalf of the Contracting Authority. Any award of notification of preferred bidder status by the Contracting Authority shall not give rise to any enforceable rights by the Tenderer. The Contracting Authority may cancel this public procurement competition at any time prior to a formal written contract being executed by or on behalf of the Contracting Authority. The Contracting Authority does not bind itself to accept the lowest priced or any tender.

The award of the contract will be on the basis of the Most Economically Advantageous Tender received, in accordance with the detailed award criteria and weightings set out below.

The lowest Ultimate Cost tender will receive the maximum score achievable under this criterion.

| Pricing Score = | Lowest Tendered Price x Max Points |
| --- | --- |

Price being evaluated

Responses for non-cost/qualitative Award Criteria items will be scored using the following matrix:

| Scoring Methodology |  |
| --- | --- |
| Scoring Range | Meaning |
| 91% – 100% | Excellent response that fully meets and exceeds requirements, providing comprehensive, detailed, and convincing assurance that the Tenderer will deliver to an exceptional standard. |
| 80% – 90% | A very good response that demonstrates real understanding and fully meets the requirements and provides convincing assurance that the Tenderer will deliver to high standard. |
| 60% – 79% | A satisfactory response which demonstrates a reasonable understanding of requirements and provides reasonable confidence that the minimum requirements will be delivered. However, it does not provide sufficiently convincing assurance to award a higher mark. |
| 30% - 59% | A response where reservations exist due to a lack of convincing details and comprehensive credibility. There is a significant risk that the requirements will not be delivered to a satisfactory standard. |
| 1% – 29% | A response that raises significant reservations, due to insufficient detail, critical or notable deficiencies, that seriously lacks credibility, presenting a high risk of non-delivery. |
| 0% | No Response or partial Response that completely fails to address the criterion under consideration |

-Appen

## B3: Appendix 1-Technical Specification

| Lot No. | Title | Tendering (Y/N) |  |
| --- | --- | --- | --- |
| 1 | Plasma Etcher (Silicon /Dielectrics) |  |  |
| 2 | Plasma Etcher (Metals / Piezoelectrics |  |  |
| 3 | Plasma Etcher (Slow & Controllable Rate / Atomic Layer Etch for Silicon / Dielectrics) |  |  |
| 4 | Plasma Deposition (PECVD) |  |  |

dix 1-Technical Specification

Suppliers MUST provide all information relating to the Equipment and the associated consumables in the response boxes below. Responses will be scored accordingly by the Evaluation team as outlined in the Award Criteria in the RFT document.

Where an additional document is required to support your response, the document MUST be attached and the relevant section within the document MUST be clearly referenced in your answer.

Please note that website links are not sufficient as a supporting document.

Only referenced documents will form part of the evaluation. Brochures that have not been referenced will NOT be evaluated.

Please ensure all cost submissions are kept separate and included in the Excel file named: Appendix 2 – Pricing Schedule.

Pass/Fail Criteria

LA3215C – Request for Tenders under Open (OJEU) Procedure for the Supply, Delivery, Installation, Commissioning Supply, Delivery, Installation and Commissioning of suite of Plasma Etcher & Plasma Deposition Tools, [4 Lots], for Tyndall National Institute, University College Cork (UCC)

# Lot 1 Plasma Etcher (Silicon /Dielectrics)

| - Technical Merit |  |  |  |  |
| --- | --- | --- | --- | --- |
| General/Functional Requirements |  |  |  |  |
|  | The Tenderer MUST supply a single-wafer processing tool for etching of Silicon / Silicon-based and Dielectric ( e.g. silicon oxide aka SiO and silicon nitride aka SiN) materials. Please confirm and detail costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
|  | UCC will consider tenders offering new instrumentation or refurbished and factory calibrated instrumentation. Tenderers MUST state whether new or ex demo/used systems is proposed. If the Instrument is ex demo it MUST be no more than three (3) years old and all parts and components MUST be available for at least the next ten (10) years. It MUST include a full refurbishment history with appropriate calibration certificates and full warranty. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The tool MUST be capable of etching the following materials: - Silicon / Silicon-based materials (e.g. polysilicon, silicon-on-insulator aka SOI, amorphous silicon aka a-Si etc.) and - Dielectric materials (e.g. Silicon Dioxide aka SiO2 / Silicon Nitride aka SiN etc.) Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The tool MUST be capable of etching the materials listed in A3 above for a variety of micro and/or nanoscale structures with nanometre precision and excellent wafer-scale uniformity (Within-Wafer-Uniformity aka WiW) with the following parameters; - Photoresist-patterned structures with dimensions - min. ≥ 100 nm - Etch depths - min. ≥ 20 nm - max. ≤ 10 µm | Confirm (Yes / No) | Insert Comment |  |
|  | The tool MUST be suitable to run both flourine-based and chlorine-based chemistries. Any additional hardware (e.g. removable process-chamber liners etc.) required to allow for better process integrity and rapid switch-over between the different etch chemistries MUST be included as part of the core package. Please confirm and detail costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
|  | The Tenderer MUST provide plasma cleaning & conditioning recipes and/or best-known method (BKM) procedures for conditioning the chamber when switching between fluorine-based and chlorine-based chemistries as part of the core package. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The tool MUST include a vacuum loadlock chamber as part of the core package. [Tools with open-loading process-chambers will NOT be considered suitable for this tender.] Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The tool MUST be capable of etching silicon-based & dielectric materials on silicon and other substrates in wafer form (flatted or notched as-per SEMI M1-1117) for the following; - 100 mm diameter - 200 mm diameter Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The tool MUST be capable of handling the following substrate types on 100mm diameter wafers in flatted or notched form as-per SEMI M1-1117 a minimum requirement: - Silicon (Si) - Silicon-on-Insulator (SOI) - Glass / Fused Silica - Silicon Carbide (SiC) - Germanium (Ge) Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| A10 | The required etch performance of the tool MUST be achieved by inductively-coupled plasma (ICP) configuration with independent RF controls between the two sources. Bias control with high accuracy is considered essential for low-damage processes (e.g. polysilicon gate etch onto SiO2 etc.). Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The tool MUST be equipped with endpoint detection (EPD) hardware & software (e.g. Optical Emission Spectroscopy aka OES or laser interferometry etc.) as part of the core package. The proposed endpoint solution(s) MUST have the capability to monitor / control both etching and chamber cleaning & conditioning. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
| Performance Requirements |  |  |  |  |
|  | The tool MUST be capable of etching Silicon / Silicon-based materials & Dielectric materials with excellent wafer-scale uniformity. Within-Wafer-Uniformity (aka WiW) of the etch depth as measured at centre / edge points of wafer (with 10 mm edge-exclusion on both 100 mm and 200 mm diameter wafers) to be < ± 5%. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The tool MUST be capable of using the following mask materials for etching Silicon / Silicon-based materials & Dielectric materials as a minimum requirement: - Photoresist (e.g. optical and electron-beam), - Dielectric hard-masks (e.g. silicon-dioxide, silicon-nitride, amorphous-silicon) It is expected that photoresist will be used as the mask material for most applications. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | Recipe # 1 (Polysilicon gate etch) The tool MUST be capable of demonstration and a process recipe MUST be provided with the tool as part of the validation tests post at the Tyndall National Institute site for: Etching of ≤ 2 µm linewidth microstructures in polysilicon with soft-landing / minimal loss on silicon-dioxide (SiO2) under-layer on both 100 mm and 200 mm silicon wafers. The sidewall profile is to be highly anisotropic (≥ 85°). Tenderers MUST clearly demonstrate how they meet the requirements outlined with results documented in the attached Appendix C – Etch Recipe Results Template, this must be clearly named as “Lot 1 – Recipe # 1 -Results”. The Tenderer is permitted to use its own test mask / layout and substrates and / or own etching results to demonstrate compliance. Requirements: - Mask layout: open area > 80% (light-field pattern) - Feature size & type: ≤ 2 µm linewidth gate (isolated and nested) - Mask materials: Photoresist (PR) - Aspect ratio: > 1:1 - Depth: < 500 nm - Etched profile: anisotropic ≥ 85° and ≤ 90° or equivalent to the photoresist sidewall angle - Mask undercut: < 100 nm - Etched sidewall: smooth - Etch rate Si: > 100 nm/min - Etch rate SiO2: < 5 nm/min - Selectivity to Photoresist: > 2:1 (Si etch-rate ~ 2x photoresist etch-rate) - Non-uniformity depth (100 mm wafer): < 5 (± %) - Non-uniformity depth (200 mm wafer): < 5 (± %) Tenderer is to specify the recommended process temperature and gases for this recipe Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | Recipe # 2 (SiO2 contacts etch) The tool MUST be capable of demonstration and a process recipe MUST be provided with the tool as part of the validation tests post at the Tyndall National Institute site for: etching of ≤ 2 µm linewidth contact / via microstructures in silicon-dioxide (SiO2) with soft-landing / minimal loss on silicon / polysilicon under-layer on both 100 mm and 200 mm silicon wafers. The sidewall profile is to be tapered (≤ 70°) to improve gapfill for subsequent metal deposition. Tenderers MUST clearly demonstrate how they meet the requirements outlined with results documented in the attached Appendix C – Etch Recipe Results Template, this must be clearly named as “Lot 1 – Recipe # 2 -Results”. The Tenderer is permitted to use its own test mask / layout and substrates and / or own etching results to demonstrate compliance. Requirements: - Mask layout: open area < 10% (dark-field pattern) - Feature size & type: ≤ 2 µm linewidth contacts / vias - Mask materials: Photoresist (PR) - Aspect ratio: > 1:1 - Depth: ≤ 1 µm - Etched profile: anisotropic ≤ 70° or equivalent to the photoresist sidewall angle - Mask undercut: < 100 nm - Etched sidewall: smooth - Etch rate SiO2: ≥ 200 nm/min - Etch rate Si / polysilicon: < 10 nm/min - Selectivity to Photoresist: > 1:1 (SiO2 etch-rate ~ photoresist etch-rate) - Non-uniformity depth (100 mm wafer): < 5 (± %) - Non-uniformity depth (200 mm wafer): < 5 (± %) Tenderer is to specify the recommended process temperature and gases for this recipe. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | Recipe # 3 (SiN waveguide etch) The tool MUST be capable of demonstrating and a process recipe MUST be provided with the tool as part of the validation tests post at the Tyndall National Institute site for etching of: ≤ 2 µm linewidth microstructures in silicon nitride SiN) with good selectivity to silicon-dioxide (SiO2) hard-mask on 100 mm and 200 mm silicon wafers. The sidewall profile is to be highly anisotropic (≥ 85°). Tenderers MUST clearly demonstrate how they meet the requirements outlined with results documented in attached Appendix C – Etch Recipe Results Template, this must be clearly named as “Lot 1 – Recipe # 3 -Results”. The Tenderer is permitted to use its own test mask / layout and substrates and / or own etching results to demonstrate compliance. Requirements: - Mask layout: open area > 80% (light-field pattern) - Feature size & type: ≤ 2 µm linewidth waveguide - Mask materials: Photoresist (PR) - Aspect ratio: > 1:1 - Depth: ~500 nm - Etched profile: sloped > 85° and < 90° or equivalent to photoresist sidewall angle. - Mask undercut: < 100 nm - Etched sidewall: smooth - Etch rate SiN: > 200 nm/min - Etch rate SiO2: < 20 nm/min - Selectivity to Photoresist: > 1:1 (Si etch-rate ~ photoresist etch-rate) - Non-uniformity depth (100 mm wafer): < 5 (± %) - Non-uniformity depth (200 mm wafer): < 5 (± %) Tenderer is to specify the recommended process temperature and gases for this recipe. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The tool SHOULD have the capability to run the etch recipes #1, #2 & #3 listed above at room temperature (20°C) and this will be considered an advantage and scored higher. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The tool SHOULD be capable of etching the following materials with the same configuration as for the recipes #1, #2 & #3 listed above and this will be considered an advantage and scored higher. - Silicon on Insulator (SOI) - Aluminium Oxide (Al2O3) - Hafnium Oxide (HfO2) - Germanium (Ge) - Gallium Nitride (GaN) If any additional hardware (e.g. chamber-liners etc.) is required for etching these materials, the Tenderer MUST identify clearly in their response and cost separately. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |  |
|  | The Tenderer MUST provide plasma cleaning recipes for inter-wafer and inter-batch cleaning / polymer removal in the process-chamber as part of the core package. Please state if separate clean recipes for fluorine-based and chlorine-based chemistries are required and this MUST be clearly specified by the tenderer. Tenderer is to specify the recommended process temperature and gases for these cleaning recipes. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | It will be considered an advantage and scored higher if one or more of the gases required for process-chamber cleaning can also be used as an etch process gases, (e.g. CF4 / O2 / SF6 etc.), however if other gases are recommended for more efficient cleaning recipes then the tenderer SHOULD specify those gases as their preferred option. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | Recipe Support Tenderer MUST have ready access on request to a library of characterised etch / cleaning / conditioning recipes for the tool and assistance in the form of BKMs / recipes / technical support by phone & email when required in setting-up new recipes. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | Roadmap & Flexibility: Clarity of roadmap for the Tenderer’s proposed tool is required, indicating the potential future upgrade options that will be available as well as highlighting existing upgrade options for the system proposed. As with all pieces of equipment in UCC, the ability to upgrade in the future to add additional functionality not currently specified is very important, i.e., the tool SHOULD have a future roadmap of development. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The tenderer MUST state when the model was introduced and when the system is likely to be superseded. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| Hardware/Software Requirements |  |  |  |  |
|  | The system MUST be equipped with an aluminium process-chamber capable of being heated to > 50°C. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The process-chamber MUST be equipped with viewing port(s) for e.g. optical emission and laser reflection endpoint detection (EPD) systems at e.g. side and top chamber positions, and to directly view the plasma for stability. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The proposed endpoint solution(s) MUST be fully integrated with the system software. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | Tenderers SHOULD also supply further options of other EPD systems available. Please confirm and detail if any additional cost applies to additional EPD options that can be provided. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |  |
|  | Tenderers MUST identify clearly in their response if removable shields / liners for the process-chamber are required to allow for better process integrity / performance and to allow for rapid switch-over between fluorine-based and chlorine-based chemistries. Tenderers MUST include a minimum of two (2) sets of removable shields / liners as part of the core package if these are required for running the etch recipes #1, #2 & #3 listed above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
|  | The system MUST have a turbo pump to sustain low pressure (e.g. ≥ 1 mTorr) etch recipes / high flow-rate (e.g. > 100 sccm) cleaning recipes as part of the core package. Tenderers MUST specify the turbo pump brands and model numbers which will be provided with the tool as part of the core package. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | Tenders SHOULD also provide information on further options of other turbo pumps available for the process-chamber. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |  |
|  | The system MUST have a dry backing pump for the process-chamber. Tenderers MUST specify the pump brands and model numbers which will be provided with the tool as part of the core package. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The dry backing pump for the process-chamber MUST be located remotely, positioned in a sub-fab with up-to a 10m vertical drop and up-to a 10m horizontal run to the point-of-use at the tool. Please see attached Appendix B – Sub-Fab Footprint for further details on the sub-fab footprint. The dry backing pump MUST have sufficient performance when working in-conjunction with the turbo pump to sustain low pressure (e.g. ≥ 1 mTorr) etch recipes #1, #2 & #3 listed above and all necessary cleaning / conditioning recipes under these conditions. Sufficient length (minimum = 10 m) of electrical cable / plumbing etc. MUST be supplied to achieve this. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  |  |  |
|  | Tenders SHOULD also provide information on further options of other dry backing pumps available for the process-chamber. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |  |
|  | The process-chamber MUST recover to a pressure of ≤ 1 mTorr and a leak-up rate ≤ 2 mTorr/min after a wet-clean in less than two (2) hours. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The system MUST have a dry backing pump for the loadlock. Tenderers MUST specify the pump brands and model numbers which will be provided with the tool as part of the core package. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The dry backing pump for the loadlock MUST be located remotely, positioned in a sub-fab with up-to a 10m vertical drop and up-to a 10m horizontal run to the point-of-use at the tool. Please see attached Appendix B – Sub-Fab Footprint for further details on the sub-fab footprint. The dry backing pump MUST have sufficient performance to sustain low pressure (e.g. ≤ 100 mTorr) for wafer transfer into the process-chamber under these conditions. Sufficient length (minimum = 10 m) of electrical cable / plumbing etc. MUST be supplied to achieve this. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
|  | Tenders SHOULD also provide information on further options of other dry backing pumps available for the loadlock. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation | Confirm (Yes / No) | Insert Comment |  |
|  | Loadlock pump-down time SHOULD be less than 120 seconds, and no longer than 300 seconds. Faster pump-down time will be considered an advantage and scored higher. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The system MUST have a heated foreline to remove etch and cleaning plasma by-products from the process-chamber. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The system MUST have an automatic throttle valve for pressure control in the process-chamber. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The system MUST be water-cooled. Process cooling-water supply is provided by Tyndall, see section D. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The process-chamber platen / wafer-electrode MUST be heater-chiller controlled to allow operation in the range 0°C and +50°C with control of ±2°C from set point. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The heater-chiller unit MUST be located remotely, positioned in a sub-fab with up-to a 10 m vertical drop and up-to a 10 m horizontal run to the point-of-use at the tool. Please see attached Appendix B – Sub-Fab Footprint for further details on the sub-fab footprint. The heater-chiller unit MUST have sufficient performance to sustain platen / wafer-electrode temperature in the range 0°C and +50°C with control of ±2°C under these conditions. Sufficient length (minimum = 10 m) of electrical cable / plumbing etc. MUST be supplied to achieve this. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
|  | A non-return chiller valve MUST be included to prevent back-flow to reservoir. Tenderer MUST specify what model heater/chiller will be included and the coolant type (e.g. Galden or equivalent) recommended for the heater-chiller units. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | System MUST be equipped with a fully auto-matched RF generator for each plasma source. The RF generators MUST have capability for plasma strike and sustain tuned plasmas operating over a wide processing pressure range ≥ 1 mTorr. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | Tenderers MUST specify details of the RF generator brand & model, frequency, power and operating range for each plasma source which will be provided as part of the core package. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The RF generators / racks may be located remotely, positioned in a sub-fab with up-to a 10 m vertical drop and up-to a 10 m horizontal run to the point-of-use at the tool. Please see attached Appendix B – Sub-Fab Footprint for further details on the sub-fab footprint. Sufficient length (minimum = 10 m) of electrical cable / plumbing etc. MUST be supplied to achieve this. The RF generators / racks MUST have sufficient performance for plasma strike and sustain tuned plasmas under these conditions. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
|  | RF Generator racks MUST be supplied by the Tenderer, if required for the assembly of the system. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
|  | Tenderers SHOULD also provide information on further options of other RF generators available. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. If any options to purchase are available, these are optional to purchase by the Contracting Authority and will not form part of the Cost evaluation. | Confirm (Yes / No) | Insert Comment |  |
|  | The RF generators MUST have the capability to automatically attenuate the output power as required by the process recipe (e.g. low bias RF power for high-selectivity etching, high ICP RF power for chamber-cleaning etc.). Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The tool MUST be configured for single-wafer processing (wafers are etched individually). The hardware changes / procedures required to switch from running one wafer-diameter to another MUST be outlined clearly by the tenderer with all necessary hardware provided as part of the core package. The ability to switch from e.g. 100 mm to 200 mm substrate diameter with minimal downtime and / or hardware changes will be considered an advantage and scored higher. Please confirm and detail changeover procedure and time to complete. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
|  | The tool MUST have capability to handle wafer thicknesses in the range ≥ 0.2 mm and ≤ 1 mm as a minimum requirement. An ability to also handle substrates > 1 mm thickness with the same setup will be considered an advantage and scored higher. Please specify minimum and maximum thickness in tender response and identify any hardware changes (e.g. separate wafer-carriers for different thickness) required to allow a range of thin / thick substrates to be processed. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
|  | The tool SHOULD have capability to handle bonded-wafer pairs, perforated wafers, or bowed wafers (with bow / warp > 50 µm). This will be considered an advantage and scored higher. Please identify any hardware changes (e.g. separate wafer-carriers etc.) required to allow for handling these non-standard types of substrates. If any additional hardware (e.g. wafer-carriers etc.) is required for handling these non-standard types of substrates, the Tenderer MUST identify clearly in their response and cost separately. Tenderers are to provide details and pricing for the proposed option within the core package in the Appendix 2, Pricing Schedule. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation | Confirm (Yes / No) | Insert Comment |  |
|  | The tool MUST accommodate wafers as set out in requirements A8 – A9 above. Helium gas heat-transfer for wafer back-side cooling with electrostatic or mechanical clamping MUST be included. Helium transfer MUST be pressure-controlled, and the helium flow rate SHOULD automatically be checked prior to wafer etch. The He pressure MUST be stable at e.g. 10 Torr / 15 Torr during the clamping of a wafer and He flow leakage MUST not affect the process results. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The system MUST have a minimum eight (8) gas lines. The capability to have more than eight gas lines will be considered an advantage and scored higher. Suggested process gases include, but are not limited to the following; SF6, C4F8, O2, BCl3, Ar, Cl2 The tenderer MUST specify their recommended gases for the recipes #1, #2 & #3 listed above and for plasma cleaning & conditioning of the process-chamber. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | In the specific case(s) of low vapour pressure / low boiling point gases (e.g. BCl3) which are susceptible to condensation, heated gas-delivery hardware such as gas-lines, gas-box, MFCs etc. MUST be included as part of the core package. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
|  | In the specific case of oxygen (O2) the requirements for e.g. low flow-rate for high-selectivity etching of silicon onto silicon dioxide and high flow-rate for chamber plasma cleaning / chamber plasma conditioning may necessitate two separate MFCs, so this should be clearly specified by the Tenderer. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The tenderer MUST specify recommended flow size for each gas Mass Flow Controller (MFC). MFCs supplied MUST be rated (in terms of e.g. seals etc.) and calibrated for the recommended gases as nominated by the Tenderers under requirement A55 above. Digital MFCs are preferred and will be considered an advantage and scored higher. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The capability to run automatic interleaved and / or wafer-less clean MUST be available in the software. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | Whole system MUST be fully interlocked to protect against any equipment or service failure. System MUST be designed to fail-safe with system left under vacuum. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | Electrical equipment racks MUST be supplied by the Tenderer, if required for the assembly of the system. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | Electrical equipment racks may need to be located remotely, positioned in a sub-fab with up-to a 10 m vertical drop and up-to a 10 m horizontal run to the point-of-use at the tool and MUST be capable of full functionality under these conditions. Sufficient length (minimum = 10 m) of electrical cable MUST be supplied to allow the required performance. Please see attached Appendix B – Sub-Fab Footprint for further details on the sub-fab footprint. Confirmation is required that the operation of the tool and etch quality will not be impacted by remote location of the electronic control cabinets. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
|  | A minimum of one (1) computer and console MUST be included in core package, and MUST use Windows-based OS, Windows 10 or higher. Please describe the software package proposed, including name and version. Please also specify the projected lifespan of the current control system architecture. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
|  | The tool will be installed in the cleanroom in a thru-wall configuration. ALL computer(s) / console(s) required for tool operation from the front (process bay) and rear (service chase) side MUST be included in the core package. Additional console(s) / computer(s) for tool operation from the rear (service chase) side MUST also be included in the core package, or a suitable alternative such as a rotating screen(s) that are visible from both bay and chase with associated keyboard / mouse control from both sides. Please see attached Appendix A -Cleanroom Footprint for further details on the cleanroom footprint. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
|  | The tenderer MUST provide all control and analysis software required for tool operations. All tool operations, including wafer transfer to / from the loadlock to the process-chamber MUST be software-controlled and allow for both auto & manual operation. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
|  | The Plasma Etcher equipment SHOULD have perpetual licence-free software or no restriction on the number of licences needed for our use of the system i.e., licences for multiple users and computer installations for off-site use if this is possible. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | All software updates to the system, which become available during the life of the contract, MUST be provided free of charge during the warranty period and for a minimum period of eight (8) years after warranty expiry. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
|  | Continued development of the software SHOULD be demonstrated by the Tenderer, with facilities for updating systems in use. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The system MUST be capable of remote access via network connection (using e.g. TeamViewer software or equivalent) for tool support in event of problems. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The software MUST be capable of data-logging of process parameters, and archiving data-logs. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | It MUST be easy to graphically display, and export saved etch process parameter data preferably into Excel spreadsheets. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The system MUST have an easy-to-use interface and needs to be suitable for a multi-user (i.e. different access levels) environment. Various levels of access are required e.g. Operator, Supervisor, Maintenance. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The software SHOULD have the facility to lock certain recipes i.e. not editable by users other than Process Engineer / Supervisor / Maintenance etc. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | Software MUST include backup and restore capability (export/import) for process recipes and machine settings. Software SHOULD include backup (image) capability to be used in case of hard disk drive (HDD) failure. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | Software MUST allow operator access to data files while the system is running under standard operational conditions. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| - Quality, Health & Safety & Validation |  |  |  |  |
| Quality Requirements |  |  |  |  |
|  | Compliance with all relevant health and safety standards, quality assurance and analytical protocols is MANDATORY. All Goods and Services supplied in accordance with the contract MUST comply with all necessary standards, i.e. Irish Standards Institution, European Union, CE, International Standards Organisation or other applicable standard. The Tenderer MUST, on the request of the buyer, provide a copy of the standard free of charge. Please confirm & detail. | Confirm (Yes / No) | Insert Comment |  |
|  | If there are manufacturing defects known to the tenderer, they MUST identify such defects and state their policy regarding the repair of known defects. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| Health & Safety Requirements |  |  |  |  |
|  | All goods supplied in accordance with the contract MUST be of good construction, sound material, of adequate strength, free of design defects and manufactured to be safe and without risk when properly used and software MUST allow operator access to data files while the system is running under standard operational conditions. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The tool SHOULD meet the criteria of the Semi S2-0821Eb Environmental, Health, and Safety Guideline for Semiconductor Manufacturing Equipment in order to ensure easier and safer installation and maintenance from a gas safety perspective. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The Tenderer MUST, as part of the submission, include Laser Safety Warning labels and copies of any relevant testing or compliance certificates for the Goods. Confirm compliance and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The tools MUST knowingly not present any Health and Safety hazards. Any safety & hazard warning labelling on the proposed system MUST be in the English Language. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | Tenderers MUST advise if there are any hazards associated with the operation or maintenance of this instrument? (e.g. manual handing, chemicals, etc.). Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| - Warranty/Spare Parts, After-Sales & Training Support |  |  |  |  |
| Warranty & Customer Support, Repair/Breakdown Requirements |  |  |  |  |
|  | Warranty Duration: At least one (1) year’s warranty from Site Acceptance Test (SAT) MUST be included for each lot as part of the core package and the quoted price is to include all hardware, software, spare parts, labour, call outs, visit, travel and parts. Warranties of greater than 1 year are considered highly desirable and will be scored higher. Detail specific inclusions and exclusions of warranty. Preventative Maintenance MUST be provided for at least one (1) year and a list of checks carried out at Preventative Maintenance MUST be provided, along with any information about which (if any) parts are routinely replaced. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
|  | If UCC decide to extend the warranty period, Tenderers are asked to outline costs of any additional warranty. The Tenderer MUST state the details of any extended warranty and details of Preventative Maintenance and calibration & servicing check and any associated costs. Detail specific inclusions and exclusions of extended warranty and a list of checks carried out at Preventative Maintenance SHOULD be provided, along with any information about which (if any) parts are routinely replaced. Please include price per annum, details of coverage and maximum term of coverage. Please include price per annum for year 2, 3, & 4, details of coverage and maximum term of coverage. Please confirm and detail and specify what is included in the warranty options outlined in the Appendix 2 pricing schedule, Optional to Purchase. Detail specific inclusions and exclusions of extended warranty. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |  |
|  | Please provide separately a list and approximate cost of any consumable items/parts that are to be replaced during each service visit. Please update Appendix 2 Pricing Schedule – Consumables Parts sheet with detail. Prices provided for consumable items/parts MUST be valid for a minimum of one (1) year. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |  |
|  | Please advise of any items/part of the instrument that require routine replacement during the normal lifetime of the instrument, the timeframe in which they require replacing and if such items will be supplied with the instrument or as part of the warranty. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The tenderer MUST provide a description of the arrangements the arrangements and lead times for guaranteed responses times in the event of equipment breakdown both in relation to phone support, online/remote support, the provision of technical support, annual service/ maintenance, emergency call-out response, and spare parts. Please detail lead time for arrival on site by a fully trained service engineer during the warranty period(s) and post-warranty. This response time SHOULD be no more than 24 hours via phone call or remote service, and within seven (7) working days for onsite repair. Please confirm and detail information on remote diagnostic / telecom-based service support. | v | Insert Comment |  |
|  | Outside of Warranty Period: The Tenderer MUST provide typical prices and lead time for servicing and repairs (including travel, subsistence and price per day) charged outside a warranty/ maintenance contract. Detail specific inclusions and exclusions outside of warranty period. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |  |
|  | If UCC decide to take out an additional Preventative Maintenance Agreement / SLA, tenderers are asked to outline Options for years 2, 3 & 4. Option to include: - On-going Calibration to appropriate schedule - preventative maintenance - remote diagnostics of faults, if applicable - repair of the Instrument in case of breakdown) Tenderers are asked to outline full details of what is included and to outline Cost of maintenance and calibration for each year option offered. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |  |
|  | The tenderer MUST have dedicated fully trained and qualified service engineering specialists at their disposal who can respond to all service calls and any form of breakdown call. This remote response time SHOULD be the same day via phone call or remote service, and within seven (7) working days for onsite repair. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | Please outline the company’s Customer Support Proposal and escalation policy. Tenderers MUST describe what process is in place for escalating a call-out, if a breakdown cannot be resolved on the first visit. Tenderer SHOULD provide a list of approved service and calibration agents. Please confirm and detail customer support arrangements, service and calibration agents. Additional support and service options should also be outlined and pricing noted within the Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |  |
|  | The Tenderer MUST indicate the arrangements for and specify the level of experience of applications specialists and engineers who will be in charge of after sales service. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The Tenderer SHOULD specify the lead time for repairs during the following scenarios. - warranty period(s) - post-warranty, - with service contract - without service contract Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | Please note SHOULD installation, commissioning, service visits or availability of spare parts be affected in the event of a Pandemic crisis it is expected that the warranty period will be extended to facilitate the service downtime endured by the Contracting Authority. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | Warranty commences post full installation and sign-off of tool acceptance by nominated UCC representative. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| Replacement & Spare Parts & Equipment Consumables |  |  |  |  |
|  | Spare Parts: Availability of all spare parts SHOULD be maintained 10 years after the warranty period. Please indicate the availability and lead time of spare part stock centres, with maximum lead-time for any part. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | Critical spares MUST be held at a location no more than 7 Days delivery time from Tyndall. Spare parts to be held offsite with the tenderer and essential parts if necessary to be held Free of Charge (FOC) at Tyndall. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | Consumable Parts: Sufficient consumable parts MUST be supplied for the initial installation, commissioning and training. These consumable parts MUST be provided free of charge and at the point of commissioning. Full one (1) year of consumable parts to be supplied post commissioning, including any removable shields / liners for process-chamber. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
|  | Recommended Spare Lists: The tenderer MUST supply a full list of recommended spare parts and consumables for the proposed equipment, based on knowledge of equipment design and component lifetimes. This SHOULD include typical consumables and high turnover parts. Please detail in Appendix 2, Pricing Schedule- Consumable Parts. Prices provided for consumable items/parts MUST be valid for a minimum of one (1) year. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |  |
|  | The tenderer SHOULD supply, where information is available, to the buyer a schedule of spare parts showing in tabular form or, otherwise, the following information on OEM parts: - Names and address of all OEM suppliers - Part Name - Part Description - OEM Supplier Part Number. Please confirm and detail, attaching information requested above. | Confirm (Yes / No) | Insert Comment |  |
|  | The tenderer MUST provide confirmation of the availability of service for at least five (5) years after commissioning. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| Training & Technical Support Proposal |  |  |  |  |
|  | The successful tenderer MUST provide an initial training programme for UCC personnel on-site at Tyndall, for a minimum of one (1) day each for operations and maintenance, with a minimum of three (3) users fully trained to safely operate and maintain the system as part of the core package. The training programme MUST be carried out by a suitably qualified specialist at the tool after the initial installation, commissioning is completed. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
|  | Initial training provided MUST include, but not be limited to: - Demonstration of the correct operation of all relevant equipment of the systems - Routine Maintenance requirements - Basic Troubleshooting - Data processing and management - Software and hardware familiarisation Dates, timing and method of training are to be confirmed between UCC and the winning tenderer. The process validation tests may also be used for on-site training purposes. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | The tenderer MUST specifically outline the level and duration of application support offered with the instrument where “application support” refers to the expert technical support (on-line or on-site) available from the manufacturer for using the instrument in novel research applications. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  | This equipment will have multiple users over the lifetime of the tool. Tenderers MUST provide details and pricing for their proposed option for Operations / Applications / Maintenance Training at all levels (e.g. Level1, Level2, Level3 etc.) on-site or at their facilities. Training offered over the full lifetime of the equipment will be considered an advantage and scored higher. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |  |
|  | Relevant training and maintenance manuals in English MUST also be provided with the instrument. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| - Delivery Plan |  |  |  |  |
| Lead Time & Delivery |  |  |  |  |
| D1 | The equipment SHOULD be delivered within twelve (12) months of receipt of Purchase Order. The equipment SHOULD be delivered within the stated lead time requirement from receipt of Purchase Order. (Successful tenderer for each lot may be required to liaise with each other for coordination of delivery and installation of the equipment) - Tenderers MUST provide full details of lead times for the delivery of equipment including installation and commissioning (if applicable), from receipt of the purchase order. - Tenderers MUST provide the facility requirements to UCC a minimum of 16 weeks in advance of the equipment delivery. - Tenderers MUST provide advance notice (two weeks) of any requirements for tool unloading, the number of containers and weight of each. Height, length and width of containers/crates to be provided as well. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| D2 | Delivery MUST include all shipping charges, insurance, customs clearance and all aspects as outlined in DAP (Delivery At Place) incoterms. Delivery MUST be to Tyndall National Institute, UCC: - Silicon Fabrication Laboratory (Room A1.1.04), - Tyndall National Institute, - University College Cork, - Lee Maltings, Dyke Parade, - Cork T12 R5CP, - Ireland. Advance notice MUST be given (2 weeks) of any requirements for tool unloading, number of containers and weight of each. Height, length and width of containers/crates MUST be provided as well. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
| D3 | The tenderer MUST provide full details of lead times for the delivery of equipment including installation and commissioning (if applicable), from receipt of the purchase order. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| D4 | The tenderer MUST provide details of the logistics/supply chain to the point of customer delivery, i.e. where the goods will come from, how the goods will be transported and what borders the goods will cross. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| D5 | The tenderer MUST provide details of any actions that have been taken to ensure the future timely supply and delivery of goods and services, including ongoing supply of replacement/spare parts and consumables. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| D6 | The tenderer MUST provide details of their contingency plans (alternative sources/routes of supply) for any delays in delivery should they arise. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| Installation, Commissioning, Calibration & Validation Requirements |  |  |  |  |
| D7 | Pre-shipment test SHOULD demonstrate the system’s capability to achieve all process parameters outlined above. Pre-Shipment test results should be shared as agreed between Tyndall and the Tenderer, prior to shipment. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| D8 | The tenderer MUST commission the system and verify full functionality at the Tyndall National Institute site. The buyer's engineering staff (both process and equipment) will provide reasonably requested assistance to the supplier in this task. Tyndall will also provide for any labour, chemicals, wafers, power, lubricants, water, fuel and materials reasonably requested for the purpose of the process installation. Tenderer has to make a written request for this support at least four weeks in advance of installation. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| D9 | The facility requirements for this piece of equipment MUST be supplied to Tyndall a minimum of sixteen (16) weeks in advance of the equipment delivery. These facilities changes will be installed to the specifications supplied to Tyndall by the tenderer. Any deviation to the supplied specification on the delivered equipment causing a reconfiguration of the facility supply will be at the tenderer’s cost. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| D10 | The proposed system MUST be suitable for use with the 400 V, 50 Hz, 3-phase electrical supply at the proposed installation site. If a transformer unit is required to achieve this, it MUST be supplied as part of the core package. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| D11 | Tenderers SHOULD also note that the following facilities will be available at the installation site: - 3-phase electrical supply - Process cooling water supply - De-ionised water supply - Waste water drainage - Compressed air supply - Air exhaust. | Confirm (Yes / No) | Insert Comment |  |
| D12 | Tyndall Cleanroom The tenderer MUST provide footprint drawings for tool and ancillary equipment (e.g. control consoles etc.) required for operating the tool which MUST be located in the cleanroom. The available cleanroom space will be limited so tools and ancillary equipment MUST fit within a space approx. 5 m2 per tool (total footprint for the suite of tools being tendered-for is approx. 20 m2. Vacuum pumps / chillers / equipment racks etc. will be located remotely, positioned in a sub-fab with up-to a 10m vertical drop and up-to a 10m horizontal run to the point-of-use at the tool. This equipment MUST fit within a space approx. 4 m2 per tool (total footprint for the suite of tools being tendered-for is approx. 16 m2 Please see guideline footprint drawings of the cleanroom and sub-fab locations where the suite of tools will be installed in attached Appendices; Appendix A- Cleanroom Footprint & Appendix B – Sub-Fab Footprint. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| D13 | The tools MUST have thru-wall configuration for installation in the Silicon Fabrication cleanroom. All hardware (e.g. front panel, keyboard(s), monitor(s), etc.) required for thru-wall configuration MUST be included in the core package, as outlined in requirement A62 above. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| D14 | The installation, commissioning, and calibration of the equipment MUST be carried out on-site by a qualified engineer provided by the Tenderer. This entails the system being transported to the lab, unpacked, set up and demonstrated to meet the requirements as outlined in section A. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
| D15 | Prior to installation, the machine may be inspected at the Tenderer site by a Tyndall engineer to ensure that it meets the agreed specifications. Particular attention will be paid to the facilities links to ensure that the equipment is in a state ready for connection to the installed facilities at the proposed location in Tyndall. This inspection will also ensure that the machine as built corresponds exactly to the supplied drawings. Please confirm compliance. | Confirm (Yes / No) | Insert Comment |  |
| D16 | Equipment start-up is tenderer’s responsibility. Prior to start up, the supplier may inspect the facilities to determine whether they meet requirements. Start-up of the equipment shall indicate the tenderer’s acceptance of the facilities installation. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| D17 | Tyndall will carry out a Functional Acceptance Test. The equipment functionality test consists of the tenderer’s equipment start-up test. Prerequisites for the functional acceptance are: - Complete Delivery and installation - Resolution of all outstanding issues resulting from pre-shipment test, if applicable - Fulfilment of the following tests: - Shut-down test (ESO operation and re-start) Please confirm acceptance of this. | Confirm (Yes / No) | Insert Comment |  |
| D18 | The Tenderer MUST provide a qualified engineer to perform the SAT (Site Acceptance Test) post commissioning. The validation tests will be done in partnership with experienced Tyndall process engineers to ensure that the system performance has not been affected by shipping. These tests may also be used for on-site training purposes. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| D19 | The acceptance tests SHOULD include substrate handling tests in sequence without alarm conditions; the number of handling cycles to be mutually agreed between buyer and tenderer based on the system spec. - Recovery from power outage, extract outage, vacuum outage, nitrogen outage. - Delivery of all required documentation. - A copy of the tenderer’s pre-shipment final systems test procedure. - Delivery of supplier specified consumables. If the functional acceptance test produces negative results, Tyndall shall be entitled to request that the equipment will be brought into the guaranteed condition within two (2) weeks of the failed test. Please confirm compliance and detail. | Confirm (Yes / No) | Insert Comment |  |
| D20 | SITE Acceptance Test After functional acceptance, the Site Acceptance of the equipment shall be done jointly by engineers assigned by Tyndall and by the tenderer. A final acceptance document shall be prepared with the following content: - The purchase order reference. - The procurement specification reference. - A list of outstanding issues, including an action time frame for resolving these issues and the division of responsibility for this work. - The signatures and functions of the accepting engineers on each side, the Head of Process and Product Development group and the Tyndall Maintenance Manager. - The date of acceptance. - Serial number of the equipment. If the items in the list are satisfied, the equipment will be accepted. If one or more items are not accepted, the tenderer will be required to rectify the deficiency or, alternatively, a partial acceptance may be agreed combined with an extension of the warranty period. Terms shall be negotiated between Tyndall and tenderer based on the results of the acceptance tests. If there is a serious problem on the system, a mutually agreed plan of correction will be generated to correct the problem. Tyndall may waive any portion of the plan. The plan may include the replacement of the system at the discretion of the tenderer. Please confirm compliance. | Confirm (Yes / No) | Insert Comment |  |
| D21 | Tenderers are requested to provide a summary of scheduled maintenance and calibration required to keep instrument working to the full specification as per manufactures recommendations. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| D22 | A signed certificate for successful site acceptance test MUST be provided at the end of the installation and training. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| Documentation |  |  |  |  |
| D23 | A detailed User Operational Manual for the unit and control system MUST be provided. This SHOULD include the following; - Full electrical and design specifications. - Manuals MUST be supplied in English language. Manuals MUST be provided in soft copy. | Confirm (Yes / No) | Insert Comment |  |
| D24 | A Maintenance Manual for the system MUST be provided to include the following; - Cleaning requirements, service - Maintenance & Calibration requirements. - Manuals MUST be supplied in English language. - Manuals MUST be provided in softcopy where possible. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| D25 | A Technical Specification for the Unit MUST be provided covering all major components of the system. The specification MUST be supplied in the English language and in (if possible) in soft copy. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| D26 | A Technical Specification for all instrument documentation relating to installation of hardware and software, hardware/software interfacing, service and maintenance, parts list, and software operation and editing documentation MUST be supplied in English language. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| D27 | Factory Acceptance Test The equipment operation MUST be validated FAT (Factory Acceptance Test) prior to shipment to Tyndall to ensure it meets performance specifications. - The equipment will be shipped to Tyndall National Institute, Cork when Manufacturer validation tests have proved the equipment performs to the required specifications. Validation certificates MUST be provided during installation and commissioning. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| - Energy, Environmental & Sustainable Considerations |  |  |  |  |
| Environmental & Sustainable Considerations |  |  |  |  |
| E1 | Please outline your company’s commitment to the sustainable life cycle of the tendered equipment in a manner that minimises adverse environmental impacts. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| E2 | Energy & Environmental Requirements: The Contracting Authority requires the equipment provided to be as energy and economically efficient as possible. The system SHOULD have low running costs allowing the equipment to be as energy & economically efficient as possible. Where applicable, and if possible, the equipment SHOULD include the following: - Use of water-cooled units, e.g. water cooled chillers / pumps instead of air cooled chillers/ pumps as the site recovers heat from the processing cooling load to heat other areas. - Eco mode vacuum pumps to enable nitrogen saving when system in standby, control of both the nitrogen flow and eco mode activation SHOULD be from the system software - Where possible, use of a standby mode/reduced power mode of operation when the system is not in use. - Programmable capability, i.e. if possible, equipment SHOULD come with the option of programmable timer so that processes can start overnight and the site can benefit from the reduced night-time electrical rates. Please provide any details on the above or other measures taken to improve energy and efficiency of the tool. Please provide details on any “warm-up” times required for the equipment to operate to specification. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| E3 | Tenderers SHOULD provide details of any takeback arrangements they may offer when the equipment reaches end of life and how that may be availed of. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| E4 | Tenderers SHOULD indicate Typical Energy Consumption during operational time while the equipment is in use. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| E5 | Where possible, use of a standby mode/reduced power mode of operation when the system is not in use. The tenderer SHOULD specify the stand-by energy consumption. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| E6 | It is desirable that all Tender requirement is delivered using environmentally friendly Packaging Materials; packaging, tape, labels and stretch wraps. Ideally, this packaging SHOULD avoid any single use plastic and be recovered or at a minimum reusable. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| E7 | Tenderers SHOULD provide details of expected lifetime of the equipment required. The Tenderer MUST state how long the model of equipment has been on the market and indicate the expected working life of the proposed models. Longevity of equipment etc. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
| - Ultimate Cost. (Refer to Appendix 2: Pricing Schedule for Completion) |  |  |  |

# Lot 2 Plasma Etcher (Metals / Piezoelectrics)

| - Technical Merit |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| General/Functional Requirements |  |  |  |  |  |
|  |  | The Tenderer MUST supply a single-wafer processing tool for etching of metallic (e.g. Aluminium, Molybdenum etc.) and Piezoelectric (e.g. Aluminium Nitride, Scandium Aluminium Nitride etc.) materials. Please confirm and detail costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
|  |  | UCC will consider tenders offering new instrumentation or refurbished and factory calibrated instrumentation. Tenderers MUST state whether new or ex demo/used systems is proposed. If the Instrument is ex demo it MUST be no more than three (3) years old and all parts and components MUST be available for at least the next ten (10) years. It MUST include a full refurbishment history with appropriate calibration certificates and full warranty. Please confirm and detail | Confirm (Yes / No) | Insert Comment |  |
|  |  | The tool MUST be capable of etching the following materials: - Metallic materials (e.g. Aluminium aka Al, Molybdenum aka Mo, etc.) and - Piezoelectric materials (e.g. Aluminium Nitride aka AlN, Scandium Aluminium Nitride aka ScAlN etc.) Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  |  | The tool MUST be capable of etching the materials listed in A3 above for a variety of micro and/or nanoscale structures with nanometre precision and excellent wafer-scale uniformity (Within-Wafer-Uniformity aka WiW) with the following parameters; - Photoresist-patterned structures with dimensions - min. ≥ 100 nm - Etch depths - min. ≥ 20 nm - max. ≤ 10 µm | Confirm (Yes / No) | Insert Comment |  |
|  |  | The tool MUST be suitable to run both flourine-based and chlorine-based chemistries. Any additional hardware (e.g. removable process-chamber liners etc.) required to allow for better process integrity and rapid switch-over between the different etch chemistries MUST be included as part of the core package. Please confirm and detail costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
|  |  | The Tenderer MUST provide plasma cleaning & conditioning recipes and/or best-known method (BKM) procedures for conditioning the chamber when switching between fluorine-based and chlorine-based chemistries as part of the core package. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  |  | The tool MUST include a vacuum loadlock chamber as part of the core package. Tools with open-loading process-chambers will NOT be considered suitable for this tender. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  |  | The tool MUST be capable of etching silicon-based & dielectric materials on silicon and other substrates in wafer form (flatted or notched as-per SEMI M1-1117) for the following; - 100 mm diameter - 200 mm diameter Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  |  | The tool MUST be capable of handling the following substrate types in wafer form and for the sizes listed in A8 above as a minimum requirement: - Silicon (Si) - Silicon-on-Insulator (SOI) - Glass / Fused Silica - Silicon Carbide (SiC) - Germanium (Ge) Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  |  | The required etch performance of the tool MUST be achieved by inductively-coupled plasma (ICP) configuration with independent RF controls between the two sources. Bias control with high accuracy is considered essential for low-damage processes (e.g. polysilicon gate etch onto SiO2 etc.). Please confirm and detail. | Confirm (Yes / No) | Insert Comment |  |
|  |  | The tool MUST be equipped with endpoint detection (EPD) hardware & software (e.g. Optical Emission Spectroscopy aka OES or laser interferometry etc.) as part of the core package. The proposed endpoint solution(s) MUST have the capability to monitor / control both etching and chamber cleaning & conditioning. Please confirm and detail costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |  |
| Performance Requirements |  |  |  |  |  |
| - 12 | The tool MUST be capable of etching the materials listed in A3 above with excellent wafer-scale uniformity. Within-wafer variation (WiW) of the etch depth as measured at centre / edge points of wafer (with 10 mm edge-exclusion on both 100 mm and 200 mm diameter wafers) to be < ± 5%. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| - 13 | The tool MUST be capable of using the following mask materials for etching the materials listed in Recipes #4, #5 & #6 below as a minimum requirement: - Photoresist (e.g. optical and electron-beam), - Dielectric hard-masks (e.g. silicon-dioxide, silicon-nitride, amorphous-silicon) It is expected that photoresist will be used as the mask material for most applications. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| - 14 | Recipe # 4 (Al-Si1% interconnect etch) The tool MUST be capable of and a process recipe MUST be provided with the tool and demonstrated as part of the validation tests post at the Tyndall National Institute site for etching of ≤ 2 µm linewidth microstructures in Al-Si1% with soft-landing / minimal loss onto silicon-oxide (SiO) under-layer on both 100 mm and 200 mm silicon wafers. The sidewall profile is to be highly anisotropic (≥ 85°). Tenderers MUST clearly demonstrate how they meet the requirements outlined with results documented in the attached Appendix C – Etch Recipe Results Template, this must be clearly named as “Lot 2 – Recipe # 4 -Results”. The Tenderer is permitted to use its own test mask / layout and substrates and / or own etching results to demonstrate compliance. Requirements: - Mask layout: open area > 80% (light-field pattern) - Feature size & type: ≤ 2 µm linewidth interconnect - Mask materials: Photoresist (PR) - Aspect ratio: > 1:1 - Depth: ≤ 1 µm - Etched profile: anisotropic ≥ 85° and ≤ 90° or equivalent to the photoresist sidewall angle - Mask undercut: < 100 nm - Etched sidewall: smooth - Etch rate Al-Si1%: > 100 nm/min - Etch rate SiO: < 100 nm/min - Selectivity to Photoresist: > 2:1 (Al-Si1% etch-rate ~ 2x photoresist etch-rate) - Non-uniformity depth (100 mm wafer): < 5 (± %) - Non-uniformity depth (200 mm wafer): < 5 (± %) Tenderer is to specify the recommended process temperature and gases for this recipe Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| - 15 | Recipe # 5 (Mo top-electrode etch) The tool MUST be capable of and a process recipe MUST be provided with the tool and demonstrated as part of the validation tests post at the Tyndall National Institute site for etching of ≤ 2 µm linewidth electrode microstructures in Mo with soft-landing / minimal loss on Sc6%AlN under-layer on both 100 mm and 200 mm silicon wafers. The sidewall profile is to be highly anisotropic (≥ 85°). Tenderers MUST clearly demonstrate how they meet the requirements outlined with results documented in the attached Appendix C – Etch Recipe Results Template, this must be clearly named as “Lot 2 – Recipe # 5 -Results”. The Tenderer is permitted to use its own test mask / layout and substrates and / or own etching results to demonstrate compliance. Requirements: - Mask layout: open area > 80% (light-field pattern) - Feature size & type: ≤ 2 µm linewidth interdigitated electrode - Mask materials: Photoresist (PR) - Aspect ratio: > 1:1 - Depth: ≤ 500 nm - Etched profile: anisotropic ≥ 85° and ≤ 90° or equivalent to the photoresist sidewall angle - Mask undercut: < 100 nm - Etched sidewall: smooth - Etch rate Mo: ≥ 100 nm/min - Etch rate Sc6%AlN: < 50 nm/min - Selectivity to Photoresist: > 1:1 (SiO2 etch-rate ~ photoresist etch-rate) - Non-uniformity depth (100 mm wafer): < 5 (± %) - Non-uniformity depth (200 mm wafer): < 5 (± %) Tenderer is to specify the recommended process temperature and gases for this recipe. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| - 16 | Recipe # 6 (Sc6%AlN resonator etch) The tool MUST be capable of and a process recipe MUST be provided with the tool and demonstrated as part of the validation tests post at the Tyndall National Institute site for etching of ≤ 5 µm linewidth microstructures in Sc6%AlN with good selectivity to Mo under-layer on 100 mm and 200 mm silicon wafers. The sidewall profile is to be highly anisotropic (≥ 85°). Tenderers MUST clearly demonstrate how they meet the requirements outlined with results documented in the attached Appendix C – Etch Recipe Results Template, this must be clearly named as “Lot 2 – Recipe # 6 -Results”. The Tenderer is permitted to use its own test mask / layout and substrates and / or own etching results to demonstrate compliance. Requirements: - Mask layout: open area > 80% (light-field pattern) - Feature size & type: ≤ 5 µm linewidth resonators - Mask materials: Photoresist (PR) - Aspect ratio: > 1:1 - Depth: ≥ 1 µm - Etched profile: sloped > 85° and < 90° - Mask undercut: < 100 nm - Etched sidewall: smooth - Etch rate Sc6%AlN: > 100 nm/min - Etch rate Mo: < 50 nm/min - Selectivity to Photoresist: > 1:1 (Sc6%AlN etch-rate ~ photoresist etch-rate) - Non-uniformity depth (100 mm wafer): < 5 (± %) - Non-uniformity depth (200 mm wafer): < 5 (± %) Tenderer is to specify the recommended process temperature and gases for this recipe. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| - 17 | The tool SHOULD have the capability to run the etch process recipes as listed in Recipes #4, #5 & #6 above at room temperature (20°C) and this will be considered an advantage and scored higher. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| - 18 | The tool SHOULD be capable of etching the following materials with the same configuration as the recipes listed in Recipes #4, #5 & #6 above and this will be considered an advantage and scored higher. - Titanium (Ti) - Titanium Nitride (TiN) - Tungsten (W) - Tungsten-Titanium alloys - Aluminium Oxide (Al2O3) - Hafnium Oxide (HfO2) If any additional hardware (e.g. chamber-liners etc.) is required for etching these materials, the Tenderer MUST identify clearly in their response and cost separately. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) | Insert Comment |  |
| - 19 | It is desirable that the system be capable of etching the following materials, which do not readily form volatile compounds, with the same configuration as for the recipes listed in Recipes #4, #5 & #6 above and will be considered an advantage and scored higher. - Platinum (Pt) - Nickel (Ni) - Chrome (Cr) If any additional hardware (e.g. chamber-liners etc.) is required for etching these materials, the Tenderer MUST identify clearly in their response and cost separately. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) | Insert Comment |  |
| - 20 | The Tenderer MUST provide plasma cleaning recipes for inter-wafer and inter-batch cleaning / polymer removal in the process-chamber as part of the core package. Please state if separate clean recipes for fluorine-based and chlorine-based chemistries are required and this MUST be clearly specified by the tenderer. Tenderer is to specify the recommended process temperature and gases for these cleaning recipes. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| - 21 | It will be considered an advantage and scored higher if one or more of the gases required for process-chamber cleaning can also be used as an etch process gases (e.g. CF4 / O2 / SF6 etc.), however if other gases are recommended for more efficient cleaning recipes, then the tenderer SHOULD specify those gases as their preferred option. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| - 22 | Recipe Support Tenderer MUST have ready access on request to a library of characterised etch / cleaning / conditioning recipes for the tool and assistance in the form of BKMs / recipes / technical support by phone & email when required in setting-up new recipes. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| - 23 | Roadmap & Flexibility: Clarity of roadmap for the Tenderer’s proposed tool is required, indicating the potential future upgrade options that will be available as well as highlighting existing upgrade options for the system proposed. As with all pieces of equipment in UCC, the ability to upgrade in the future to add additional functionality not currently specified is very important, i.e., the tool SHOULD have a future roadmap of development. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| - 24 | The tenderer MUST state when the model was introduced and when the system is likely to be superseded. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| Hardware/Software Requirements |  |  |  |  |  |
|  | The system MUST be equipped with an aluminium process-chamber capable of being heated to > 50°C. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The process-chamber MUST be equipped with viewing port(s) for e.g. optical emission and laser reflection endpoint detection (EPD) systems at e.g. side and top chamber positions, and to directly view the plasma for stability. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The proposed endpoint solution(s) MUST be fully integrated with the system software. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | Tenderers SHOULD also supply further options of other EPD systems available. Please confirm and detail if any additional cost applies to additional EPD options that can be provided. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) | Insert Comment |  |
|  | Tenderers MUST identify clearly in their response if removable shields / liners for the process-chamber are required to allow for better process integrity / performance and to allow for rapid switch-over between fluorine-based and chlorine-based chemistries. Tenderers MUST include a minimum of two (2) sets of removable shields / liners as part of the core package if these are required for running the etch recipes as listed in Recipes #4, #5 & #6 above. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Any such additional costs must be detailed in Appendix 2. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The system MUST have a turbo pump to sustain low pressure (e.g. ≥ 1 mTorr) etch recipes / high flow-rate (e.g. > 100 sccm) cleaning recipes as part of the core package. Tenderers MUST specify the turbo pump brands and model numbers which will be provided with the tool as part of the core package. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | Tenders SHOULD also provide information on further options of other turbo pumps available for the process-chamber. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The system MUST have a dry backing pump for the process-chamber. Tenderers MUST specify the pump brands and model numbers which will be provided with the tool as part of the core package. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The dry backing pump for the process-chamber MUST be located remotely, positioned in a sub-fab with up-to a 10m vertical drop and up-to a 10m horizontal run to the point-of-use at the tool. Please see attached Appendix B – Sub-Fab Footprint for further details on the sub-fab footprint. The dry backing pump MUST have sufficient performance when working in-conjunction with the turbo pump to sustain low pressure (e.g. ≥ 1 mTorr) etch recipes as outlined in Recipes #4, #5 & #6 above and all necessary cleaning / conditioning recipes under these conditions. Sufficient length (minimum = 10 m) of electrical cable / plumbing etc. MUST be supplied to achieve this. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) | Insert Comment |  |
|  | Tenders SHOULD also provide information on further options of other dry backing pumps available for the process-chamber. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The process-chamber MUST recover to a pressure of ≤ 1 mTorr and a leak-up rate ≤ 2 mTorr/min after a wet-clean in less than two (2) hours. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The system MUST have a dry backing pump for the loadlock. Tenderers MUST specify the pump brands and model numbers which will be provided with the tool as part of the core package. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The dry backing pump for the loadlock MUST be located remotely, positioned in a sub-fab with up-to a 10m vertical drop and up-to a 10m horizontal run to the point-of-use at the tool. Please see attached Appendix B – Sub-Fab Footprint for further detailed on the sub-fab footprint. The dry backing pump MUST have sufficient performance to sustain low pressure (e.g. ≤ 100 mTorr) for wafer transfer into the process-chamber under these conditions. Sufficient length (minimum = 10 m) of electrical cable / plumbing etc. MUST be supplied to achieve this. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) | Insert Comment |  |
|  | Tenderers SHOULD also provide information on further options of other dry backing pumps available for the loadlock. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation |  | Confirm (Yes / No) | Insert Comment |  |
|  | Loadlock pump-down time SHOULD be less than 120 seconds, and no longer than 300 seconds. Faster pump-down time will be considered an advantage and scored higher. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The system MUST have a heated foreline to remove etch and cleaning plasma by-products from the process-chamber. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The system MUST have an automatic throttle valve for pressure control in the process-chamber. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The system MUST be water-cooled. Process cooling-water supply is provided by Tyndall, see section D. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The process-chamber platen / wafer-electrode MUST be heater-chiller controlled to allow operation in the range 0°C and +50°C with control of ±2°C from set point. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The heater-chiller unit MUST be located remotely, positioned in a sub-fab with up-to a 10 m vertical drop and up-to a 10 m horizontal run to the point-of-use at the tool. Please see attached Appendix B – Sub-Fab Footprint for further details on the sub-fab footprint. The heater-chiller unit MUST have sufficient performance to sustain platen / wafer-electrode temperature in the range 0°C and +50°C with control of ±2°C under these conditions. Sufficient length (minimum = 10 m) of electrical cable / plumbing etc. MUST be supplied to achieve this. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) | Insert Comment |  |
|  | A non-return chiller valve MUST be included to prevent back-flow to reservoir. Tenderer MUST specify what model heater/chiller will be included and the coolant type (e.g. Galden or equivalent) recommended for the heater-chiller units. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | System MUST be equipped with a fully auto-matched RF generator for each plasma source. The RF generators MUST have capability for plasma strike and sustain tuned plasmas operating over a wide processing pressure range ≥ 1 mTorr. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | Tenderers MUST specify details of the RF generator brand & model, frequency, power and operating range for each plasma source which will be provided as part of the core package. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The RF generators / racks may be located remotely, positioned in a sub-fab with up-to a 10 m vertical drop and up-to a 10 m horizontal run to the point-of-use at the tool. Please see attached Appendix B – Sub-Fab Footprint for further details on the sub-fab footprint. Sufficient length (minimum = 10 m) of electrical cable / plumbing etc. MUST be supplied to achieve this. The RF generators / racks MUST have sufficient performance for plasma strike and sustain tuned plasmas under these conditions. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) | Insert Comment |  |
|  | RF Generator racks MUST be supplied by the Tenderer, if required for the assembly of the system. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) | Insert Comment |  |
|  | Tenderers SHOULD also provide information on further options of other RF generators available. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. If any options to purchase are available, these are optional to purchase by the Contracting Authority and will not form part of the Cost evaluation. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The RF generators MUST have the capability to automatically attenuate the output power as required by the process recipe (e.g. low bias RF power for high-selectivity etching, high ICP RF power for chamber-cleaning etc.). Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The tool MUST be configured for single-wafer processing (wafers are etched individually). The hardware changes / procedures required to switch from running one wafer-diameter to another MUST be outlined clearly by the tenderer with all necessary hardware provided as part of the core package. The ability to switch from e.g. 100 mm to 200 mm substrate diameter with minimal downtime and / or hardware changes will be considered an advantage and scored higher. Please confirm and detail changeover procedure and time to complete. Any such additional costs must be detailed in Appendix 2. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The tool MUST have capability to handle wafer thicknesses in the range ≥ 0.2 mm and ≤ 1 mm as a minimum requirement. An ability to also handle substrates > 1 mm thickness with the same setup will be considered an advantage and scored higher. Please specify minimum and maximum thickness in tender response and identify any hardware changes (e.g. separate wafer-carriers for different thickness) required to allow a range of thin / thick substrates to be processed. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The tool SHOULD have capability to handle bonded-wafer pairs, perforated wafers, or bowed wafers (with bow / warp > 50 µm). This will be considered an advantage and scored higher. Please identify any hardware changes (e.g. separate wafer-carriers etc.) required to allow for handling these non-standard types of substrates. If any additional hardware (e.g. wafer-carriers etc.) is required for handling these non-standard types of substrates, the Tenderer MUST identify clearly in their response and cost separately. Tenderers are to provide details and pricing for the proposed option within the core package in the Appendix 2, Pricing Schedule. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation |  | Confirm (Yes / No) | Insert Comment |  |
|  | The tool MUST accommodate wafers as set out in requirements A8 - A9 above. Helium gas heat-transfer for wafer back-side cooling with electrostatic or mechanical clamping MUST be included. Helium transfer MUST be pressure-controlled and the helium flow rate SHOULD automatically be checked prior to wafer etch. The He pressure MUST be stable at e.g. 10 Torr / 15 Torr during the clamping of a wafer and the He flow leakage may not affect the process's results. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The system MUST have a minimum eight (8) gas lines. The capability to have more than eight gas lines will be considered an advantage and scored higher. Suggested process gases include, but are not limited to the following; SF6, C4F8, O2, BCl3, Ar, Cl2 The tenderer MUST specify their recommended gases for the recipes outlined in requirements Recipes #4, #5 & #6 above and for plasma cleaning & conditioning of the process-chamber. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | In the specific case(s) of low vapour pressure / low boiling point gases (e.g. BCl3) which are susceptible to condensation, heated gas-delivery hardware such as gas-lines, gas-box, MFCs etc. MUST be included as part of the core package. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) | Insert Comment |  |
|  | In the specific case of oxygen (O2) the requirements for e.g. low flow-rate for high-selectivity etching of silicon onto silicon dioxide and high flow-rate for chamber plasma cleaning / chamber plasma conditioning may necessitate two separate MFCs, so this should be clearly specified by the Tenderer. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The tenderer MUST specify recommended flow size for each gas Mass Flow Controller (MFC). MFCs supplied MUST be rated (in terms of e.g. seals etc.) and calibrated for the recommended gases as nominated by the Tenderers under requirement A55. Digital MFCs are preferred and will be considered an advantage and scored higher. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The capability to run automatic interleaved and / or wafer-less clean MUST be available in the software. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | Whole system MUST be fully interlocked to protect against any equipment or service failure. System MUST be designed to fail-safe with system left under vacuum. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | Electrical equipment racks may need to be located remotely, positioned in a sub-fab with up-to a 10 m vertical drop and up-to a 10 m horizontal run to the point-of-use at the tool and MUST be capable of full functionality under these conditions. Sufficient length (minimum = 10 m) of electrical cable MUST be supplied to allow the required performance. Please see Appendix B – Sub-Fab Footprint for further detailed on the sub-fab footprint. Confirmation is required that the operation of the tool and etch quality will not be impacted by remote location of the electronic control cabinets. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) | Insert Comment |  |
|  | Electrical equipment racks MUST be supplied by the Tenderer, if required for the assembly of the system. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) | Insert Comment |  |
|  | A minimum of one (1) computer and console MUST be included in core package, and MUST use Windows-based OS, Windows 10 or higher. Please describe the software package proposed, including name and version. Please also specify the projected lifespan of the current control system architecture. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The tool will be installed in the cleanroom in a thru-wall configuration. ALL computer(s) / console(s) required for tool operation from the front (process bay) and rear (service chase) side MUST be included in the core package. Additional console(s) / computer(s) for tool operation from the rear (service chase) side MUST also be included in the core package, or a suitable alternative such as a rotating screen(s) that are visible from both bay and chase with associated keyboard / mouse control from both sides. Please see attached Appendix A -Cleanroom Footprint for further details on the cleanroom footprint. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The tenderer MUST provide all control and analysis software required for tool operations. All tool operations, including wafer transfer to / from the loadlock to the process-chamber MUST be software-controlled and allow for both auto & manual operation. Any such additional costs must be detailed in Appendix 2. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The Plasma Etcher equipment SHOULD have perpetual licence-free software or no restriction on the number of licences needed for our use of the system i.e., licences for multiple users and computer installations for off-site use if this is possible. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | All software updates to the system, which become available during the life of the contract, MUST be provided free of charge during the warranty period and for a minimum period of eight (8) years after warranty expiry. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) | Insert Comment |  |
|  | Continued development of the software SHOULD be demonstrated by the Tenderer, with facilities for updating systems in use. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The system MUST be capable of remote access via network connection (using e.g. TeamViewer software or equivalent) for tool support in event of problems. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The software MUST be capable of data-logging of process parameters, and archiving data-logs. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | It MUST be easy to graphically display and export saved etch process parameter data preferably into Excel spreadsheets. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The system MUST have an easy-to-use interface and needs to be suitable for a multi-user (i.e. different access levels) environment. Various levels of access are required e.g. Operator, Supervisor, Maintenance. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | The software SHOULD have the facility to lock certain recipes i.e. not editable by users other than Process Engineer / Supervisor / Maintenance etc. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | Software MUST include backup and restore capability (export/import) for process recipes and machine settings. Software SHOULD include backup (image) capability to be used in case of hard disk drive (HDD) failure. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
|  | Software MUST allow operator access to data files while the system is running under standard operational conditions. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| - Quality, Health & Safety & Validation |  |  |  |  |  |
| Quality Requirements |  |  |  |  |  |
|  | Compliance with all relevant European health and safety standards, quality assurance and protocols are MANDATORY. All Goods and Services supplied in accordance with the contract MUST comply with all necessary standards, i.e. Irish Standards Institution, European Union, CE, International Standards Organisation or other applicable standard. The Tenderer MUST, on the request of the buyer, provide a copy of the standard free of charge. Please confirm & detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  | If there are manufacturing defects known to the tenderer, they MUST identify such defects and state their policy regarding the repair of known defects. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  | The Tenderer SHOULD, as part of the submission, include copies of any relevant testing or compliance certificates for the Goods. (For example, Factory acceptance tests, Quality assurance test etc.) Please confirm and detail |  | Confirm (Yes / No) |  | Insert Comment |
| Health & Safety Requirements |  |  |  |  |  |
|  | All goods supplied in accordance with the contract MUST be of good construction, sound material, of adequate strength, free of design defects and manufactured to be safe and without risk when properly used and software MUST allow operator access to data files while the system is running under standard operational conditions. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  | The tool SHOULD meet the criteria of the Semi S2-0821Eb Environmental, Health, and Safety Guideline for Semiconductor Manufacturing Equipment in order to ensure easier and safer installation and maintenance from a gas safety perspective. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  | The Tenderer MUST, as part of the submission, include Laser Safety Warning labels and copies of any relevant testing or compliance certificates for the Goods. Confirm compliance and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  | The tools MUST knowingly not present any Health and Safety hazards. Any safety & hazard warning labelling on the proposed system MUST be in the English Language. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  | Tenderers MUST advise if there are any hazards associated with the operation or maintenance of this instrument? (e.g. manual handing, chemicals, etc.). Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| C. Warranty/Spare Parts, After-Sales & Training Support |  |  |  |  |  |
| Warranty & Customer Support, Repair/Breakdown Requirements |  |  |  |  |  |
|  | Warranty Duration: At least one (1) year’s warranty from Site Acceptance Test (SAT) MUST be included for each lot as part of the core package and the quoted price is to include all hardware, software, spare parts, labour, call outs, visit, travel and parts. Warranties of greater than 1 year are considered highly desirable and will be scored higher. Detail specific inclusions and exclusions of warranty. Preventative Maintenance MUST be provided for at least one (1) year and a list of checks carried out at Preventative Maintenance MUST be provided, along with any information about which (if any) parts are routinely replaced. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  | Insert Comment |
|  | If UCC decide to extend the warranty period, Tenderers are asked to outline costs of any additional warranty. The Tenderer MUST state the details of any extended warranty and details of Preventative Maintenance and calibration & servicing check and any associated costs. Detail specific inclusions and exclusions of extended warranty and a list of checks carried out at Preventative Maintenance SHOULD be provided, along with any information about which (if any) parts are routinely replaced. Please include price per annum, details of coverage and maximum term of coverage. Please include price per annum for year 2, 3, & 4, details of coverage and maximum term of coverage. Please confirm and detail and specify what is included in the warranty options outlined in the Appendix 2 pricing schedule, Optional to Purchase. Detail specific inclusions and exclusions of extended warranty. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) |  | Insert Comment |
|  | Please provide separately a list and approximate cost of any consumable items/parts that are to be replaced during each service visit. Please update Appendix 2 Pricing Schedule – Consumables Parts sheet with detail. Prices provided for consumable items/parts MUST be valid for a minimum of one (1) year. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) |  | Insert Comment |
|  | Please advise of any items/part of the instrument that require routine replacement during the normal lifetime of the instrument, the timeframe in which they require replacing and if such items will be supplied with the instrument or as part of the warranty. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  | The tenderer MUST provide a description of the arrangements the arrangements and lead times for guaranteed responses times in the event of equipment breakdown both in relation to phone support, online/remote support, the provision of technical support, annual service/ maintenance, emergency call-out response, and spare parts. Please detail lead time for arrival on site by a fully trained service engineer during the warranty period(s) and post-warranty. This response time SHOULD be no more than 24 hours via phone call or remote service, and within seven (7) working days for onsite repair. Please confirm and detail information on remote diagnostic / telecom-based service support. |  | Confirm (Yes / No) |  | Insert Comment |
|  | Outside of Warranty Period: The Tenderer MUST provide typical prices and lead time for servicing and repairs (including travel, subsistence and price per day) charged outside a warranty/ maintenance contract. Detail specific inclusions and exclusions outside of warranty period. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) |  | Insert Comment |
|  | If UCC decide to take out an additional Preventative Maintenance Agreement / SLA, tenderers are asked to outline Options for years 2, 3 & 4. Option to include: - On-going Calibration to appropriate schedule - preventative maintenance - remote diagnostics of faults, if applicable - repair of the Instrument in case of breakdown) Tenderers are asked to outline full details of what is included and to outline Cost of maintenance and calibration for each year option offered. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) |  | Insert Comment |
|  | The tenderer MUST have dedicated fully trained and qualified service engineering specialists at their disposal who can respond to all service calls and any form of breakdown call. This remote response time SHOULD be the same day via phone call or remote service, and within seven (7) working days for onsite repair. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  | Please outline the company’s Customer Support Proposal and escalation policy. Tenderers MUST describe what process is in place for escalating a call-out, if a breakdown cannot be resolved on the first visit. Tenderer SHOULD provide a list of approved service and calibration agents. Please confirm and detail customer support arrangements, service and calibration agents. Additional support and service options should also be outlined and pricing noted within the Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) |  | Insert Comment |
|  | The Tenderer MUST indicate the arrangements for and specify the level of experience of applications specialists and engineers who will be in charge of after sales service. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  | The Tenderer SHOULD specify the lead time for repairs during the following scenarios. - warranty period(s) - post-warranty, - with service contract - without service contract Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  | Please note SHOULD installation, commissioning, service visits or availability of spare parts be affected in the event of a Pandemic crisis it is expected that the warranty period will be extended to facilitate the service downtime endured by the Contracting Authority. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  | Warranty commences post full installation and sign-off of tool acceptance by nominated UCC representative. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| Replacement & Spare Parts & Equipment Consumables |  |  |  |  |  |
|  | Spare Parts: Availability of all spare parts SHOULD be maintained 10 years after the warranty period. Please indicate the availability and lead time of spare part stock centres, with maximum lead-time for any part. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  | Critical spares MUST be held at a location no more than 7 Days delivery time from Tyndall. Spare parts to be held offsite with the tenderer and essential parts if necessary to be held Free of Charge (FOC) at Tyndall. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  | Consumable Parts: Sufficient consumable parts MUST be supplied for the initial installation, commissioning and training. These consumable parts MUST be provided free of charge and at the point of commissioning. Full one (1) year of consumable parts to be supplied post commissioning, including any removable shields / liners for process-chamber. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  | Insert Comment |
|  | Recommended Spare Lists: The tenderer MUST supply a full list of recommended spare parts and consumables for the proposed equipment, based on knowledge of equipment design and component lifetimes. This SHOULD include typical consumables and high turnover parts. Please detail in Appendix 2, Pricing Schedule- Consumable Parts. Prices provided for consumable items/parts MUST be valid for a minimum of one (1) year. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) |  | Insert Comment |
|  | The tenderer SHOULD supply, where information is available, to the buyer a schedule of spare parts showing in tabular form or, otherwise, the following information on OEM parts: - Names and address of all OEM suppliers - Part Name - Part Description - OEM Supplier Part Number. Please confirm and detail, attaching information requested above. |  | Confirm (Yes / No) |  | Insert Comment |
|  | The tenderer MUST provide confirmation of the availability of service for at least five (5) years after commissioning. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| Training & Technical Support Proposal |  |  |  |  |  |
|  | The successful tenderer MUST provide an initial training programme for UCC personnel on-site at Tyndall, for a minimum of one (1) day each for operations and maintenance, with a minimum of three (3) users fully trained to safely operate and maintain the system as part of the core package. The training programme MUST be carried out by a suitably qualified specialist at the tool after the initial installation, commissioning is completed. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  | Insert Comment |
|  | Initial training provided MUST include, but not be limited to: - Demonstration of the correct operation of all relevant equipment of the systems - Routine Maintenance requirements - Basic Troubleshooting - Data processing and management - Software and hardware familiarisation Dates, timing and method of training are to be confirmed between UCC and the winning tenderer. The process validation tests may also be used for on-site training purposes. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  | The tenderer MUST specifically outline the level and duration of application support offered with the instrument where “application support” refers to the expert technical support (on-line or on-site) available from the manufacturer for using the instrument in novel research applications. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  | This equipment will have multiple users over the lifetime of the tool. Tenderers MUST provide details and pricing for their proposed option for Operations / Applications / Maintenance Training at all levels (e.g. Level1, Level2, Level3 etc.) on-site or at their facilities. Training offered over the full lifetime of the equipment will be considered an advantage and scored higher. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) |  | Insert Comment |
|  | Relevant training and maintenance manuals in English MUST also be provided with the instrument. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| - Delivery Plan |  |  |  |  |  |
| Lead Time & Delivery |  |  |  |  |  |
| D1 | The equipment SHOULD be delivered within twelve (12) months of receipt of Purchase Order. The equipment SHOULD be delivered within the stated lead time requirement from receipt of Purchase Order. (Successful tenderer for each lot may be required to liaise with each other for coordination of delivery and installation of the equipment) - Tenderers MUST provide full details of lead times for the delivery of equipment including installation and commissioning (if applicable), from receipt of the purchase order. - Tenderers MUST provide the facility requirements to UCC a minimum of 16 weeks in advance of the equipment delivery. - Tenderers MUST provide advance notice (two weeks) of any requirements for tool unloading, number of containers and weight of each. Height, length and width of containers/crates to be provided as well. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| D2 | Delivery MUST include all shipping charges, insurance, customs clearance and all aspects as outlined in DAP (Delivery At Place) incoterms. Delivery MUST be to Tyndall National Institute, UCC: - Silicon Fabrication Laboratory (Room A1.1.04), - Tyndall National Institute, - University College Cork, - Lee Maltings, Dyke Parade, - Cork T12 R5CP, - Ireland. Advance notice MUST be given (2 weeks) of any requirements for tool unloading, number of containers and weight of each. Height, length and width of containers/crates MUST be provided as well. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) | Insert Comment |  |
| D3 | The tenderer MUST provide full details of lead times for the delivery of equipment including installation and commissioning (if applicable), from receipt of the purchase order. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| D4 | The tenderer MUST provide details of the logistics/supply chain to the point of customer delivery, i.e. where the goods will come from, how the goods will be transported and what borders the goods will cross. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| D5 | The tenderer MUST provide details of any actions that have been taken to ensure the future timely supply and delivery of goods and services, including ongoing supply of replacement/spare parts and consumables. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| D6 | The tenderer MUST provide details of their contingency plans (alternative sources/routes of supply) for any delays in delivery should they arise. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| Installation, Commissioning, Calibration & Validation Requirements |  |  |  |  |  |
| D7 | Pre-shipment test SHOULD demonstrate the system’s capability to achieve all process parameters outlined above. Pre-Shipment test results should be shared as agreed between Tyndall and the Tenderer, prior to shipment. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| D8 | The tenderer MUST commission the system and verify full functionality at the Tyndall National Institute site. The buyer's engineering staff (both process and equipment) will provide reasonably requested assistance to the supplier in this task. Tyndall will also provide for any labour, chemicals, wafers, power, lubricants, water, fuel and materials reasonably requested for the purpose of the process installation. Tenderer has to make a written request for this support at least four weeks in advance of installation. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| D9 | The facility requirements for this piece of equipment MUST be supplied to Tyndall a minimum of sixteen (16) weeks in advance of the equipment delivery. These facilities changes will be installed to the specifications supplied to Tyndall by the tenderer. Any deviation to the supplied specification on the delivered equipment causing a reconfiguration of the facility supply will be at the tenderer’s cost. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| D10 | The proposed system MUST be suitable for use with the 400 V, 50 Hz, 3-phase electrical supply at the proposed installation site. If a transformer unit is required to achieve this, it MUST be supplied as part of the core package. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| D11 | Tenderers SHOULD also note that the following facilities will be available at the installation site: - 3-phase electrical supply as per D10 above - Process cooling water supply - De-ionised water supply - Waste water drainage - Compressed air supply - Air exhaust. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| D12 | Tyndall Cleanroom The tenderer MUST provide footprint drawings for tool and ancillary equipment (e.g. control consoles etc.) required for operating the tool which MUST be located in the cleanroom. The available cleanroom space will be limited so tools and ancillary equipment MUST fit within a space approx. 5 m2 per tool (total footprint for the suite of tools being tendered-for is approx. 20 m2. Vacuum pumps / chillers / equipment racks etc. will be located remotely, positioned in a sub-fab with up-to a 10m vertical drop and up-to a 10m horizontal run to the point-of-use at the tool. This equipment MUST fit within a space approx. 4 m2 per tool (total footprint for the suite of tools being tendered-for is approx. 16 m2 Please see guideline footprint drawings of the cleanroom and sub-fab locations where the suite of tools will be installed in attached Appendices; Appendix A- Cleanroom Footprint & Appendix B – Sub-Fab Footprint. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| D13 | The tools MUST have thru-wall configuration for installation in the Silicon Fabrication cleanroom. All hardware (e.g. front panel, keyboard(s), monitor(s), etc.) required for thru-wall configuration MUST be included in the core package, as outlined in requirement A62 above. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| D14 | The installation, commissioning, and calibration of the equipment MUST be carried out on-site by a qualified engineer provided by the Tenderer. This entails the system being transported to the lab, unpacked, set up and demonstrated to meet the requirements as outlined in section A. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) | Insert Comment |  |
| D15 | Prior to installation, the machine may be inspected at the Tenderer site by a Tyndall engineer to ensure that it meets the agreed specifications. Particular attention will be paid to the facilities links to ensure that the equipment is in a state ready for connection to the installed facilities at the proposed location in Tyndall. This inspection will also ensure that the machine as built corresponds exactly to the supplied drawings. Please confirm compliance. |  | Confirm (Yes / No) | Insert Comment |  |
| D16 | Equipment start-up is tenderer’s responsibility. Prior to start up, the supplier may inspect the facilities to determine whether they meet requirements. Start-up of the equipment shall indicate the tenderer’s acceptance of the facilities installation. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| D17 | Tyndall will carry out a Functional Acceptance Test. The equipment functionality test consists of the tenderer’s equipment start-up test. Prerequisites for the functional acceptance are: - Complete Delivery and installation - Resolution of all outstanding issues resulting from pre-shipment test, if applicable - Fulfilment of the following tests: - Shut-down test (ESO operation and re-start) Please confirm acceptance of this. |  | Confirm (Yes / No) | Insert Comment |  |
| D18 | The Tenderer MUST provide a qualified engineer to perform the SAT (Site Acceptance Test) post commissioning. The validation tests will be done in partnership with experienced Tyndall process engineers to ensure that the system performance has not been affected by shipping. These tests may also be used for on-site training purposes. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| D19 | The acceptance tests SHOULD include substrate handling tests in sequence without alarm conditions; the number of handling cycles to be mutually agreed between buyer and tenderer based on the system spec. - Recovery from power outage, extract outage, vacuum outage, nitrogen outage. - Delivery of all required documentation. - A copy of the tenderer’s pre-shipment final systems test procedure. - Delivery of supplier specified consumables. If the functional acceptance test produces negative results, Tyndall shall be entitled to request that the equipment will be brought into the guaranteed condition within 2 (two) weeks of the failed test. Please confirm compliance and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| D20 | Site Acceptance Test After functional acceptance, the Site Acceptance of the equipment shall be done jointly by engineers assigned by Tyndall and by the tenderer. A final acceptance document shall be prepared with the following content: - The purchase order reference. - The procurement specification reference. - A list of outstanding issues, including an action time frame for resolving these issues and the division of responsibility for this work. - The signatures and functions of the accepting engineers on each side, the Head of Process and Product Development group and the Tyndall Maintenance Manager. - The date of acceptance. - Serial number of the equipment. If the items in the list are satisfied, the equipment will be accepted. If one or more items are not accepted, the tenderer will be required to rectify the deficiency or, alternatively, a partial acceptance may be agreed combined with an extension of the warranty period. Terms shall be negotiated between Tyndall and tenderer based on the results of the acceptance tests. If there is a serious problem on the system, a mutually agreed plan of correction will be generated to correct the problem. Tyndall may waive any portion of the plan. The plan may include the replacement of the system at the discretion of the tenderer. Please confirm compliance. |  | Confirm (Yes / No) | Insert Comment |  |
| D21 | Tenderers are requested to provide a summary of scheduled maintenance and calibration required to keep instrument working to the full specification as per manufactures recommendations. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| D22 | A signed certificate for successful site acceptance test MUST be provided at the end of the installation and training. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| Documentation |  |  |  |  |  |
| D23 | A detailed User Operational Manual for the unit and control system MUST be provided. This SHOULD include the following; - Full electrical and design specifications. - Manuals MUST be supplied in English language. Manuals MUST be provided in soft copy. |  | Confirm (Yes / No) | Insert Comment |  |
| D24 | A Maintenance Manual for the system MUST be provided to include the following; - Cleaning requirements, service - Maintenance & Calibration requirements. - Manuals MUST be supplied in English language. - Manuals MUST be provided in softcopy where possible. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| D25 | A Technical Specification for the Unit MUST be provided covering all major components of the system. The specification MUST be supplied in the English language and in (if possible) in soft copy. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| D26 | A Technical Specification for all instrument documentation relating to installation of hardware and software, hardware/software interfacing, service and maintenance, parts list, and software operation and editing documentation MUST be supplied in English language. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| D27 | Factory Acceptance Test The equipment operation MUST be validated FAT (Factory Acceptance Test) prior to shipment to Tyndall to ensure it meets performance specifications. - The equipment will be shipped to Tyndall National Institute, Cork when Manufacturer validation tests have proved the equipment performs to the required specifications. Validation certificates MUST be provided during installation and commissioning. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| - Energy, Environmental & Sustainable Considerations |  |  |  |  |  |
| Environmental & Sustainable Considerations |  |  |  |  |  |
| E1 | Please outline your company’s commitment to the sustainable life cycle of the tendered equipment in a manner that minimises adverse environmental impacts. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| E2 | Energy & Environmental Requirements: The Contracting Authority requires the equipment provided to be as energy and economically efficient as possible. The system SHOULD have low running costs allowing the equipment to be as energy & economically efficient as possible. Where applicable, and if possible, the equipment SHOULD include the following: - Use of water-cooled units, e.g. water cooled chillers / pumps instead of air cooled chillers/ pumps as the site recovers heat from the processing cooling load to heat other areas. - Eco mode vacuum pumps to enable nitrogen saving when system in standby, control of both the nitrogen flow and eco mode activation SHOULD be from the system software - Where possible, use of a standby mode/reduced power mode of operation when the system is not in use. - Programmable capability, i.e. if possible equipment SHOULD come with the option of programmable timer so that processes can start overnight and the site can benefit from the reduced night-time electrical rates. Please provide any details on the above or other measures taken to improve energy and efficiency of the tool. Please provide details on any “warm-up” times required for the equipment to operate to specification. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| E3 | Tenderers SHOULD provide details of any takeback arrangements they may offer when the equipment reaches end of life and how that may be availed of. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| E4 | Tenderers SHOULD indicate Typical Energy Consumption during operational time while the equipment is in use. |  | Confirm (Yes / No) | Insert Comment |  |
| E5 | Where possible, use of a standby mode/reduced power mode of operation when the system is not in use. The tenderer SHOULD specify the stand-by energy consumption. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| E6 | It is desirable that all Tender requirement is delivered using environmentally friendly Packaging Materials; packaging, tape, labels and stretch wraps. Ideally, this packaging SHOULD avoid any single use plastic and be recovered or at a minimum reusable. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| E7 | Tenderers SHOULD provide details of expected lifetime of the equipment required. The Tenderer MUST state how long the model of equipment has been on the market and indicate the expected working life of the proposed models. Longevity of equipment etc. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| - Ultimate Cost. (Refer to Appendix 2: Pricing Schedule for Completion) |  |  |  |  |  |

# Lot 3 Plasma Etcher (Slow & Controllable Rate / Atomic Layer Etch of Silicon / Dielectrics)

| - Technical Merit |  |  |  |
| --- | --- | --- | --- |
| General / Functional Requirements |  |  |  |
|  | The Tenderer MUST supply a single-wafer processing tool for slow & controllable rate etching / atomic layer etching of Silicon / Silicon-based and Dielectric (e.g. e.g. silicon oxide aka SiO and silicon nitride aka SiN) materials. Please confirm and detail costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |
|  | UCC will consider tenders offering new instrumentation or refurbished and factory calibrated instrumentation. Tenderers MUST state whether new or ex demo/used systems is proposed. If the Instrument is ex demo it MUST be no more than three (3) years old and all parts and components MUST be available for at least the next ten (10) years. It MUST include a full refurbishment history with appropriate calibration certificates and full warranty. Please confirm and detail | Confirm (Yes / No) | Insert Comment |
|  | The tool MUST be capable of etching the following materials at slow & controllable rates (≤ 15 nm/min) as a minimum requirement: - Silicon / Silicon-based materials (e.g. e.g. silicon-on-insulator aka SOI, polysilicon etc.) and - Dielectric materials (e.g. SiO2 / SiN / HfO2 / Al2O3 etc.) Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The tool MUST also be capable of etching the following materials at higher rates ≥ 100 nm/min with the same configuration: - Silicon / Silicon-based materials (e.g. e.g. silicon-on-insulator aka SOI, polysilicon etc.) and - Dielectric materials (e.g. SiO2 / SiN / HfO2 / Al2O3 etc.) Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The tool MUST be capable of etching the materials listed in A3 and A4 above for a variety of micro and/or nanoscale structures with nanometre precision and excellent wafer-scale uniformity (Within-Wafer-Uniformity aka WiW) with the following parameters; - Photoresist-patterned structures with dimensions - min. ≥ 100 nm - Etch depths - min. ≥ 20 nm - max. ≤ 10 µm | Confirm (Yes / No) | Insert Comment |
|  | The tool MUST be suitable to run both flourine-based and chlorine-based chemistries. Any additional hardware (e.g. removable process-chamber liners etc.) required to allow for better process integrity and rapid switch-over between the different etch chemistries MUST be included as part of the core package. Please confirm and detail costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |
|  | The Tenderer MUST provide plasma cleaning & conditioning recipes and/or best-known method (BKM) procedures for conditioning the chamber when switching between fluorine-based and chlorine-based chemistries as part of the core package. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The tool MUST include a vacuum loadlock chamber as part of the core package. Tools with open-loading process-chambers will NOT be considered suitable for this tender. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The tool MUST be capable of etching silicon-based & dielectric materials on silicon and other substrates in wafer form (flatted or notched as-per SEMI M1-1117) for the following; - 100 mm diameter - 200 mm diameter Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The tool MUST be capable of handling the following substrate types in wafer form and for the sizes listed in A9 above as a minimum requirement: - Silicon (Si) - Silicon-on-Insulator (SOI) - Glass / Fused Silica - Silicon Carbide (SiC) - Germanium (Ge) Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The required etch performance of the tool MUST be achieved by inductively-coupled plasma (ICP) configuration with independent RF controls between the two sources. Bias control with high accuracy is considered essential for low-damage processes (e.g. polysilicon gate etch onto SiO2 etc.). Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The tool MUST be equipped with endpoint detection (EPD) hardware & software (e.g. Optical Emission Spectroscopy aka OES or laser interferometry etc.) as part of the core package. The proposed endpoint solution(s) MUST have the capability to monitor / control both etching and chamber cleaning & conditioning. Please confirm and detail costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |
| Performance Requirements |  |  |  |
|  | The tool MUST be capable of etching the materials listed in A3 and A4 above with excellent wafer-scale uniformity. Within-wafer variation (WiW) of the etch depth as measured at centre / edge points of wafer (with 10 mm edge-exclusion on both 100 mm and 200 mm diameter wafers) to be < ± 5%. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The tool MUST be able to use the following etch mask materials as a minimum requirement: - Photoresist (e.g. optical and electron-beam), - Dielectric hard-masks (e.g. silicon-dioxide, silicon-nitride, amorphous-silicon) It is expected that photoresist will be used as the mask material for most applications Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | Recipe # 7 (SOI) The tool MUST be capable of and a process recipe MUST be provided with the tool and demonstrated as part of the validation tests post at the Tyndall National Institute site for slow & controllable rate etching (≤ 15 nm/min) of < 100 µm linewidth microstructures in silicon-on-insulator (SOI) with soft-landing / minimal loss on silicon-dioxide (SiO2) under-layer on both 100 mm and 200 mm silicon wafers. The sidewall profile is to be anisotropic (≥ 75°). Tenderers MUST clearly demonstrate how they meet the requirements outlined with results documented in the attached Appendix C – Etch Recipe Results Template, this must be clearly named as “Lot 3 – Recipe #7-Results”. The Tenderer is permitted to use its own test mask / layout and substrates and / or own etching results to demonstrate compliance. Requirements: - Mask layout: open area < 10% (dark-field pattern) - Feature size & type: < 100 µm linewidth opening - Mask materials: Photoresist (PR) - Aspect ratio: > 1:1 - Depth: < 100 nm - Etched profile: anisotropic ≥ 75° and ≤ 90° or equivalent to photoresist sidewall profile - Mask undercut: < 100 nm - Etched sidewall: smooth - Etch rate Si: ≤ 15 nm/min - Etch rate SiO2: < 5 nm/min - Selectivity to Photoresist: > 1:1 (Si etch-rate ~ photoresist etch-rate) - Non-uniformity depth (100 mm wafer): < 5 (± %) - Non-uniformity depth (200 mm wafer): < 5 (± %) Tenderer is to specify the recommended process temperature and gases for this recipe Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | Recipe # 8 (SiO2) The tool MUST be capable of and a process recipe MUST be provided with the tool and demonstrated as part of the validation tests post at the Tyndall National Institute site for slow & controllable rate etching (≤ 15 nm/min) of < 100 µm linewidth contact / via microstructures in silicon-dioxide (SiO2) with soft-landing / minimal loss on silicon / polysilicon under-layer on both 100 mm and 200 mm silicon wafers. The sidewall profile is to be anisotropic (≥ 75°C) Tenderers MUST clearly demonstrate how they meet the requirements outlined with results documented in attached Appendix C – Etch Recipe Results Template, this must be clearly named as “Lot 3 – Recipe #8-Results”. The Tenderer is permitted to use its own test mask / layout and substrates and / or own etching results to demonstrate compliance. Requirements: - Mask layout: open area < 10% (dark-field pattern) - Feature size & type: < 100 µm linewidth contact / via - Mask materials: Photoresist (PR) - Aspect ratio: > 1:1 - Depth: < 100 nm - Etched profile: anisotropic ≥ 75° and ≤ 90° or equivalent to photoresist sidewall profile - Mask undercut: < 100 nm - Etched sidewall: smooth - Etch rate SiO2: < 15 nm/min - Etch rate Si / polysilicon: < 5 nm/min - Selectivity to Photoresist: > 1:1 (SiO2 etch-rate ~ photoresist etch-rate) - Non-uniformity depth (100 mm wafer): < 5 (± %) - Non-uniformity depth (200 mm wafer): < 5 (± %) Tenderer is to specify the recommended process temperature and gases for this recipe. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | Recipe # 9 (Silicon trench) The tool MUST be capable of and a process recipe MUST be provided with the tool and demonstrated as part of the validation tests post at the Tyndall National Institute site for higher rate etching (> 100 nm/min) of < 5 µm linewidth microstructures in silicon / SOI with good selectivity to silicon-dioxide (SiO2) hard-mask on 100 mm and 200 mm silicon wafers. The sidewall profile is to be highly anisotropic (≥ 85°C) Tenderers MUST clearly demonstrate how they meet the requirements outlined with results documented in in an attached Appendix, this must be clearly named as “Lot 3 – Recipe #9-Results”. The Tenderer is permitted to use its own test mask / layout and substrates and / or own etching results to demonstrate compliance. Requirements: - Mask layout: open area < 25% (dark-field pattern) - Feature size & type: 5 µm line / space - Mask materials: Photoresist (PR) - Aspect ratio: > 1:1 - Depth: ~ 1000 nm - Etched profile: anisotropic ≥ 85° and ≤ 90° or equivalent to photoresist sidewall profile - Mask undercut: < 100 nm - Etched sidewall: smooth - Etch rate Si: > 100 nm/min - Etch rate SiO2: < 20 nm/min - Selectivity to Photoresist: > 2:1 (Si etch-rate ~ 2x photoresist etch-rate) - Non-uniformity depth (100 mm wafer): < 5 (± %) - Non-uniformity depth (200 mm wafer): < 5 (± %) Tenderer is to specify the recommended process temperature and gases for this recipe. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The tool SHOULD have the capability to run the etch process recipes as listed in Recipes #7, #8 & #9 above at room temperature (20°C) and this will be considered an advantage and scored higher. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The tool SHOULD be capable of etching the following materials with the same configuration as for the recipes listed in for the recipes listed in Recipes #7, #8 & #9 above and will be considered an advantage and scored higher. - Silicon Nitride (SiN) - Aluminium Oxide (Al2O3) - Hafnium Oxide (HfO2) - Germanium (Ge) - Gallium Nitride (GaN) If any additional hardware (e.g. chamber-liners etc.) is required for etching these materials, the Tenderer MUST identify clearly in their response and cost separately. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |
|  | The is desirable that the tool SHOULD have capability to perform cyclical atomic-layer etching (ALE), removing as few as one monolayer of material per-cycle, and this will be considered an advantage and scored higher. Additional hardware which enables slow & controllable rate / atomic-layer etch processes (e.g. close-coupled MFCs, ALD valves, attenuated / pulsed RF generators), MUST be identified clearly by the Tenderer in their response. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation | Confirm (Yes / No) | Insert Comment |
|  | The Tenderer MUST provide plasma cleaning recipes for inter-wafer and inter-batch cleaning / polymer removal in the process-chamber as part of the core package. Please state if separate clean recipes for fluorine-based and chlorine-based chemistries are required and this MUST be clearly specified by the tenderer. Tenderer is to specify the recommended process temperature and gases for these cleaning recipes. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | It will be considered an advantage and scored higher if one or more of the gases required for process-chamber cleaning can also be used as an etch process gases (e.g. CF4 / O2 / SF6 etc.), however if other gases are recommended for more efficient cleaning recipes, then the tenderer SHOULD specify those gases as their preferred option. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | Recipe Support Tenderer MUST have ready access on request to a library of characterised etch / cleaning / conditioning recipes for the tool and assistance in the form of BKMs / recipes / technical support by phone & email when required in setting-up new recipes. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | Roadmap & Flexibility: Clarity of roadmap for the Tenderer’s proposed tool is required, indicating the potential future upgrade options that will be available as well as highlighting existing upgrade options for the system proposed. As with all pieces of equipment in UCC, the ability to upgrade in the future to add additional functionality not currently specified is very important, i.e., the tool SHOULD have a future roadmap of development. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The tenderer MUST state when the model was introduced and when the system is likely to be superseded. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| Hardware/Software Requirements |  |  |  |
|  | The system MUST be equipped with an aluminium process-chamber capable of being heated to > 50°C. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The process-chamber MUST be equipped with viewing port(s) for e.g. optical emission and laser reflection endpoint detection (EPD) systems at e.g. side and top chamber positions, and to directly view the plasma for stability. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The proposed endpoint solution(s) MUST be fully integrated with the system software. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | Tenderers SHOULD also supply further options of other EPD systems available. Please confirm and detail if any additional cost applies to additional EPD options that can be provided. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |
|  | Tenderers MUST identify clearly in their response if removable shields / liners for the process-chamber are required to allow for better process integrity / performance and to allow for rapid switch-over between fluorine-based and chlorine-based chemistries. Tenderers MUST include a minimum of two (2) sets of removable shields / liners as part of the core package if these are required for running the etch process recipes as listed in Recipes #7, #8 & #9 above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |
|  | The system MUST have a turbo pump to sustain low pressure (e.g. ≥ 1 mTorr) etch recipes / high flow-rate (e.g. > 100 sccm) cleaning recipes as part of the core package. Tenderers MUST specify the turbo pump brands and model numbers which will be provided with the tool as part of the core package. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | Tenders SHOULD also provide information on further options of other turbo pumps available for the process-chamber. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |
|  | The system MUST have a dry backing pump for the process-chamber. Tenderers MUST specify the pump brands and model numbers which will be provided with the tool as part of the core package. | Confirm (Yes / No) | Insert Comment |
|  | The dry backing pump for the process-chamber MUST be located remotely, positioned in a sub-fab with up-to a 10m vertical drop and up-to a 10m horizontal run to the point-of-use at the tool. Please see attached Appendix B – Sub-Fab Footprint for further details on the sub-fab footprint. The dry backing pump MUST have sufficient performance when working in-conjunction with the turbo pump to sustain low pressure (e.g. ≥ 1 mTorr) etch recipes as outlined in Recipes #7, #8 & #9 above and all necessary cleaning / conditioning recipes under these conditions. Sufficient length (minimum = 10 m) of electrical cable / plumbing etc. MUST be supplied to achieve this. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |
|  | Tenders SHOULD also provide information on further options of other dry backing pumps available for the process-chamber. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |
|  | The process-chamber MUST recover to a pressure of ≤ 1 mTorr and a leak-up rate ≤ 2 mTorr/min after a wet-clean in less than two (2) hours. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The system MUST have a dry backing pump for the loadlock. Tenderers MUST specify the pump brands and model numbers which will be provided with the tool as part of the core package. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The dry backing pump for the loadlock MUST be located remotely, positioned in a sub-fab with up-to a 10 m vertical drop and up-to a 10 m horizontal run to the point-of-use at the tool. Please see attached Appendix B – Sub-Fab Footprint for further details on the sub-fab footprint. The dry backing pump MUST have sufficient performance to sustain low pressure (e.g. ≤ 100 mTorr) for wafer transfer into the process-chamber under these conditions. Sufficient length (minimum = 10 m) of electrical cable / plumbing etc. MUST be supplied to achieve this. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |
|  | Tenderers SHOULD also provide information on further options of other dry backing pumps available for the loadlock. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation | Confirm (Yes / No) | Insert Comment |
|  | Loadlock pump-down time SHOULD be less than 120 seconds, and no longer than 300 seconds. Faster pump-down time will be considered an advantage and scored higher. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The system MUST have a heated foreline to remove etch and cleaning plasma by-products from the process-chamber. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The system MUST have an automatic throttle valve for pressure control in the process-chamber. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The system MUST be water-cooled. Process cooling-water supply is provided by Tyndall, see section D. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The process-chamber platen / wafer-electrode MUST be heater-chiller controlled to allow operation in the range 0°C and +50°C with control of ±2°C from set point. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The heater-chiller unit MUST be located remotely, positioned in a sub-fab with up-to a 10 m vertical drop and up-to a 10 m horizontal run to the point-of-use at the tool. Please attached Appendix B – Sub-Fab Footprint for further details on the sub-fab footprint. The heater-chiller unit MUST have sufficient performance to sustain platen / wafer-electrode temperature in the range 0°C and +50°C with control of ±2°C under these conditions. Sufficient length (minimum = 10 m) of electrical cable / plumbing etc. MUST be supplied to achieve this. Please confirm and detail costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |
|  | A non-return chiller valve MUST be included to prevent back-flow to reservoir. Tenderer MUST specify what model heater/chiller will be included and the coolant type (e.g. Galden or equivalent) recommended for the heater-chiller units. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | System MUST be equipped with a fully auto-matched RF generator for each plasma source. The RF generators MUST have capability for plasma strike and sustain tuned plasmas operating over a wide processing pressure range ≥ 1 mTorr. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | Tenderers MUST specify details of the RF generator brand & model, frequency, power and operating range for each plasma source which will be provided as part of the core package. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The RF generators / racks may be located remotely, positioned in a sub-fab with up-to a 10 m vertical drop and up-to a 10 m horizontal run to the point-of-use at the tool. Please see attached Appendix B – Sub-Fab Footprint for further details on the sub-fab footprint. Sufficient length (minimum = 10 m) of electrical cable / plumbing etc. MUST be supplied to achieve this. The RF generators / racks MUST have sufficient performance for plasma strike and sustain tuned plasmas under these conditions. Please confirm and detail costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |
|  | RF Generator racks MUST be supplied by the Tenderer, if required for the assembly of the system. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | Tenderers SHOULD also provide information on further options of other RF generators available. The tool SHOULD have the capability for Pulsed RF / Fast-Strike plasma-matching for slow & controllable rate / atomic layer etching and this will be considered an advantage and scored higher. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. If any options to purchase are available, these are optional to purchase by the Contracting Authority and will not form part of the Cost evaluation. | Confirm (Yes / No) | Insert Comment |
|  | The tool MUST be configured for single-wafer processing (wafers are etched individually). The hardware changes / procedures required to switch from running one wafer-diameter to another MUST be outlined clearly by the tenderer with all necessary hardware provided as part of the core package. The ability to switch from e.g. 100 mm to 200 mm substrate diameter with minimal downtime and / or hardware changes will be considered an advantage and scored higher. Please confirm and detail changeover procedure and time to complete. | Confirm (Yes / No) | Insert Comment |
|  | The tool MUST have capability to handle wafer thicknesses in the range ≥ 0.2 mm and ≤ 1 mm as a minimum requirement. An ability to also handle substrates > 1 mm thickness with the same setup will be considered an advantage and scored higher. Please specify minimum and maximum thickness in tender response and identify any hardware changes (e.g. separate wafer-carriers for different thickness) required to allow a range of thin / thick substrates to be processed. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |
|  | The tool SHOULD have capability to handle bonded-wafer pairs, perforated wafers, or bowed wafers (with bow / warp > 50 µm). This will be considered an advantage and scored higher. Please identify any hardware changes (e.g. separate wafer-carriers etc.) required to allow for handling these non-standard types of substrates. If any additional hardware (e.g. wafer-carriers etc.) is required for handling these non-standard types of substrates, the Tenderer MUST identify clearly in their response and cost separately. Tenderers are to provide details and pricing for the proposed option within the core package in the Appendix 2, Pricing Schedule. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |
|  | The tool MUST accommodate wafers as set out in requirements A9 – A10 above. Helium gas heat-transfer for wafer back-side cooling with electrostatic or mechanical clamping MUST be included. Helium transfer MUST be pressure-controlled and the helium flow rate SHOULD automatically be checked prior to wafer etch. The He pressure MUST be stable at e.g. 10 Torr / 15 Torr during the clamping of a wafer and the He flow leakage may not affect the process's results. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The system MUST have a minimum eight (8) gas lines. The capability to have more than eight gas lines will be considered an advantage and scored higher. Suggested process gases include, but are not limited to the following; SF6, C4F8, O2, BCl3, Ar, Cl2 The tenderer MUST specify their recommended gases for the recipes outlined in requirements Recipes #7, #8 & #9 above and for plasma cleaning & conditioning of the process-chamber. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | In the specific case(s) of low vapour pressure / low boiling point gases (e.g. BCl3) which are susceptible to condensation, heated gas-delivery hardware such as gas-lines, gas-box, MFCs etc. MUST be included as part of the core package. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |
|  | In the specific case of oxygen (O2) the requirements for e.g. low flow-rate for high-selectivity etching of silicon onto silicon dioxide and high flow-rate for chamber plasma cleaning / chamber plasma conditioning may necessitate two separate MFCs, so this should be clearly specified by the Tenderer. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The tenderer MUST specify recommended flow size for each gas Mass Flow Controller (MFC). MFCs supplied MUST be rated (in terms of e.g. seals etc.) and calibrated for the recommended gases as nominated by the Tenderers under requirement A55. Digital MFCs are preferred and will be considered an advantage and scored higher. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The capability to run automatic interleaved and / or wafer-less clean MUST be available in the software. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | Whole system MUST be fully interlocked to protect against any equipment or service failure. System MUST be designed to fail-safe with system left under vacuum. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | Electrical equipment racks may need to be located remotely, positioned in a sub-fab with up-to a 10 m vertical drop and up-to a 10 m horizontal run to the point-of-use at the tool and MUST be capable of full functionality under these conditions. Sufficient length (minimum = 10 m) of electrical cable MUST be supplied to allow the required performance. Please see attached Appendix B – Sub-Fab Footprint for further details on the sub-fab footprint. Confirmation is required that the operation of the tool and etch quality will not be impacted by remote location of the electronic control cabinets. Please confirm and detail costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |
|  | Electrical equipment racks MUST be supplied by the Tenderer, if required for the assembly of the system. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |
|  | A minimum of one (1) computer and console MUST be included in core package, and MUST use Windows-based OS, Windows 10 or higher. Please describe the software package proposed, including name and version. Please also specify the projected lifespan of the current control system architecture. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |
|  | The tool will be installed in the cleanroom in a thru-wall configuration. ALL computer(s) / console(s) required for tool operation from the front (process bay) and rear (service chase) side MUST be included in the core package. Additional console(s) / computer(s) for tool operation from the rear (service chase) side MUST also be included in the core package, or a suitable alternative such as a rotating screen(s) that are visible from both bay and chase with associated keyboard / mouse control from both sides. Please see attached Appendix A -Cleanroom Footprint for further details on the cleanroom footprint. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No | Insert Comment |
|  | The tenderer MUST provide all control and analysis software required for tool operations. All tool operations, including wafer transfer to / from the loadlock to the process-chamber MUST be software-controlled and allow for both auto & manual operation. Please confirm and detail costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |
|  | The Plasma Etcher equipment SHOULD have perpetual licence-free software or no restriction on the number of licences needed for our use of the system i.e., licences for multiple users and computer installations for off-site use if this is possible. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | All software updates to the system, which become available during the life of the contract, MUST be provided free of charge during the warranty period and for a minimum period of eight (8) years after warranty expiry. Please confirm and detail costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |
|  | Continued development of the software SHOULD be demonstrated by the Tenderer, with facilities for updating systems in use. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The system MUST be capable of remote access via network connection (using e.g. TeamViewer software or equivalent) for tool support in event of problems. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The software MUST be capable of data-logging of process parameters, and archiving data-logs. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | It MUST be easy to graphically display and export saved etch process parameter data preferably into Excel spreadsheets. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The system MUST have an easy-to-use interface and needs to be suitable for a multi-user (i.e. different access levels) environment. Various levels of access are required e.g. Operator, Supervisor, Maintenance. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The software SHOULD have the facility to lock certain recipes i.e. not editable by users other than Process Engineer / Supervisor / Maintenance etc. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | Software MUST include backup and restore capability (export/import) for process recipes and machine settings. Software SHOULD include backup (image) capability to be used in case of hard disk drive (HDD) failure. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | Software MUST allow operator access to data files while the system is running under standard operational conditions. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| - Quality, Health & Safety & Validation |  |  |  |
| Quality Requirements |  |  |  |
|  | Compliance with all relevant European health and safety standards, quality assurance and protocols are MANDATORY. All Goods and Services supplied in accordance with the contract MUST comply with all necessary standards, i.e. Irish Standards Institution, European Union, CE, International Standards Organisation or other applicable standard. The Tenderer MUST, on the request of the buyer, provide a copy of the standard free of charge. Please confirm & detail. | Confirm (Yes / No) | Insert Comment |
|  | If there are manufacturing defects known to the tenderer, they MUST identify such defects and state their policy regarding the repair of known defects. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The Tenderer SHOULD, as part of the submission, include copies of any relevant testing or compliance certificates for the Goods. (For example, Factory acceptance tests, Quality assurance test etc.) Please confirm and detail | Confirm (Yes / No) | Insert Comment |
| Health & Safety Requirements |  |  |  |
|  | All goods supplied in accordance with the contract MUST be of good construction, sound material, of adequate strength, free of design defects and manufactured to be safe and without risk when properly used and software MUST allow operator access to data files while the system is running under standard operational conditions. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The tool SHOULD meet the criteria of the Semi S2-0821Eb Environmental, Health, and Safety Guideline for Semiconductor Manufacturing Equipment in order to ensure easier and safer installation and maintenance from a gas safety perspective. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The Tenderer MUST, as part of the submission, include Laser Safety Warning labels and copies of any relevant testing or compliance certificates for the Goods. Confirm compliance and detail. | Confirm (Yes / No) | Insert Comment |
|  | The tools MUST knowingly not present any Health and Safety hazards. Any safety & hazard warning labelling on the proposed system MUST be in the English Language. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | Tenderers MUST advise if there are any hazards associated with the operation or maintenance of this instrument? (e.g. manual handing, chemicals, etc.). Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| - Warranty/Spare Parts, After-Sales & Training Support |  |  |  |
| Warranty & Customer Support, Repair/Breakdown Requirements |  |  |  |
|  | Warranty Duration: At least one (1) year’s warranty from Site Acceptance Test (SAT) MUST be included for each lot as part of the core package and the quoted price is to include all hardware, software, spare parts, labour, call outs, visit, travel and parts. Warranties of greater than 1 year are considered highly desirable and will be scored higher. Detail specific inclusions and exclusions of warranty. Preventative Maintenance MUST be provided for at least one (1) year and a list of checks carried out at Preventative Maintenance MUST be provided, along with any information about which (if any) parts are routinely replaced. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |
|  | If UCC decide to extend the warranty period, Tenderers are asked to outline costs of any additional warranty. The Tenderer MUST state the details of any extended warranty and details of Preventative Maintenance and calibration & servicing check and any associated costs. Detail specific inclusions and exclusions of extended warranty and a list of checks carried out at Preventative Maintenance SHOULD be provided, along with any information about which (if any) parts are routinely replaced. Please include price per annum, details of coverage and maximum term of coverage. Please include price per annum for year 2, 3, & 4, details of coverage and maximum term of coverage. Please confirm and detail and specify what is included in the warranty options outlined in the Appendix 2 pricing schedule, Optional to Purchase. Detail specific inclusions and exclusions of extended warranty. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |
|  | Please provide separately a list and approximate cost of any consumable items/parts that are to be replaced during each service visit. Please update Appendix 2 Pricing Schedule – Consumables Parts sheet with detail. Prices provided for consumable items/parts MUST be valid for a minimum of one (1) year. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |
|  | Please advise of any items/part of the instrument that require routine replacement during the normal lifetime of the instrument, the timeframe in which they require replacing and if such items will be supplied with the instrument or as part of the warranty. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The tenderer MUST provide a description of the arrangements the arrangements and lead times for guaranteed responses times in the event of equipment breakdown both in relation to phone support, online/remote support, the provision of technical support, annual service/ maintenance, emergency call-out response, and spare parts. Please detail lead time for arrival on site by a fully trained service engineer during the warranty period(s) and post-warranty. This response time SHOULD be no more than 24 hours via phone call or remote service, and within seven (7) working days for onsite repair. Please confirm and detail information on remote diagnostic / telecom-based service support. | Confirm (Yes / No) | Insert Comment |
|  | Outside of Warranty Period: The Tenderer MUST provide typical prices and lead time for servicing and repairs (including travel, subsistence and price per day) charged outside a warranty/ maintenance contract. Detail specific inclusions and exclusions outside of warranty period. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |
|  | If UCC decide to take out an additional Preventative Maintenance Agreement / SLA, tenderers are asked to outline Options for years 2, 3 & 4. Option to include: - On-going Calibration to appropriate schedule - preventative maintenance - remote diagnostics of faults, if applicable - repair of the Instrument in case of breakdown) Tenderers are asked to outline full details of what is included and to outline Cost of maintenance and calibration for each year option offered. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |
|  | The tenderer MUST have dedicated fully trained and qualified service engineering specialists at their disposal who can respond to all service calls and any form of breakdown call. This remote response time SHOULD be the same day via phone call or remote service, and within seven (7) working days for onsite repair. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | Please outline the company’s Customer Support Proposal and escalation policy. Tenderers MUST describe what process is in place for escalating a call-out, if a breakdown cannot be resolved on the first visit. Tenderer SHOULD provide a list of approved service and calibration agents. Please confirm and detail customer support arrangements, service and calibration agents. Additional support and service options should also be outlined and pricing noted within the Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |
|  | The Tenderer MUST indicate the arrangements for and specify the level of experience of applications specialists and engineers who will be in charge of after sales service. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The Tenderer SHOULD specify the lead time for repairs during the following scenarios. - warranty period(s) - post-warranty, - with service contract - without service contract Please confirm and detail. | Confirm (Yes / No)) | Insert Comment |
|  | Please note SHOULD installation, commissioning, service visits or availability of spare parts be affected in the event of a Pandemic crisis it is expected that the warranty period will be extended to facilitate the service downtime endured by the Contracting Authority. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | Warranty commences post full installation and sign-off of tool acceptance by nominated UCC representative. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| Replacement & Spare Parts & Equipment Consumables |  |  |  |
|  | Spare Parts: Availability of all spare parts SHOULD be maintained 10 years after the warranty period. Please indicate the availability and lead time of spare part stock centres, with maximum lead-time for any part. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | Critical spares MUST be held at a location no more than 7 Days delivery time from Tyndall. Spare parts to be held offsite with the tenderer and essential parts if necessary to be held Free of Charge (FOC) at Tyndall. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | Consumable Parts: Sufficient consumable parts MUST be supplied for the initial installation, commissioning and training. These consumable parts MUST be provided free of charge and at the point of commissioning. Full one (1) year of consumable parts to be supplied post commissioning, including any removable shields / liners for process-chamber. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |
|  | Recommended Spare Lists: The tenderer MUST supply a full list of recommended spare parts and consumables for the proposed equipment, based on knowledge of equipment design and component lifetimes. This SHOULD include typical consumables and high turnover parts. Please detail in Appendix 2, Pricing Schedule- Consumable Parts. Prices provided for consumable items/parts MUST be valid for a minimum of one (1) year. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |
|  | The tenderer SHOULD supply, where information is available, to the buyer a schedule of spare parts showing in tabular form or, otherwise, the following information on OEM parts: - Names and address of all OEM suppliers - Part Name - Part Description - OEM Supplier Part Number. Please confirm and detail, attaching information requested above. | Confirm (Yes / No) | Insert Comment |
|  | The tenderer MUST provide confirmation of the availability of service for at least five (5) years after commissioning. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| Training & Technical Support Proposal |  |  |  |
|  | The successful tenderer MUST provide an initial training programme for UCC personnel on-site at Tyndall, for a minimum of one (1) day each for operations and maintenance, with a minimum of three (3) users fully trained to safely operate and maintain the system as part of the core package. The training programme MUST be carried out by a suitably qualified specialist at the tool after the initial installation, commissioning is completed. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |
|  | Initial training provided MUST include, but not be limited to: - Demonstration of the correct operation of all relevant equipment of the systems - Routine Maintenance requirements - Basic Troubleshooting - Data processing and management - Software and hardware familiarisation Dates, timing and method of training are to be confirmed between UCC and the winning tenderer. The process validation tests may also be used for on-site training purposes. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | The tenderer MUST specifically outline the level and duration of application support offered with the instrument where “application support” refers to the expert technical support (on-line or on-site) available from the manufacturer for using the instrument in novel research applications. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
|  | This equipment will have multiple users over the lifetime of the tool. Tenderers MUST provide details and pricing for their proposed option for Operations / Applications / Maintenance Training at all levels (e.g. L1, L2, L3 etc.) on-site or at their facilities. Training offered over the full lifetime of the equipment will be considered an advantage and scored higher. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. | Confirm (Yes / No) | Insert Comment |
|  | Relevant training and maintenance manuals in English MUST also be provided with the instrument. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| - Delivery Plan |  |  |  |
| Lead Time & Delivery |  |  |  |
| D1 | The equipment SHOULD be delivered within twelve (12) months of receipt of Purchase Order. The equipment SHOULD be delivered within the stated lead time requirement from receipt of Purchase Order. (Successful tenderer for each lot may be required to liaise with each other for coordination of delivery and installation of the equipment) - Tenderers MUST provide full details of lead times for the delivery of equipment including installation and commissioning (if applicable), from receipt of the purchase order. - Tenderers MUST provide the facility requirements to UCC a minimum of 16 weeks in advance of the equipment delivery. - Tenderers MUST provide advance notice (two weeks) of any requirements for tool unloading, number of containers and weight of each. Height, length and width of containers/crates to be provided as well. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| D2 | Delivery MUST include all shipping charges, insurance, customs clearance and all aspects as outlined in DAP (Delivery At Place) incoterms. Delivery MUST be to Tyndall National Institute, UCC: - Silicon Fabrication Laboratory (Room A1.1.04), - Tyndall National Institute, - University College Cork, - Lee Maltings, Dyke Parade, - Cork T12 R5CP, - Ireland. Advance notice MUST be given (2 weeks) of any requirements for tool unloading, number of containers and weight of each. Height, length and width of containers/crates MUST be provided as well. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |
| D3 | The tenderer MUST provide full details of lead times for the delivery of equipment including installation and commissioning (if applicable), from receipt of the purchase order. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| D4 | The tenderer MUST provide details of the logistics/supply chain to the point of customer delivery, i.e. where the goods will come from, how the goods will be transported and what borders the goods will cross. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| D5 | The tenderer MUST provide details of any actions that have been taken to ensure the future timely supply and delivery of goods and services, including ongoing supply of replacement/spare parts and consumables. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| D6 | The tenderer MUST provide details of their contingency plans (alternative sources/routes of supply) for any delays in delivery should they arise. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| Installation, Commissioning, Calibration & Validation Requirements |  |  |  |
| D7 | Pre-shipment test SHOULD demonstrate the system’s capability to achieve all process parameters outlined above. Pre-Shipment test results should be shared as agreed between Tyndall and the Tenderer, prior to shipment. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| D8 | The tenderer MUST commission the system and verify full functionality at the Tyndall National Institute site. The buyer's engineering staff (both process and equipment) will provide reasonably requested assistance to the supplier in this task. Tyndall will also provide for any labour, chemicals, wafers, power, lubricants, water, fuel and materials reasonably requested for the purpose of the process installation. Tenderer has to make a written request for this support at least four weeks in advance of installation. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| D9 | The facility requirements for this piece of equipment MUST be supplied to Tyndall a minimum of sixteen (16) weeks in advance of the equipment delivery. These facilities changes will be installed to the specifications supplied to Tyndall by the tenderer. Any deviation to the supplied specification on the delivered equipment causing a reconfiguration of the facility supply will be at the tenderer’s cost. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| D10 | The proposed system MUST be suitable for use with the 400 V, 50 Hz, 3-phase electrical supply at the proposed installation site. If a transformer unit is required to achieve this, it MUST be supplied as part of the core package. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| D11 | Tenderers SHOULD also note that the following facilities will be available at the installation site: - 3-phase electrical supply as per D10 above - Process cooling water supply - De-ionised water supply - Waste water drainage - Compressed air supply - Air exhaust. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| D12 | Tyndall Cleanroom The tenderer MUST provide footprint drawings for tool and ancillary equipment (e.g. control consoles etc.) required for operating the tool which MUST be located in the cleanroom. The available cleanroom space will be limited so tools and ancillary equipment MUST fit within a space approx. 5 m2 per tool (total footprint for the suite of tools being tendered-for is approx. 20 m2. Vacuum pumps / chillers / equipment racks etc. will be located remotely, positioned in a sub-fab with up-to a 10m vertical drop and up-to a 10m horizontal run to the point-of-use at the tool. This equipment MUST fit within a space approx. 4 m2 per tool (total footprint for the suite of tools being tendered-for is approx. 16 m2 Please see guideline footprint drawings of the cleanroom and sub-fab locations where the suite of tools will be installed in attached Appendices; Appendix A- Cleanroom Footprint & Appendix B – Sub-Fab Footprint. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| D13 | The tools MUST have thru-wall configuration for installation in the Silicon Fabrication cleanroom. All hardware (e.g. front panel, keyboard(s), monitor(s), etc.) required for thru-wall configuration MUST be included in the core package, as outlined in requirement A63 above. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| D14 | The installation, commissioning, and calibration of the equipment MUST be carried out on-site by a qualified engineer provided by the Tenderer. This entails the system being transported to the lab, unpacked, set up and demonstrated to meet the requirements as outlined in section A. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. | Confirm (Yes / No) | Insert Comment |
| D15 | Prior to installation, the machine may be inspected at the Tenderer site by a Tyndall engineer to ensure that it meets the agreed specifications. Particular attention will be paid to the facilities links to ensure that the equipment is in a state ready for connection to the installed facilities at the proposed location in Tyndall. This inspection will also ensure that the machine as built corresponds exactly to the supplied drawings. Please confirm compliance. | Confirm (Yes / No) | Insert Comment |
| D16 | Equipment start-up is tenderer’s responsibility. Prior to start up, the supplier may inspect the facilities to determine whether they meet requirements. Start-up of the equipment shall indicate the tenderer’s acceptance of the facilities installation. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| D17 | Tyndall will carry out a Functional Acceptance Test. The equipment functionality test consists of the tenderer’s equipment start-up test. Prerequisites for the functional acceptance are: - Complete Delivery and installation - Resolution of all outstanding issues resulting from pre-shipment test, if applicable - Fulfilment of the following tests: - Shut-down test (ESO operation and re-start) Please confirm acceptance of this. | Confirm (Yes / No) | Insert Comment |
| D18 | The Tenderer MUST provide a qualified engineer to perform the SAT (Site Acceptance Test) post commissioning. The validation tests will be done in partnership with experienced Tyndall process engineers to ensure that the system performance has not been affected by shipping. These tests may also be used for on-site training purposes. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| D19 | The acceptance tests SHOULD include substrate handling tests in sequence without alarm conditions; the number of handling cycles to be mutually agreed between buyer and tenderer based on the system spec. - Recovery from power outage, extract outage, vacuum outage, nitrogen outage. - Delivery of all required documentation. - A copy of the tenderer’s pre-shipment final systems test procedure. - Delivery of supplier specified consumables. If the functional acceptance test produces negative results, Tyndall shall be entitled to request that the equipment will be brought into the guaranteed condition within 2 (two) weeks of the failed test. Please confirm compliance and detail. | Confirm (Yes / No) | Insert Comment |
| D20 | Site Acceptance Test After functional acceptance, the Site Acceptance of the equipment shall be done jointly by engineers assigned by Tyndall and by the tenderer. A final acceptance document shall be prepared with the following content: - The purchase order reference. - The procurement specification reference. - A list of outstanding issues, including an action time frame for resolving these issues and the division of responsibility for this work. - The signatures and functions of the accepting engineers on each side, the Head of Process and Product Development group and the Tyndall Maintenance Manager. - The date of acceptance. - Serial number of the equipment. If the items in the list are satisfied, the equipment will be accepted. If one or more items are not accepted, the tenderer will be required to rectify the deficiency or, alternatively, a partial acceptance may be agreed combined with an extension of the warranty period. Terms shall be negotiated between Tyndall and tenderer based on the results of the acceptance tests. If there is a serious problem on the system, a mutually agreed plan of correction will be generated to correct the problem. Tyndall may waive any portion of the plan. The plan may include the replacement of the system at the discretion of the tenderer. Please confirm compliance. | Confirm (Yes / No) | Insert Comment |
| D21 | Tenderers are requested to provide a summary of scheduled maintenance and calibration required to keep instrument working to the full specification as per manufactures recommendations. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| D22 | A signed certificate for successful site acceptance test MUST be provided at the end of the installation and training. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| Documentation |  |  |  |
| D23 | A detailed User Operational Manual for the unit and control system MUST be provided. This SHOULD include the following; - Full electrical and design specifications. - Manuals MUST be supplied in English language. Manuals MUST be provided in soft copy. | Confirm (Yes / No) | Insert Comment |
| D24 | A Maintenance Manual for the system MUST be provided to include the following; - Cleaning requirements, service - Maintenance & Calibration requirements. - Manuals MUST be supplied in English language. - Manuals MUST be provided in softcopy where possible. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| D25 | A Technical Specification for the Unit MUST be provided covering all major components of the system. The specification MUST be supplied in the English language and in (if possible) in soft copy. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| D26 | A Technical Specification for all instrument documentation relating to installation of hardware and software, hardware/software interfacing, service and maintenance, parts list, and software operation and editing documentation MUST be supplied in English language. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| D27 | Factory Acceptance Test The equipment operation MUST be validated FAT (Factory Acceptance Test) prior to shipment to Tyndall to ensure it meets performance specifications. - The equipment will be shipped to Tyndall National Institute, Cork when Manufacturer validation tests have proved the equipment performs to the required specifications. Validation certificates MUST be provided during installation and commissioning. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| - Energy, Environmental & Sustainable Considerations |  |  |  |
| Environmental & Sustainable Considerations |  |  |  |
| E1 | Please outline your company’s commitment to the sustainable life cycle of the tendered equipment in a manner that minimises adverse environmental impacts. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| E2 | Energy & Environmental Requirements: The Contracting Authority requires the equipment provided to be as energy and economically efficient as possible. The system SHOULD have low running costs allowing the equipment to be as energy & economically efficient as possible. Where applicable, and if possible, the equipment SHOULD include the following: - Use of water-cooled units, e.g. water cooled chillers / pumps instead of air cooled chillers/ pumps as the site recovers heat from the processing cooling load to heat other areas. - Eco mode vacuum pumps to enable nitrogen saving when system in standby, control of both the nitrogen flow and eco mode activation SHOULD be from the system software - Where possible, use of a standby mode/reduced power mode of operation when the system is not in use. - Programmable capability, i.e. if possible, equipment SHOULD come with the option of programmable timer so that processes can start overnight and the site can benefit from the reduced night-time electrical rates. Please provide any details on the above or other measures taken to improve energy and efficiency of the tool. Please provide details on any “warm-up” times required for the equipment to operate to specification. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| E3 | Tenderers SHOULD provide details of any takeback arrangements they may offer when the equipment reaches end of life and how that may be availed of. | Confirm (Yes / No) | Insert Comment |
| E4 | Tenderers SHOULD indicate Typical Energy Consumption during operational time while the equipment is in use. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| E5 | Where possible, use of a standby mode/reduced power mode of operation when the system is not in use. The tenderer SHOULD specify the stand-by energy consumption. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| E6 | It is desirable that all Tender requirement is delivered using environmentally friendly Packaging Materials; packaging, tape, labels and stretch wraps. Ideally, this packaging SHOULD avoid any single use plastic and be recovered or at a minimum reusable. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| E7 | Tenderers SHOULD provide details of expected lifetime of the equipment required. The Tenderer MUST state how long the model of equipment has been on the market and indicate the expected working life of the proposed models. Longevity of equipment etc. Please confirm and detail. | Confirm (Yes / No) | Insert Comment |
| - Ultimate Cost. (Refer to Appendix 2: Pricing Schedule for Completion) |  |  |  |

# Lot 4 Plasma Deposition (PECVD)

| - Technical Merit |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| General/Functional Requirements |  |  |  |  |  |  |
|  | The Tenderer MUST supply a single-wafer processing tool for Plasma-Enhanced Chemical Vapour Deposition (PECVD) tool of thin films of Dielectric materials (e.g. silicon oxide aka SiO and silicon nitride aka SiN) and silicon-based (e.g. amorphous silicon aka a-Si) materials. Please confirm and detail costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | UCC will consider tenders offering new instrumentation or refurbished and factory calibrated instrumentation. Tenderers MUST state whether new or ex demo/used systems is proposed. If the Instrument is ex demo it MUST be no more than three (3) years old and all parts and components MUST be available for at least the next ten (10) years. It MUST include a full refurbishment history with appropriate calibration certificates and full warranty. Please confirm and detail |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The tool MUST be capable of depositing thin films of the materials listed in A5 below with excellent wafer-scale uniformity. For deposition of thin films wafer-scale uniformity (Within-Wafer-Uniformity aka WiW) is defined as: Max-MinMean x 100 (%) for a 5-point measurement (N, S, E, W, C) across the full wafer diameter on 100 mm / 200 mm wafers with a 5 mm edge-exclusion. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The tool MUST be capable of depositing thin films of the materials listed in A5 below with excellent stress control. For deposition of thin films, stress is measured using contact profilometry (pre / post deposition) and radial scanning (30° interval) across the full wafer diameter on 100 mm / 200 mm wafers with a 10 mm edge-exclusion. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The tool MUST be capable of depositing thin films of - Silicon-based dielectric materials (e.g. SiO / SiN) and - Silicon-based amorphous materials (e.g. a-Si) at a process temperature of 300°C with nanometre precision and wafer-scale uniformity (WiW) < ± 5%. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The tool MUST be capable of depositing thin films of the materials listed in A5 above with the following parameters; - Deposition Rate (300°C) - SiO min. ≥ 300 nm/min - SiN min. ≥ 100 nm/min - a-Si min. ≥ 50 nm/min - Stress (300°C) - all films ± 200 MPa - WiW (300°C) - all films < ± 5% Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The tool MUST be capable of depositing thin films of - Silicon-based dielectric materials (e.g. SiO / SiN) and - Silicon-based amorphous materials (e.g. a-Si) at a process temperature of ≤ 200°C. with nanometre precision and wafer-scale uniformity (WiW) < ± 5% Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The tool MUST be capable of depositing thin films of the materials listed in A7 above with the following parameters; - Deposition Rate (≤ 200°C) - SiO min. ≥ 20 nm/min - SiN min. ≥ 20 nm/min - a-Si min. ≥ 20 nm/min - Stress (≤ 200°C) - all films ± 200 MPa - WiW (≤ 200°C) - all films < ± 5% Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The Tenderer MUST provide plasma cleaning & conditioning recipes and best-known method (BKM) procedures for conditioning the process chamber after wet-cleaning and when switching between deposition of different films as part of the core package. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The tool MUST include a vacuum loadlock chamber as part of the core package. Tools with open-loading process-chambers will NOT be considered suitable for this tender. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The tool MUST be configured for automated loading of individual substrates / carriers from a cassette to and from the process-chamber. Operators will load wafers / carriers into a cassette (e.g. separate 25-slot cassettes for both 100 mm and 200 mm diameter wafers or 12/13-slot double-spaced cassettes for both 200 mm diameter wafers and 200 mm diameter wafer-carriers etc.) and the material type for the cassettes is to be specified by tenderer). The wafers / carriers MUST be transferred to & from the process-chamber via robotic-arm. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The tool MUST be capable of depositing dielectric & silicon-based materials on silicon and other substrates in wafer form (flatted or notched as-per SEMI M1-1117) for the following; - 100 mm diameter - 200 mm diameter Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The tool MUST be capable of handling the following substrate types in wafer form and for the sizes listed in A12 above as a minimum requirement: - Silicon (Si) - Silicon-on-Insulator (SOI) - Glass / Fused Silica - Silicon Carbide (SiC) - Germanium (Ge) Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The tool MUST be equipped with endpoint detection (EPD) hardware & software (e.g. Optical Emission Spectroscopy aka OES / reflectometry etc.) as part of the core package. The proposed endpoint solution(s) MUST have the capability to monitor / control process-chamber plasma cleaning & conditioning. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  |  | Insert Comment |
| Performance Requirements |  |  |  |  |  |  |
|  | Recipe # 10 (Sil-Ox @ 300°C) The tool MUST be capable of and a process recipe MUST be provided with the tool and demonstrated as part of the validation tests at the Tyndall National Institute site for deposition of silane-based oxide (Sil-Ox) films on both 100 mm and 200 mm silicon wafers at a deposition temperature of 300°C ± 20°C. Tenderers MUST clearly demonstrate how they meet the requirements outlined with results documented in the attached Appendix D – Deposition Recipe Results Template, this must be clearly named as “Lot 4 – Recipe # 10 -Results”. The Tenderer is permitted to use its own substrates and / or own deposition results to demonstrate compliance. Requirements: Dep. Rate Sil-Ox: > 300 nm/min Stress range Sil-Ox: < ± 200 MPa Refractive Index: 1.45 - 1.47 is desirable - Non-uniformity film thickness (WiW) on 100 mm wafer with 5 mm edge exclusion: < ± 5% - Non-uniformity film thickness (WiW) on 200 mm wafer with 5 mm edge-exclusion: < ± 5% Tenderer to specify process temperature and recommended gases for their recipe. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | Recipe # 11 (TEOS-Ox @ 300°C) The tool MUST be capable of and a process recipe MUST be provided with the tool and demonstrated as part of the validation tests at the Tyndall National Institute site for deposition of TEOS-based oxide (TEOS-Ox) films on both 100 mm and 200 mm silicon wafers. Tenderers MUST clearly demonstrate how they meet the requirements outlined with results documented in the attached Appendix D – Deposition Recipe Results Template, this must be clearly named as “Lot 4 – Recipe # 11 -Results”. The Tenderer is permitted to use its own substrates and / or own deposition results to demonstrate compliance. Requirements: Dep. Rate TEOS-Ox: > 300 nm/min Stress range TEOS-Ox: < ± 200 MPa Refractive Index: 1.45 - 1.47 is desirable - Non-uniformity film thickness (WiW) on 100 mm wafer with 5 mm edge exclusion: < ± 5% - Non-uniformity film thickness (200 mm wafer with 5 mm edge-exclusion): < ±5 % Tenderer to specify process temperature and recommended gases for their recipe. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | Recipe # 12 (SiN @ 300°C) The tool MUST be capable of and a process recipe MUST be provided with the tool and demonstrated as part of the validation tests at the Tyndall National Institute site for deposition of Silane-based nitride (SiN) films on both 100 mm and 200 mm silicon wafers. Tenderers MUST clearly demonstrate how they meet the requirements outlined with results documented in the attached Appendix D – Deposition Recipe Results Template, this must be clearly named as “Lot 4 – Recipe # 12 -Results”. The Tenderer is permitted to use its own substrates and / or own deposition results to demonstrate compliance. Requirements: Dep. Rate SiN : > 100 nm/min Stress range SiN: < ± 200 MPa Refractive Index: 1.98 - 2.01 is desirable - Non-uniformity film thickness (WiW) on 100 mm wafer with 5 mm edge exclusion: < ± 5% - Non-uniformity film thickness (200 mm wafer with 5 mm edge-exclusion): < ± 5% Tenderer to specify process temperature and recommended gases for their recipe. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | Recipe # 13 (a-Si @ 300°C) The tool MUST be capable of and a process recipe MUST be provided with the tool and demonstrated as part of the validation tests at the Tyndall National Institute site for deposition of silane-based amorphous silicon (a-Si) films on both 100 mm and 200 mm silicon wafers. Tenderers MUST clearly demonstrate how they meet the requirements outlined with results documented in the attached Appendix D – Deposition Recipe Results Template, this must be clearly named as “Lot 4 – Recipe # 13 -Results”. The Tenderer is permitted to use its own substrates and / or own deposition results to demonstrate compliance. Requirements: Dep. Rate a-Si : > 50 nm/min Stress range a-Si : < ± 200 MPa Refractive Index : 3.7 - 4.0 is desirable - Non-uniformity film thickness (WiW) on 100 mm wafer with 5 mm edge exclusion: < ± 5% - Non-uniformity film thickness on 200 mm wafer with 5 mm edge-exclusion: < ± 5% Tenderer to specify process temperature and recommended gases for their recipe. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | Recipe - #14 SiN resistance to 5% TMAH solution For the SiN films, a key processing application is the encapsulation of MEMS devices and deposition of hard-mask layers for deep trench wet-etching (e.g. KOH, TMAH etc.) in silicon The SiN film (thickness ≥ 500 nm) as deposited over a single / double step of 2 µm height MUST be able to withstand immersion in a solution of 5% TMAH for at-least 4 hours without ingress through the deposited layer. Tenderers MUST clearly demonstrate how they meet the requirements outlined with results documented in the attached Appendix D – Deposition Recipe Results Template, this must be clearly named as “Lot 4 – Recipe # 14-Results”. The Tenderer is permitted to use its own substrates and / or own deposition results to demonstrate compliance. This MUST be demonstrated as part of the validation tests at the Tyndall National Institute site Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | Plasma Cleaning / Conditioning recipes The Tenderer MUST provide plasma cleaning recipes for inter-wafer / inter-batch cleaning and chamber-conditioning as part of the core package. Tenderer is to specify the recommended process temperature and gases for these cleaning recipes. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | Higher deposition-rate recipes It is desirable that the tool SHOULD have the capability to achieve deposition rates > 500 nm/min for Sil-Ox / TEOS-Ox / SiN deposition recipes at 300°C ± 20°C process temperature and this will be considered an advantage and scored higher. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | Lower-temperature deposition recipes It is desirable that the tool SHOULD have the capability to deposit the same Sil-Ox / TEOS-Ox / SiN / a-Si films as listed in Recipes #10, #11, #12, #13 & #14 above at < 200°C process temperature and this will be considered an advantage and scored higher. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | Recipe Support Tenderer MUST have ready access on request to a library of characterised deposition / cleaning / conditioning recipes for the tool and assistance in the form of BKMs / recipes / technical support by phone & email when required in setting-up new recipes. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | Roadmap & Flexibility: Clarity of roadmap for the Tenderer’s proposed tool is required, indicating the potential future upgrade options that will be available as well as highlighting existing upgrade options for the system proposed. As with all pieces of equipment in UCC, the ability to upgrade in the future to add additional functionality not currently specified is very important, i.e., the tool SHOULD have a future roadmap of development. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
| Hardware/Software Requirements |  |  |  |  |  |  |
|  | The system MUST be equipped with an aluminium process-chamber capable of deposition temperatures from 100°C to 350°C as a minimum requirement. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The process-chamber MUST be equipped with viewing port(s) for e.g. optical emission endpoint detection (EPD) and in-situ film thickness measurement systems at e.g. side and top chamber positions, and to directly view the plasma for stability. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The proposed endpoint solution(s) MUST be fully integrated with the system software. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | Tenderers SHOULD also supply further options of other EPD systems available for the tool. Please confirm and detail if any additional cost applies to additional EPD options that can be provided. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | In-Situ film thickness measurement It is desirable that the tool SHOULD have the capability to measure the thickness of deposited films in-situ for the Sil-Ox / TEOS-Ox / SiN / a-Si films as listed in Recipes #10, #11, #12, #13 & #14 above and this will be considered an advantage and scored higher. Please confirm and detail if any additional cost applies to in-situ film thickness measurement options that can be provided. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | Tenderers MUST identify clearly in their response if removable shields / liners for the process-chamber are required to allow for better process integrity / performance and to allow for rapid switch-over between the different deposition process recipes as listed in Recipes #10, #11, #12, #13 & #14 above. Tenderers MUST include a minimum of two (2) sets of removable shields / liners as part of the core package if these are required for running these deposition process recipes. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The system MUST have a dry backing pump for the process-chamber. Tenderers MUST specify the pump brands and model numbers which will be provided with the tool as part of the core package. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The dry backing pump for the process-chamber MUST be located remotely, positioned in a sub-fab with up-to a 10 m vertical drop and up-to a 10 m horizontal run to the point-of-use at the tool. Please see attached Appendix B – Sub-Fab Footprint for further details on the sub-fab footprint. The dry backing pump MUST have sufficient performance to sustain low pressure deposition recipes (< 10 Torr) as outlined in Recipes #10, #11, #12, #13 & #14 above and all necessary cleaning / conditioning recipes under these conditions. Sufficient length (minimum = 10 m) of electrical cable / plumbing etc. MUST be supplied to achieve this. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | Tenders SHOULD also provide information on further options of other dry backing pumps available for the process-chamber. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The process-chamber MUST recover to a pressure of ≤ 10 mTorr and a leak-up rate ≤ 2 mTorr / min after a wet-clean in less than twelve (12) hours. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The system MUST have a dry backing pump for the loadlock. Tenderers MUST specify the pump brands and model numbers which will be provided with the tool as part of the core package. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The dry backing pump for the loadlock MUST be located remotely, positioned in a sub-fab with up-to a 10 m vertical drop and up-to a 10 m horizontal run to the point-of-use at the tool. Please see attached Appendix B – Sub-Fab Footprint for further details on the sub-fab footprint. The dry backing pump MUST have sufficient performance to sustain low pressure (e.g. ≤ 1 Torr) for wafer transfer into the process-chamber under these conditions. Sufficient length (minimum = 10 m) of electrical cable / plumbing etc. MUST be supplied to achieve this. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | Tenderers SHOULD also provide information on further options of other dry backing pumps available for the loadlock. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | Loadlock pump-down time SHOULD be less than 120 seconds, and no longer than 300 seconds. Faster pump-down time will be considered an advantage and scored higher. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The system MUST have a heated foreline to remove deposition and cleaning plasma by-products from the process-chamber. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The system MUST have an automatic throttle valve for pressure control in the process-chamber. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The system MUST be water-cooled. Process cooling-water supply is provided by Tyndall, see section D. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The process-chamber platen / wafer-electrode temperature MUST be controllable (using e.g. resistive-heating / chiller / air-cooling etc.) to allow deposition temperature in the range 100°C to 350°C with control of ± 2°C from set point. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The required deposition performance of the tool MUST be achieved by either dual high-frequency (e.g. 13.56 MHz) and low-frequency (e.g. 100-500 kHz) plasma configuration with independent RF controls between the two (2) plasma sources or with one (1) single-frequency plasma source. The capability of having an additional low frequency RF generator is considered highly desirable for tailoring of film stress for SiN and this will be considered an advantage and scored higher. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | System MUST be equipped with a fully auto-matched RF generator for each plasma source. The RF generators MUST have capability for plasma strike and sustain tuned plasmas operating over a wide processing pressure range < 10 Torr. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | Tenderers MUST specify details of the RF generator brand & model, frequency, power and operating range for each plasma source which will be provided as part of the core package. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The RF generators / racks may be located remotely, positioned in a sub-fab with up-to a 10 m vertical drop and up-to a 10 m horizontal run to the point-of-use at the tool. Please see attached Appendix B – Sub-Fab Footprint for further details on the sub-fab footprint. Sufficient length (minimum = 10 m) of electrical cable / plumbing etc. MUST be supplied to achieve this. The RF generators / racks MUST have sufficient performance for plasma strike and sustain tuned plasmas under these conditions. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | RF Generator racks MUST be supplied by the Tenderer, if required for the assembly of the system. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. Please confirm and detail. |  |  |  |  |  |
|  | Tenderers SHOULD also provide information on further options of other RF generators available. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. If any options to purchase are available, these are optional to purchase by the Contracting Authority and will not form part of the Cost evaluation. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | Enhanced Chamber Cleaning The capability for enhanced (faster) cleaning of the chamber post-deposition will be considered an advantage and scored higher. Tenderer is to specify the recommended gases for these cleaning recipes. If additional hardware (e.g. remote plasma source NF3 cleaning, high-power clean) is required for enhanced chamber-cleaning, the Tenderer MUST identify clearly in their response and cost separately. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The tool MUST accommodate wafers as set out in requirements A11 – A13 above as a minimum requirement. The tool MUST be configured for automated loading of individual substrates / carriers from a cassette to and from the process-chamber (e.g. separate 25-slot cassettes for both 100 mm and 200 mm diameter wafers or 12/13-slot double-spaced cassettes for both 200 mm diameter wafers and 200 mm diameter wafer-carriers). Operators will load wafers / carriers into the cassette and the wafers / carriers MUST be transferred from the cassette to & from the process-chamber via robotic-arm. Tenderers MUST include a minimum of two (2) sets of cassettes for each diameter (if required) as part of the core package if these are required for running the deposition process recipes on both 100 mm and 200 mm diameter wafers. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The hardware changes / procedures required to switch from running one wafer-diameter to another MUST be outlined clearly by the tenderer with all necessary hardware (e.g. robotic end-effectors, cassettes, wafer-carriers etc.) MUST be provided as part of the core package. The ability to switch from e.g. 100 mm to 200 mm substrate diameter with minimal downtime and / or hardware changes will be considered an advantage and scored higher. Please confirm and detail changeover procedure and time to complete. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | If wafer-carriers are required to enable running deposition recipes on both 100 mm and 200 mm diameter wafers with minimal downtime and / or hardware changes, Tenderers MUST include a minimum of two (2) sets of these wafer-carriers (e.g. for 12/13-slot cassettes, one (1) set = 12/13 x 200 mm diameter wafer-carriers) to be provided as part of the core package. Tenderer should also specify the material type of the wafer-carrier and this MUST be compatible with silicon CMOS device processing. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  |  |  |  |  |
|  | The tool MUST have capability to handle wafer thicknesses in the range ≥ 0.2 mm and ≤ 1 mm as a minimum requirement. An ability to also handle substrates > 1 mm thickness with the same setup will be considered an advantage and scored higher. Please specify minimum and maximum thickness in tender response and identify any hardware changes (e.g. separate wafer-carriers for different thickness wafers at 100 mm / 200 mm diameters) required to allow a range of thin / thick substrates to be processed. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The tool SHOULD have capability to handle bonded-wafer pairs, perforated wafers, or bowed wafers (with bow / warp > 50 µm). This will be considered an advantage and scored higher. Please identify any hardware changes (e.g. separate wafer-carriers etc.) required to allow for handling these non-standard types of substrates. If any additional hardware (e.g. wafer-carriers etc.) is required for handling these non-standard types of substrates, the Tenderer MUST identify clearly in their response and cost separately. Tenderers are to provide details and pricing for the proposed option within the core package in the Appendix 2, Pricing Schedule. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The system MUST have a minimum of seven (7) gas lines plus one (1) carrier-gas / vapour delivery line for TEOS. An ability to have more than eight gas / carrier-gas lines will be considered an advantage and scored higher. Suggested process gases / liquid sources include, but are not limited to the following; SiH4, O2, N2O, N2, NH3, Ar, and TEOS. The tenderer MUST specify their recommended gases for the recipes outlined in requirements Recipes #10, #11, #12, #13 & #14 above and for plasma cleaning & conditioning of the process-chamber. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The TEOS liquid / vapour delivery hardware must be fully integrated with the tool system and software and ALL necessary hardware for TEOS vapour delivery MUST be included as part of the core package. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | In the specific case(s) of liquid sources (e.g. TEOS) which are susceptible to condensation, ALL necessary vapour-delivery hardware such as heated gas-lines, heated gas-box, heated manifolds / MFCs etc. MUST be included as part of the core package. Tenderer MUST also specify their preferred supplier for TEOS liquid and standard canister specification. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | In the specific case of oxygen (O2) the requirements for e.g. low flow-rate for deposition of thin layers of silicon dioxide and high flow-rate for chamber plasma cleaning / chamber plasma conditioning may necessitate two separate MFCs, so this MUST be clearly specified by the Tenderer. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The tenderer MUST specify recommended flow size for each gas Mass Flow Controller (MFC). MFCs supplied MUST be rated (in terms of e.g. seals etc.) and calibrated for the recommended gases as nominated by the Tenderers under requirement A55. Digital MFCs are preferred and will be considered an advantage and scored higher. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The capability to run automatic interleaved and / or wafer-less clean MUST be available in the software. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | Whole system MUST be fully interlocked to protect against any equipment or service failure. System MUST be designed to fail-safe with system left under vacuum. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | Electrical equipment racks may need to be located remotely, positioned in a sub-fab with up-to a 10 m vertical drop and up-to a 10 m horizontal run to the point-of-use at the tool and MUST be capable of full functionality under these conditions. Sufficient length (minimum = 10 m) of electrical cable MUST be supplied to allow the required performance. Please see attached Appendix B – Sub-Fab Footprint for further details on the sub-fab footprint. Confirmation is required that the operation of the tool and etch quality will not be impacted by remote location of the electronic control cabinets. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | Electrical equipment racks MUST be supplied by the Tenderer, if required for the assembly of the system. Please confirm and detail. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | A minimum of one (1) computer and console MUST be included in core package, and MUST use Windows-based OS, Windows 10 or higher. Please describe the software package proposed, including name and version. Please also specify the projected lifespan of the current control system architecture. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The tool will be installed in the cleanroom in a thru-wall configuration. ALL computer(s) / console(s) required for tool operation from the front (process bay) and rear (service chase) side MUST be included in the core package. Additional console(s) / computer(s) for tool operation from the rear (service chase) side MUST also be included in the core package, or a suitable alternative such as a rotating screen(s) that are visible from both bay and chase with associated keyboard / mouse control from both sides. Please see attached Appendix A -Cleanroom Footprint for further details on the cleanroom footprint. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The tenderer MUST provide all control and analysis software required for tool operations. All tool operations, including wafer transfer to / from the cassette to / from the loadlock to / from the process-chamber MUST be software-controlled and allow for both automatic & manual operation. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The Plasma Etcher equipment SHOULD have perpetual licence-free software or no restriction on the number of licences needed for our use of the system i.e., licences for multiple users and computer installations for off-site use if this is possible. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | All software updates to the system, which become available during the life of the contract, MUST be provided free of charge during the warranty period and for a minimum period of eight (8) years after warranty expiry. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | Continued development of the software SHOULD be demonstrated by the Tenderer, with facilities for updating systems in use. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The system MUST be capable of remote access via network connection (using e.g. TeamViewer software or equivalent) for tool support in event of problems. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The software MUST be capable of data-logging of process parameters, and archiving data-logs. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | It MUST be easy to graphically display and export saved process parameter data preferably into Excel spreadsheets. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The system MUST have an easy-to-use interface and needs to be suitable for a multi-user (i.e. different access levels) environment. Various levels of access are required e.g. Operator, Supervisor, Maintenance. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | The software SHOULD have the facility to lock certain recipes i.e. not editable by users other than Process Engineer / Supervisor / Maintenance etc. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | Software MUST include backup and restore capability (export/import) for process recipes and machine settings. Software SHOULD include backup (image) capability to be used in case of hard disk drive (HDD) failure. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
|  | Software MUST allow operator access to data files while the system is running under standard operational conditions. Please confirm and detail. |  | Confirm (Yes / No) |  |  | Insert Comment |
| - Quality, Health & Safety & Validation |  |  |  |  |  |  |
| Quality Requirements |  |  |  |  |  |  |
|  |  | Compliance with all relevant European health and safety standards, quality assurance and protocols are MANDATORY. All Goods and Services supplied in accordance with the contract MUST comply with all necessary standards, i.e. Irish Standards Institution, European Union, CE, International Standards Organisation or other applicable standard. The Tenderer MUST, on the request of the buyer, provide a copy of the standard free of charge. Please confirm & detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | If there are manufacturing defects known to the tenderer, they MUST identify such defects and state their policy regarding the repair of known defects. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | The Tenderer SHOULD, as part of the submission, include copies of any relevant testing or compliance certificates for the Goods. (For example, Factory acceptance tests, Quality assurance test etc.) Please confirm and detail |  | Confirm (Yes / No) |  | Insert Comment |
| Health & Safety Requirements |  |  |  |  |  |  |
|  |  | All goods supplied in accordance with the contract MUST be of good construction, sound material, of adequate strength, free of design defects and manufactured to be safe and without risk when properly used and software MUST allow operator access to data files while the system is running under standard operational conditions. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | The tool SHOULD meet the criteria of the Semi S2-0821Eb Environmental, Health, and Safety Guideline for Semiconductor Manufacturing Equipment in order to ensure easier and safer installation and maintenance from a gas safety perspective. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | The Tenderer MUST, as part of the submission, include Laser Safety Warning labels and copies of any relevant testing or compliance certificates for the Goods. Confirm compliance and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | The tools MUST knowingly not present any Health and Safety hazards. Any safety & hazard warning labelling on the proposed system MUST be in the English Language. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | Tenderers MUST advise if there are any hazards associated with the operation or maintenance of this instrument? (e.g. manual handing, chemicals, etc.). Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| - Warranty/Spare Parts, After-Sales & Training Support |  |  |  |  |  |  |
| Warranty & Customer Support, Repair/Breakdown Requirements |  |  |  |  |  |  |
|  |  | Warranty Duration: At least one (1) year’s warranty from Site Acceptance Test (SAT) MUST be included for each lot as part of the core package and the quoted price is to include all hardware, software, spare parts, labour, call outs, visit, travel Warranties of greater than 1 year are considered highly desirable and will be scored higher. Detail specific inclusions and exclusions of warranty. Preventative Maintenance MUST be provided for at least one (1) year and a list of checks carried out at Preventative Maintenance MUST be provided, along with any information about which (if any) parts are routinely replaced. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | If UCC decide to extend the warranty period, Tenderers are asked to outline costs of any additional warranty. The Tenderer MUST state the details of any extended warranty and details of Preventative Maintenance and calibration & servicing check and any associated costs. Detail specific inclusions and exclusions of extended warranty and a list of checks carried out at Preventative Maintenance SHOULD be provided, along with any information about which (if any) parts are routinely replaced. Please include price per annum, details of coverage and maximum term of coverage. Please include price per annum for year 2, 3, & 4, details of coverage and maximum term of coverage. Please confirm and detail and specify what is included in the warranty options outlined in the Appendix 2 pricing schedule, Optional to Purchase. Detail specific inclusions and exclusions of extended warranty. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | Please provide separately a list and approximate cost of any consumable items/parts that are to be replaced during each service visit. Please update Appendix 2 Pricing Schedule – Consumables Parts sheet with detail. Prices provided for consumable items/parts MUST be valid for a minimum of one (1) year. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | Please advise of any items/part of the instrument that require routine replacement during the normal lifetime of the instrument, the timeframe in which they require replacing and if such items will be supplied with the instrument or as part of the warranty. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | The tenderer MUST provide a description of the arrangements the arrangements and lead times for guaranteed responses times in the event of equipment breakdown both in relation to phone support, online/remote support, the provision of technical support, annual service/ maintenance, emergency call-out response, and spare parts. Please detail lead time for arrival on site by a fully trained service engineer during the warranty period(s) and post-warranty. This response time SHOULD be no more than 24 hours via phone call or remote service, and within seven (7) working days for onsite repair. Please confirm and detail information on remote diagnostic / telecom-based service support. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | Outside of Warranty Period: The Tenderer MUST provide typical prices and lead time for servicing and repairs (including travel, subsistence and price per day) charged outside a warranty/ maintenance contract. Detail specific inclusions and exclusions outside of warranty period. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | If UCC decide to take out an additional Preventative Maintenance Agreement / SLA, tenderers are asked to outline Options for years 2, 3 & 4. Option to include: - On-going Calibration to appropriate schedule - preventative maintenance - remote diagnostics of faults, if applicable - repair of the Instrument in case of breakdown) Tenderers are asked to outline full details of what is included and to outline Cost of maintenance and calibration for each year option offered. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | The tenderer MUST have dedicated fully trained and qualified service engineering specialists at their disposal who can respond to all service calls and any form of breakdown call. This remote response time SHOULD be the same day via phone call or remote service, and within seven (7) working days for onsite repair. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | Please outline the company’s Customer Support Proposal and escalation policy. Tenderers MUST describe what process is in place for escalating a call-out, if a breakdown cannot be resolved on the first visit. Tenderer SHOULD provide a list of approved service and calibration agents. Please confirm and detail customer support arrangements, service and calibration agents. Additional support and service options should also be outlined and pricing noted within the Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | The Tenderer MUST indicate the arrangements for and specify the level of experience of applications specialists and engineers who will be in charge of after sales service. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | The Tenderer SHOULD specify the lead time for repairs during the following scenarios. - warranty period(s) - post-warranty, - with service contract - without service contract Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | Please note SHOULD installation, commissioning, service visits or availability of spare parts be affected in the event of a Pandemic crisis it is expected that the warranty period will be extended to facilitate the service downtime endured by the Contracting Authority. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | Warranty commences post full installation and sign-off of tool acceptance by nominated UCC representative. Please confirm and detail. |  | Confirm (Yes / No)) |  | Insert Comment |
| Replacement & Spare Parts & Equipment Consumables |  |  |  |  |  |  |
|  |  | Spare Parts: Availability of all spare parts SHOULD be maintained 10 years after the warranty period. Please indicate the availability and lead time of spare part stock centres, with maximum lead-time for any part. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | Critical spares MUST be held at a location no more than 7 Days delivery time from Tyndall. Spare parts to be held offsite with the tenderer and essential parts if necessary to be held Free of Charge (FOC) at Tyndall. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | Consumable Parts: Sufficient consumable parts MUST be supplied for the initial installation, commissioning and training. These consumable parts MUST be provided free of charge and at the point of commissioning. Full one (1) year of consumable parts to be supplied post commissioning, including any removable shields / liners for process-chamber. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing must only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | Recommended Spare Lists: The tenderer MUST supply a full list of recommended spare parts and consumables for the proposed equipment, based on knowledge of equipment design and component lifetimes. This SHOULD include typical consumables and high turnover parts. Please detail in Appendix 2, Pricing Schedule- Consumable Parts. Prices provided for consumable items/parts MUST be valid for a minimum of one (1) year. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | The tenderer SHOULD supply, where information is available, to the buyer a schedule of spare parts showing in tabular form or, otherwise, the following information on OEM parts: - Names and address of all OEM suppliers - Part Name - Part Description - OEM Supplier Part Number. Please confirm and detail, attaching information requested above. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | The tenderer MUST provide confirmation of the availability of service for at least five (5) years after commissioning. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| Training & Technical Support Proposal |  |  |  |  |  |  |
|  |  | The successful tenderer MUST provide an initial training programme for UCC personnel on-site at Tyndall, for a minimum of one (1) day each for operations and maintenance, with a minimum of three (3) users fully trained to safely operate and maintain the system as part of the core package. The training programme MUST be carried out by a suitably qualified specialist at the tool after the initial installation, commissioning is completed. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | Initial training provided MUST include, but not be limited to: - Demonstration of the correct operation of all relevant equipment of the systems - Routine Maintenance requirements - Basic Troubleshooting - Data processing and management - Software and hardware familiarisation Dates, timing and method of training are to be confirmed between UCC and the winning tenderer. The process validation tests may also be used for on-site training purposes. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | The tenderer MUST specifically outline the level and duration of application support offered with the instrument where “application support” refers to the expert technical support (on-line or on-site) available from the manufacturer for using the instrument in novel research applications. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | This equipment will have multiple users over the lifetime of the tool. Tenderers MUST provide details and pricing for their proposed option for Operations / Applications / Maintenance Training at all levels (e.g. Level1, Level2, Level3 etc.) on-site or at their facilities. Training offered over the full lifetime of the equipment will be considered an advantage and scored higher. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |  | Confirm (Yes / No) |  | Insert Comment |
|  |  | Relevant training and maintenance manuals in English MUST also be provided with the instrument. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| - Delivery Plan |  |  |  |  |  |  |
| Lead Time & Delivery |  |  |  |  |  |  |
| D1 |  | The equipment SHOULD be delivered within twelve (12) months of receipt of Purchase Order. The equipment SHOULD be delivered within the stated lead time requirement from receipt of Purchase Order. (Successful tenderer for each lot may be required to liaise with each other for coordination of delivery and installation of the equipment) - Tenderers MUST provide full details of lead times for the delivery of equipment including installation and commissioning (if applicable), from receipt of the purchase order. - Tenderers MUST provide the facility requirements to UCC a minimum of 16 weeks in advance of the equipment delivery. - Tenderers MUST provide advance notice (two weeks) of any requirements for tool unloading, number of containers and weight of each. Height, length and width of containers/crates to be provided as well. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| D2 |  | Delivery MUST include all shipping charges, insurance, customs clearance and all aspects as outlined in DAP (Delivery At Place) incoterms. Delivery MUST be to Tyndall National Institute, UCC: - Silicon Fabrication Laboratory (Room A1.1.04), - Tyndall National Institute, - University College Cork, - Lee Maltings, Dyke Parade, - Cork T12 R5CP, - Ireland. Advance notice MUST be given (2 weeks) of any requirements for tool unloading, number of containers and weight of each. Height, length and width of containers/crates MUST be provided as well. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  | Insert Comment |
| D3 |  | The tenderer MUST provide full details of lead times for the delivery of equipment including installation and commissioning (if applicable), from receipt of the purchase order. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| D4 |  | The tenderer MUST provide details of the logistics/supply chain to the point of customer delivery, i.e. where the goods will come from, how the goods will be transported and what borders the goods will cross. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| D5 |  | The tenderer MUST provide details of any actions that have been taken to ensure the future timely supply and delivery of goods and services, including ongoing supply of replacement/spare parts and consumables. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| D6 |  | The tenderer MUST provide details of their contingency plans (alternative sources/routes of supply) for any delays in delivery should they arise. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| Installation, Commissioning, Calibration & Validation Requirements |  |  |  |  |  |  |
| D7 |  | Pre-shipment test SHOULD demonstrate the system’s capability to achieve all process parameters outlined above. Pre-Shipment test results should be shared as agreed between Tyndall and the Tenderer, prior to shipment. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| D8 |  | The tenderer MUST commission the system and verify full functionality at the Tyndall National Institute site. The buyer's engineering staff (both process and equipment) will provide reasonably requested assistance to the supplier in this task. Tyndall will also provide for any labour, chemicals, wafers, power, lubricants, water, fuel and materials reasonably requested for the purpose of the process installation. Tenderer has to make a written request for this support at least four weeks in advance of installation. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| D9 |  | The facility requirements for this piece of equipment MUST be supplied to Tyndall a minimum of sixteen (16) weeks in advance of the equipment delivery. These facilities changes will be installed to the specifications supplied to Tyndall by the tenderer. Any deviation to the supplied specification on the delivered equipment causing a reconfiguration of the facility supply will be at the tenderer’s cost. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| D10 |  | The proposed system MUST be suitable for use with the 400 V, 50 Hz, 3-phase electrical supply at the proposed installation site. If a transformer unit is required to achieve this, it MUST be supplied as part of the core package. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| D11 |  | Tenderers SHOULD also note that the following facilities will be available at the installation site: - 3-phase electrical supply as per D10 above - Process cooling water supply - De-ionised water supply - Waste water drainage - Compressed air supply - Air exhaust. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| D12 |  | Tyndall Cleanroom The tenderer MUST provide footprint drawings for tool and ancillary equipment (e.g. control consoles etc.) required for operating the tool which MUST be located in the cleanroom. The available cleanroom space will be limited so tools and ancillary equipment MUST fit within a space approx. 5 m2 per tool (total footprint for the suite of tools being tendered-for is approx. 20 m2. Vacuum pumps / chillers / equipment racks etc. will be located remotely, positioned in a sub-fab with up-to a 10m vertical drop and up-to a 10m horizontal run to the point-of-use at the tool. This equipment MUST fit within a space approx. 4 m2 per tool (total footprint for the suite of tools being tendered-for is approx. 16 m2 Please see guideline footprint drawings of the cleanroom and sub-fab locations where the suite of tools will be installed in attached Appendices; Appendix A- Cleanroom Footprint & Appendix B – Sub-Fab Footprint. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| D13 |  | The tools MUST have thru-wall configuration for installation in the Silicon Fabrication cleanroom. All hardware (e.g. front panel, keyboard(s), monitor(s), etc.) required for thru-wall configuration MUST be included in the core package. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| D14 |  | The installation, commissioning, and calibration of the equipment MUST be carried out on-site by a qualified engineer provided by the Tenderer. This entails the system being transported to the lab, unpacked, set up and demonstrated to meet the requirements as outlined in section A. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |  | Confirm (Yes / No) |  | Insert Comment |
| D15 |  | Prior to installation, the machine may be inspected at the Tenderer site by a Tyndall engineer to ensure that it meets the agreed specifications. Particular attention will be paid to the facilities links to ensure that the equipment is in a state ready for connection to the installed facilities at the proposed location in Tyndall. This inspection will also ensure that the machine as built corresponds exactly to the supplied drawings. Please confirm compliance. |  | Confirm (Yes / No) |  | Insert Comment |
| D16 |  | Equipment start-up is tenderer’s responsibility. Prior to start up, the supplier may inspect the facilities to determine whether they meet requirements. Start-up of the equipment shall indicate the tenderer’s acceptance of the facilities installation. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| D17 |  | Tyndall will carry out a Functional Acceptance Test. The equipment functionality test consists of the tenderer’s equipment start-up test. Prerequisites for the functional acceptance are: - Complete Delivery and installation - Resolution of all outstanding issues resulting from pre-shipment test, if applicable - Fulfilment of the following tests: - Shut-down test (ESO operation and re-start) Please confirm acceptance of this. |  | Confirm (Yes / No) |  | Insert Comment |
| D18 |  | The Tenderer MUST provide a qualified engineer to perform the SAT (Site Acceptance Test) post commissioning. The validation tests will be done in partnership with experienced Tyndall process engineers to ensure that the system performance has not been affected by shipping. These tests may also be used for on-site training purposes. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| D19 |  | The acceptance tests SHOULD include substrate handling tests in sequence without alarm conditions; the number of handling cycles to be mutually agreed between buyer and tenderer based on the system spec. - Recovery from power outage, extract outage, vacuum outage, nitrogen outage. - Delivery of all required documentation. - A copy of the tenderer’s pre-shipment final systems test procedure. - Delivery of supplier specified consumables. If the functional acceptance test produces negative results, Tyndall shall be entitled to request that the equipment will be brought into the guaranteed condition within 2 (two) weeks of the failed test. Please confirm compliance and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| D20 |  | Site Acceptance Test After functional acceptance, the Site Acceptance of the equipment shall be done jointly by engineers assigned by Tyndall and by the tenderer. A final acceptance document shall be prepared with the following content: - The purchase order reference. - The procurement specification reference. - A list of outstanding issues, including an action time frame for resolving these issues and the division of responsibility for this work. - The signatures and functions of the accepting engineers on each side, the Head of Process and Product Development group and the Tyndall Maintenance Manager. - The date of acceptance. - Serial number of the equipment. If the items in the list are satisfied, the equipment will be accepted. If one or more items are not accepted, the tenderer will be required to rectify the deficiency or, alternatively, a partial acceptance may be agreed combined with an extension of the warranty period. Terms shall be negotiated between Tyndall and tenderer based on the results of the acceptance tests. If there is a serious problem on the system, a mutually agreed plan of correction will be generated to correct the problem. Tyndall may waive any portion of the plan. The plan may include the replacement of the system at the discretion of the tenderer. Please confirm compliance. |  | Confirm (Yes / No) |  | Insert Comment |
| D21 |  | Tenderers are requested to provide a summary of scheduled maintenance and calibration required to keep instrument working to the full specification as per manufactures recommendations. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| D22 |  | A signed certificate for successful site acceptance test MUST be provided at the end of the installation and training. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| Documentation |  |  |  |  |  |  |
| D23 |  | A detailed User Operational Manual for the unit and control system MUST be provided. This SHOULD include the following; - Full electrical and design specifications. - Manuals MUST be supplied in English language. Manuals MUST be provided in soft copy. |  | Confirm (Yes / No) |  | Insert Comment |
| D24 |  | A Maintenance Manual for the system MUST be provided to include the following; - Cleaning requirements, service - Maintenance & Calibration requirements. - Manuals MUST be supplied in English language. - Manuals MUST be provided in softcopy where possible. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| D25 |  | A Technical Specification for the Unit MUST be provided covering all major components of the system. The specification MUST be supplied in the English language and in (if possible) in soft copy. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| D26 |  | A Technical Specification for all instrument documentation relating to installation of hardware and software, hardware/software interfacing, service and maintenance, parts list, and software operation and editing documentation MUST be supplied in English language. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| D27 |  | Factory Acceptance Test The equipment operation MUST be validated FAT (Factory Acceptance Test) prior to shipment to Tyndall to ensure it meets performance specifications. - The equipment will be shipped to Tyndall National Institute, Cork when Manufacturer validation tests have proved the equipment performs to the required specifications. Validation certificates MUST be provided during installation and commissioning. Please confirm and detail. |  | Confirm (Yes / No) |  | Insert Comment |
| - Energy, Environmental & Sustainable Considerations |  |  |  |  |  |  |
| Environmental & Sustainable Considerations |  |  |  |  |  |  |
| E1 |  | Please outline your company’s commitment to the sustainable life cycle of the tendered equipment in a manner that minimises adverse environmental impacts. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| E2 |  | Energy & Environmental Requirements: The Contracting Authority requires the equipment provided to be as energy and economically efficient as possible. The system SHOULD have low running costs allowing the equipment to be as energy & economically efficient as possible. Where applicable, and if possible, the equipment SHOULD include the following: - Use of water-cooled units, e.g. water cooled chillers / pumps instead of air cooled chillers/ pumps as the site recovers heat from the processing cooling load to heat other areas. - Eco mode vacuum pumps to enable nitrogen saving when system in standby, control of both the nitrogen flow and eco mode activation SHOULD be from the system software - Programmable capability, i.e. if possible equipment SHOULD come with the option of programmable timer so that processes can start overnight and the site can benefit from the reduced night-time electrical rates. Please provide any details on the above or other measures taken to improve energy and efficiency of the tool. Please provide details on any “warm-up” times required for the equipment to operate to specification. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| E3 |  | Tenderers SHOULD provide details of any takeback arrangements they may offer when the equipment reaches end of life and how that may be availed of. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| E4 |  | Tenderers SHOULD indicate Typical Energy Consumption during operational time while the equipment is in use. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| E5 |  | Where possible, use of a standby mode/reduced power mode of operation when the system is not in use. The tenderer SHOULD specify the stand-by energy consumption. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| E6 |  | It is desirable that all Tender requirement is delivered using environmentally friendly Packaging Materials; packaging, tape, labels and stretch wraps. Ideally, this packaging SHOULD avoid any single use plastic and be recovered or at a minimum reusable. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| E7 |  | Tenderers SHOULD provide details of expected lifetime of the equipment required. The Tenderer MUST state how long the model of equipment has been on the market and indicate the expected working life of the proposed models. Longevity of equipment etc. Please confirm and detail. |  | Confirm (Yes / No) | Insert Comment |  |
| - Ultimate Cost. (Refer to Appendix 2: Pricing Schedule for Completion) |  |  |  |  |  |  |

## B4: Appendix 2- Pricing Schedule

Please use Appendix 2, which is provided as a supplementary Excel, titled:

LA3215C Appendix 2 Pricing Schedule

The Appendix 2 Pricing Schedule provided in the Excel file MUST be completed per Lot/section(s) you are Tendering for. Please submit in Excel format. The relevant Pricing Schedule must be fully completed, totalled, and included with the Tender submission.

Tenderers are reminded that the prices/discounts quoted MUST comply with the requirements outlined in paragraph 2.10 of the RFT, i.e., they MUST be all-inclusive, fixed, expressed in Euro, exclusive of VAT, and remain unchanged for the duration specified in the RFT.

Any particulars and merits should be specified in the comments section to assist with scoring under paragraph 3.3 Award Criteria of the RFT.

## B5-Checklist for Tender Submissions

| Please confirm that all the information requested in Sections A1-A9 has been provided | Confirm Yes/No |
| --- | --- |
| Please confirm that the Appendix 2 Pricing Schedule has been returned and included in your proposal. | Confirm Yes/No |
| Please confirm that Appendix 3 [Tenderer’s Statement], (on headed paper, In TRD and RFT, has been completed and returned as requested. | Confirm Yes/No |
| Please confirm that Appendix 4 [Declaration as to Personal Circumstances of Tenderer] in TRD and RFT, has been signed and completed and returned as requested. This document MUST also be signed, stamped and dated by a Practising Solicitor / Commissioner for Oaths. | Confirm Yes/No |
| Tenderers must address each of the issues and requirements and submit a detailed description in each case which demonstrates how these issues and requirements will be dealt with / met and their approach to the proposed delivery of the Goods. A mere affirmative statement by the Tenderer that it can / will do so or a reiteration of the tender requirements is NOT sufficient in this regard. | Confirm Yes/No |
| If multiple equipment types are available which suit the specifications required, please provide details of each option on a separate TRD and Appendix 2, clearly defining each option. | Confirm Yes/No |
| Please ensure that all supporting documentation is submitted in the English language. | Confirm Yes/No |
| Please ensure all documents are named accordingly and submitted within a zipped/compressed folder. | Confirm Yes/No |

Please ensure to sign the confirmation form at the end of this document.

This form must be completed and signed by an authorised officer of the tenderer’s organisation.

- This offer will remain open for acceptance by you for a period of 180 days from the date of the deadline for submission of tenders.

- We acknowledge that you are not obliged to accept the lowest or any offer.

- I/We undertake to deliver the products/services in accordance with the terms and conditions of the tender specification.

- I/We undertake to maintain full confidentiality with regard to all aspects of this tender process

| Signed: |  | Position: |  |
| --- | --- | --- | --- |
| Print Name: |  | Phone No: |  |
| Company Name: |  | Date: |  |
| Address: |  |  |  |
| Email: |  |  |  |

Document end