Appendix A.	d33 Measurement on Clamped Thin Films

THIN-FILM PIEZOELECTRIC CHARACTERISATION ON CLAMPED STRUCTURES: Effective d33 Measurement on Clamped Thin Films

The platform MUST be capable of measuring the effective longitudinal piezoelectric coefficient of clamped thin-film capacitor structures on substrates and wafers.

Tenderers MUST complete the following table to define the performance and requirements of their proposed thin-film piezoelectric characterisation subsystem.

| Technical Attribute | Tyndall Minimum Requirement ("MUST") | Tenderer's Proposed Specification | Detailed Description / Supporting Doc Reference |
| --- | --- | --- | --- |
| Measurement Principle | Suppression of substrate bending |  | [Must detail method] |
| Min. Displacement Resolution | Better than 1 pm |  | [Reference specific sensor spec] |
| Repeatability (d33,f ) | ≤ 1% to 5% Sigma on Quartz reference |  | [Reference test data] |
| Substrate Thickness Range | 50 µm to 1,500 µm (1.5 mm) |  |  |
| Backside Optical Access | State whether required or whether not applicable for proposed method |  | [Describe sample holder access] |
| Backside Reflectivity | State if metallisation or polishing is required |  |  |
| Max Sample Size | At least 100 mm (4-inch) (Wafer Mapping) |  |  |
| Min Electrode Dimension | Compatible with laser spot size |  | [State laser spot size] |
| Bottom Electrode Access | Method for common bottom contact on wafers/coupons |  | [e.g., Conductive Chuck or Probe] |
| Top Electrode Geometry | Must state required configuration (e.g., round, rectangular, or equivalent) |  |  |