Tender Response Document

Tender Reference: LA3673C

Request-for-Tenders under Open (OJEU) Procedure for the Supply, Delivery, Installation, Commissioning & Calibration of a Fully Integrated Precision Ferroelectric and Piezoelectric Analysis Platform, for Tyndall National Institute, University College Cork (UCC)

Location:

Tyndall National Institute – Cork, Lee Maltings Complex, Dyke Parade, Cork, T12 R5CP, Ireland.

| Issue Date: | Wednesday 22nd July 2026 |
| --- | --- |
| Contact Person for Queries: | Gráinne Quinn |
| Contact Details for Queries: | E-Tenders Mailbox |
| Closing Date for Receipt of Queries: | 12.00pm On Thursday 13th August 2026 |
| Deadline For Receipt of Tenders: | 12.00pm On Monday 24th August 2026 |

Contents

1.	Instructions for the Completion of the Tender Response Document.	3

2.	Introduction	4

3.	Summary of Requirements	6

4.	Section A: Selection Criteria – RFT Section 3.2	8

A1: Insurance	8

A2: Tax Clearance	10

A3: Statement of Turnover	11

A4: Audited Accounts & Reports	11

A5: Data Protection	12

A6: Quality Assurance & Accreditation	12

A7: Organisational Summary	13

A8: Previous Experience	14

A9: Health & Safety	15

A10: Appendix 3 Tenderers’ Statement	17

A11: Appendix 4 Declaration as to Personal Circumstances of Tenderer	19

A12: Appendix 5 Declaration and Warranty on the Use of Artificial Intelligence in the Preparation of the Tenderer’s Response	22

5.	Section B: Tender Response	24

B1: Requirements & Specifications	24

B2: Award Criteria	26

B3: Appendix 1-Technical Specification	29

B4: Appendix 2- Pricing Schedule	71

B5: Checklist for Tender Submissions	71

# Instructions for the Completion of the Tender Response Document.

Tenders MUST submit the Tender Response Document (TRD) duly completed. When completing this TRD please note the following:

- Please ensure you complete all of Section A before you progress to Section B.

- Tenderers should note they MUST meet and ‘Pass’ all the requirements contained in Section A, the Selection Criteria, to be considered eligible for assessment against the award criteria, Section B.

- You need to COMPLETE AND RETURN 2 DOCUMENTS. This TRD and the accompanying Appendix 2 Pricing Schedule.

- You DO NOT need to return the RFT Document. The RFT document is for information purposes only.

TRD - Section A Selection Criteria

- A1-A4: Economic & Financial Standing

- A5-A9 Technical & Professional Ability

- A10: Appendix 3 Tenderer’s Statement

- A11: Appendix 4 Declaration as to Personal Circumstances of Tenderer

- A12: Appendix 5 Declaration and Warranty on the Use of Artificial Intelligence in the Preparation of the Tenderer’s Response

TRD - Section B: Tender Response

- B1: Requirements & Specifications

- B2: Award Criteria

- B3: Appendix 1-Technical Specification

- B4: Appendix 2- Pricing Schedule (separate Excel File)

- B5: Checklist for Tender Submissions

The completed TRD must be returned in accordance with Section 2.6 of the RFT.

Suppliers please note:

eESPD - When submitting tenders on the eTenders platform, please ensure to complete the online eESPD as part of the tender submission. Instructions are included in the tender documents.

# Introduction

University College Cork (UCC)

UCC is an award-winning institution with a history of independent thinking stretching back over 170 years. UCC is proud to be ranked in the top 2% of universities in the world. Their beautiful university opened its gates to just 115 students in 1849. They now have a student population of over 21,000.

Tyndall National Institute

Tyndall National Institute is a leading European research centre in integrated ICT (Information and Communications Technology) hardware and systems. Specialising in both electronics and photonics – materials, devices, circuits and systems – it is a global leader in the core research areas of:

•	Smart sensors and systems

•	Optical communication systems

•	Mixed signal and analog circuit design

•	Microelectronic and photonic integration

•	Semiconductor wafer fabrication

•	Nano materials and device processing

Central to our mission is delivering economic impact through research excellence. We work with industry and academia to transform research into products in our core market areas of electronics, communications, energy, health, agri-food and the environment. With a network of over 200 industry partners and customers worldwide, we are focused on delivering real impact from our excellent research. Our ambitious five-year strategic plan sets out a clear strategy to create employment and build critical mass within the Irish technology space.

Materials & Devices - CMOS++, Tyndall National Institute

The CMOS++ Materials & Devices research programme focuses on materials, emerging devices and architectures for future information processing, interfacing with CMOS and Beyond.

Our vision is to enable deep-tech innovation for next-generation and emerging information-processing technologies.

Our mission is to deliver state-of-the-art solutions for achieving:

- quantum advantage using solid-state platform technologies

- low-power, energy-efficient, high-performance devices and architectures

- increased functionality through novel materials integration.

These will advance the semiconductor industry and contribute to innovations for sustainable information processing applied to, for example, more efficient design of targeted drugs, personalising healthcare, improving logistics to protect natural resources and the environment and managing financial and personal risk.

Funding Acknowledgement

Tyndall-UCC acknowledge funding from Research Ireland under the Research Infrastructure Programme project: Advanced Semiconductor Processing and characterisation facility for multifunctional thin-film materials (ASPIRE) (25/RI/14077).

# Summary of Requirements

Tyndall National Institute, University College Cork requests the supply of a Precision Ferroelectric and Piezoelectric Analysis Platform, comprising of seven complementary subsystems with functionality requirements as follows:

- Measurement Core – Central electrical tester and control electronics for ferroelectric / piezoelectric /leakage tasks

- Optical Metrology - 	Sub-picometre displacement sensor

- Thermal Modules	 - Automated heating/cooling stages and controllers for bulk and thin films

- HV Subsystem	 - High-voltage amplifiers (up to ±10 kV) and safety/protection interfaces

- Fixturing & Probing	 - Automated wafer stages, bulk sample holders, and precision probe positioners

- Workstation	 - Pre-configured PC, high-resolution monitor(s), and full software suite

- Infrastructure - 	Vibration isolation (e.g., optical table) and environmental shielding (acoustic/light)

The Platform must have the ability to perform comprehensive, high-resolution ferroelectric and piezoelectric characterisation of ferroic materials, spanning bulk ceramics, ultrathin films and full wafers, across a broad temperature range.

The primary purpose of this platform is to provide a unified, high-performance infrastructure for the fundamental research and industrial benchmarking of advanced ferroic and piezoelectric materials. It will be utilised for the precision characterisation of material properties at various stages of development, from initial thin-film growth and bulk ceramic synthesis to full wafer-level device integration and reliability testing, supporting the development of next-generation sensors, actuators, and beyond-CMOS computing architectures.

Materials with controlled ferroic responses are key to enabling next‑generation technologies. Ferroelectric and piezoelectric materials are vital for precision actuators, sensors, photonic systems, and non‑volatile memories, while multiferroics, combining electric and magnetic order, present transformative potential for ultra‑low‑power, non‑volatile logic and memory.

Realising these technologies demands high‑performance measurement platforms capable of characterising domain dynamics, coercive fields, fatigue, and reliability, providing rapid feedback for material and device optimisation. Establishing this capability will accelerate innovation in energy‑efficient actuators, adaptive sensors, and beyond‑CMOS computing, maintaining Ireland’s leadership in materials‑enabled technology development.

At Tyndall, this infrastructure is critical for wafer‑level characterisation and mapping of ferroelectric and piezoelectric thin films, enabling process optimisation, improved device integration, and rapid prototyping of advanced microsystems. The investment will strengthen Ireland’s strategic position in emerging research areas such as neuromorphic and quantum computing, flexible and sustainable electronics, and the broader digital and green transitions, reinforcing Ireland’s status as a global hub for advanced materials research and innovation.

Budget Available

The Contracting Authority estimates that the expenditure on the Goods to be covered by the proposed Goods Contract is no greater than €435,000 ex VAT (NOT including Optional to Purchase Items)

This is to include all Costs (Capital Costs and associated Costs) covered within the total Ultimate Cost in Appendix 2.

Note: The Contracting Authority reserves the right to exclude submissions from suppliers which exceed the maximum budget available as stated above.

UCC will consider tenders offering new instrumentation. The Goods/Instrument(s) must be functional as new and supplied with a comprehensive mechanical and electrical warranty.

If multiple equipment types are available which suit the specifications required, please provide details of each option on a separate Tender Response Document and Appendix 2 Pricing Schedule, clearly defining each option.

UCC reserves the right to buy additional options available with this equipment at a later stage should the budget become available. This option will be available for 24 months.

# Section A: Selection Criteria – RFT Section 3.2

Tenderers will either Pass OR Fail each of the Selection Criteria. In the event of one or more of the Selection Criteria achieving a fail, the Tenderer will be excluded from participating in this Competition.

Economic and Financial Standing A1 – A4

## A1: Insurance

| Note – UCC reserves the right to determine whether the insurance evidence provided is sufficient to demonstrate compliance. Where certificates or broker letters do not clearly set out all required information – including limits, basis of cover, excesses, business description or indemnity to principal wording – the tender may be deemed non-compliant. UCC also reserves the right to request copies of relevant policy endorsements (e.g., indemnity to principal, additional insured wording) to verify compliance. |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Employer Liability |  | Public Liability |  | Product Liability |  |  | Professional Indemnity |  |
| Minimum Limit of Indemnity |  | €13,000,000 |  | €6,500,000 |  | €6,500,000 |  |  | €6,500,000 |  |
| I confirm that we have the following insurances in place: |  |  |  |  |  |  |  |  |  |  |
| Insurance Type |  |  | Level in Place |  | Details of Any Excess |  |  | Expiry Date |  |  |
| Employers Liability Required Level: €13 Million (on an each and every occurrence basis) - Euro or foreign currency equivalent |  |  | € |  |  |  |  |  |  |  |
| Public Liability Required Level: €6.5Million (on an each and every occurrence basis) - Euro or foreign currency equivalent |  |  | € |  |  |  |  |  |  |  |
| Product Liability Required Level: €6.5Million (on an aggregate basis) – Euro or foreign currency equivalent |  |  | € |  |  |  |  |  |  |  |
| Professional Indemnity Required Level: €6.5Million (each and every claim or in the annual aggregate) - Euro or foreign currency |  |  | € |  |  |  |  |  |  |  |
| AND |  |  |  |  |  |  |  |  |  |  |
| I confirm that if successful, where the levels required under the contract are higher than those currently in our possession, I will be in a position to put the required forms and levels of insurance required in place. |  |  |  |  |  |  | Yes/No |  |  |  |
| AND |  |  |  |  |  |  |  |  |  |  |
| I confirm that I will provide the following promptly on request: - evidence of insurance in place or - letter from Insurance Broker confirming that the required levels could be put in place if successful |  |  |  |  |  |  | Yes/No |  |  |  |
| Requirements for ALL insurance certificates: - Full legal name of insured entity and any subsidiaries involved in the contract. - Business description, confirming the policy covers the activities associated with the UCC contract. - Reference to the tender (e.g., “Tender Ref #– University College Cork”). - Policy details, including: - Policy type and number - Insurer name - Start and expiry dates - Limits of indemnity, specifying whether limits apply per occurrence or in the aggregate - Basis of cover (e.g., occurrence or claims‑made; retroactive date if claims‑made) - Excess / Deductible / SIR – explicit euro or foreign currency equivalent amounts must be stated - Territorial limits and jurisdiction - Confirmation of any contract‑specific requirements, including: Indemnity to Principal / Specific Indemnity to UCC |  |  |  |  |  |  |  |  |  |  |
| Signed: |  |  |  |  |  |  |  |  |  |  |
| Position: |  |  |  |  |  |  |  |  |  |  |

## A2: Tax Clearance

- Please provide a copy of a valid Irish Tax Clearance Certificate and complete the information below:

| Valid Tax Clearance Certificate: |  |
| --- | --- |
| Evidence Attached | Yes/No |
| Appendix Number |  |
| Expiry Date: |  |

(ii) Alternatively, please provide the following information to enable online verification.

| Tax Registration Number |  |
| --- | --- |
| Tax Clearance Access number |  |

(iii) Or - Please provide evidence that a Tax Clearance Certificate has been applied for from The Irish Revenue Commissioners

| Evidence Attached | Yes/No |
| --- | --- |
| Appendix Number |  |

Note: All non-resident companies MUST apply for a general TCC where tax clearance is required in connection with a Public-Sector Contract (formal or otherwise).

Both Resident and Non-Resident Tax Clearance Certificate information may be obtained from the Irish Revenue Commissioners http://www.revenue.ie/en/tax/it/forms/webtc1.pdf

## A3: Statement of Turnover

Applicants should provide evidence, from your company’s Auditor/Accountant, that their turnover relating to the services specified in Appendix 1 only, has exceeded €870,000 per annum averaged over the three years 2022/2023/2024.

| Lot | Yearly Turnover | 2022 | 2023 | 2024 |
| --- | --- | --- | --- | --- |
| 1 | €870,000 |  |  |  |

Please confirm below that this information has been submitted with your tender response and the Appendix number of the information provided:

| Confirm Yes/No | Yes/No |
| --- | --- |
| Appendix Number |  |

## A4: Audited Accounts & Reports

Three (3) years Audited accounts, from your company’s Auditor/Accountant, must be provided to support the turnover statement. In line with current legislation, where a company has an annual turnover of less than €12 Million then audited accounts are not necessary. In these instances, a Statement of Accounts from the company’s accountants will suffice. Please note: a link to audited accounts will not be accepted.

Please confirm below that this information has been submitted with your tender response and the Appendix number of the information provided:

| Audited Accounts, Yes/No | Yes/No |
| --- | --- |
| Statement of Accounts from Accountant (if turnover < €12M), YES/NO | Yes/No |
| Appendix Number |  |

Technical and Professional Ability A5-A9

## A5: Data Protection

The Tenderer must ensure that the security of Personal Data at all times is in line with the Data Protection Acts 1998 and 2003 and guidance issued by the Data Protection Commissioner of Ireland. Tenderers are required to be compliant under the EU General Data Protection Regulation (GDPR) replacing the Data Protection Directive 95/46/EC, which came into effect on 25/05/2018. Please provide a copy of your Data Protection Policy.

Tenderers must comply with the EU General Data Protection Regulation (GDPR), which replaces the Data Protection Directive 95/46/EC and came into effect on 25/05/2018.

Please provide a copy of your Data Protection Policy.

Please confirm below that this information has been submitted with your tender response and the Appendix number of the information provided:

| Confirm Yes/No | Yes/No |
| --- | --- |
| Appendix Number |  |

## A6: Quality Assurance & Accreditation

A certificate confirming an accreditation (e.g. accredited partnership) exists between the Instrument source manufacturer and the Tenderer for the proposed solution, confirming the Tenderer’s ability as authorised provider of the solution to the Contracting Authority.

Please confirm below that this information has been submitted with your tender response and the Appendix number of the information provided:

| Confirm Yes/No | Yes/No |
| --- | --- |
| Appendix Number |  |
| Expiry Date of Certification |  |
| We are the Source Manufacturer | Yes/No |
| Is it 3rd-party certified or in-house? | 3rd Party Certified/In-house |

Regulatory Compliance – tenderers must demonstrate adherence to all relevant Irish and European regulatory requirements, including but not limited to the following.

- Must be clearly CE marked.

If supplying through the UK, please specify the location of the awarding body of the CE mark, as since Brexit, all UK suppliers must provide evidence that the CE mark has been certified by an EU Notified Body. (if applicable).

| Confirm CE Compliance: Yes/No | Yes/No |
| --- | --- |

## A7: Organisational Summary

Please complete the following table(s) with your Company Summary.  If the Tenderer is a grouping, separate information must be completed for each group member.

| Tendering Party Name |  |
| --- | --- |
| Legal Structure – partnership, limited company, etc. |  |
| Please confirm if you are an SME (Small and Medium Enterprise) as defined in Commission Recommendation 2003/361/EC | YES / NO Definition as per 2003/361/EC The category of micro, small, and medium-sized enterprises (SMEs) comprises enterprises that employ fewer than 250 persons and have an annual turnover not exceeding EUR 50 million, and/or an annual balance sheet total not exceeding EUR 43 million. |
| Tendering Group Information – if relevant |  |
| Contact Name: |  |
| Address: |  |
| Landline telephone number: |  |
| Mobile number: |  |
| Out of hours contact telephone number: |  |
| Email address: |  |
| Please provide a brief company profile, including an organisation chart (Please expand this box as necessary) |  |

## A8: Previous Experience

Please provide details of three (3) contracts undertaken within the last five (5) years that demonstrate your experience of successfully providing a service of similar scope, scale and complexity to the service anticipated under this contract. Be very clear what your organisation delivered under this contract. Please demonstrate expertise and track record in the Installation and Implementation of a similar system or part thereof.

Please ensure permission has been sought before submitting reference information as these will be checked by the evaluation team.

Please confirm that this information has been completed and submitted with your tender.

| Confirm Yes/No | Yes/No |
| --- | --- |

| Previous Experience | Reference 1 |  |  |
| --- | --- | --- | --- |
| Title of Project: |  |  |  |
| Client Name & address: (Note: where name is confidential, please indicate nature / type of client and location) |  |  |  |
| Client contact person: |  | Phone no: |  |
| Contact E-mail Address: |  |  |  |
| Annual Value: |  |  |  |
| Any associated subcontractors |  |  |  |
| Start Date: |  | End Date: |  |
| Please describe this contract in full |  |  |  |
| Insert, expanding this table as necessary. |  |  |  |
| Please outline the extent to which this contract is similar in nature and scale to the Contracting Authority’s current requirements.. Insert, expanding this table as necessary. |

| Previous Experience | Reference 2 |  |  |
| --- | --- | --- | --- |
| Title of Project: |  |  |  |
| Client Name & address: (Note: where name is confidential, please indicate nature / type of client and location) |  |  |  |
| Client contact person: |  | Phone no: |  |
| Contact E-mail Address: |  |  |  |
| Annual Value: |  |  |  |
| Any associated subcontractors |  |  |  |
| Start Date: |  | End Date: |  |
| Please describe this contract in full |  |  |  |
| Insert, expanding this table as necessary. |  |  |  |
| Please outline the extent to which this contract is similar in nature and scale to the Contracting Authority’s current requirements.. Insert, expanding this table as necessary. |

| Previous Experience | Reference 3 |  |  |
| --- | --- | --- | --- |
| Title of Project: |  |  |  |
| Client Name & address: (Note: where name is confidential, please indicate nature / type of client and location) |  |  |  |
| Client contact person: |  | Phone no: |  |
| Contact E-mail Address: |  |  |  |
| Annual Value: |  |  |  |
| Any associated subcontractors |  |  |  |
| Start Date: |  | End Date: |  |
| Please describe this contract in full |  |  |  |
| Insert, expanding this table as necessary. |  |  |  |
| Please outline the extent to which this contract is similar in nature and scale to the Contracting Authority’s current requirements.. Insert, expanding this table as necessary. |

## A9: Health & Safety

Confirmation of Organisation’s Health and Safety statement must be provided.  All Tenderers should demonstrate any other relevant standards in relation to the supply of the tendered equipment.

Please confirm below that this information has been submitted with your tender response and the Appendix number of the information provided:

| Confirm Yes/No | Yes/No |
| --- | --- |
| Appendix Number |  |

## A10: Appendix 3 Tenderers’ Statement

[Tenderers shall complete and return the following form of Tenderers’ Statement printed on the Tenderers’ headed notepaper and signed by the Tenderer.]

TENDERERS’ STATEMENT

TO: University College Cork (the “Contracting Authority”)

RE: Request for Tenders for the Supply, Delivery, Installation, Commissioning and Calibration of a Ferroelectric and Piezoelectric Analysis Platform, for Tyndall National Institute, University College Cork (UCC)

Having examined your Request for Tenders (the “RFT”) including the Instructions to Tenderers, the Selection and Award Criteria, the Requirements and Specifications, and the Terms and Conditions of the Goods Contract, we hereby agree and declare the following:

| 1. | We understand the nature and extent of the Goods required to be delivered as described in Requirements and Specifications at Appendix 1 to the RFT. |
| --- | --- |
| 2. | We accept all of the Terms and Conditions of the RFT, the Goods Contract, and the Confidentiality Agreement and agree, if awarded a Goods Contract, to execute the Goods Contract at Appendix 5 to the RFT and the Confidentiality Agreement at Appendix 6 to the RFT. |
| 3. | We accept all the Selection and Award Criteria as set out in Part 3 of the RFT. |
| 4. | We agree to supply the Contracting Authority with the Goods in accordance with the RFT and our Tender. |
| 5. | We agree that, if awarded any Goods Contract we shall, in the performance of such contract, comply with all applicable obligations in the field of environmental, social and labour law. |
| 6. | We confirm that we have complied with all requirements as set out at Part 2 of the RFT. |
| 7. | We confirm that all prices quoted in our Tender will remain valid for the period of time, commencing from the Tender Deadline, specified at paragraph 2.10.3 of the RFT. |
| 8. | We shall, if awarded any Goods Contract under the RFT, have in place on the Effective Date of the Goods Contract all insurances (if any) as required by paragraph 2.21.1 of the RFT. |
| 9. | We confirm that all Data Subjects whose Personal Data is provided in our Tender have consented to the processing of such Personal Data by us, the Contracting Authority, the Evaluation Team and the supplier of the etenders.gov.ie website, for the purposes of our participation in this Competition or that we otherwise have a legal basis for providing such Personal Data to the Contracting Authority for the purposes of our participation in this Competition and that we will provide evidence of such consent and / or legal basis to the Contracting Authority upon request. |
| 10. | We do not come within the category of prohibited economic operators identified in Regulation (EU) No 833/2014 of 31 July 2014 (as amended by EU Regulation 2022/576 or any subsequent amendments to same). |
| 11. | The origin of goods connected to our Tender, if any, are not subject to the prohibitions set out in Regulation (EU) No 833/2014 (as amended by EU Regulation 2022/576 or any subsequent amendments to same). |
| 12. | The subcontractor(s) on whose capacity we rely as part of our Tender (where the value of that subcontract exceeds 10% of the value of the Services Contract) does not come within the category of prohibited economic operators identified in Regulation (EU) No 833/2014 of 31 July 2014 (as amended by EU Regulation 2022/576 or any subsequent amendments to same). |

| SIGNED (Authorised Signatory) | Company |
| --- | --- |
| Print name | Address |
| Date |  |

[

## A11: Appendix 4 Declaration as to Personal Circumstances of Tenderer

Re: LA3673C - Request-for-Tenders under Open (OJEU) Procedure for the Supply, Delivery, Installation, Commissioning & Calibration of a Fully Integrated Precision Ferroelectric and Piezoelectric Analysis Platform, for Tyndall National Institute, University College Cork (UCC)

NAME:   	[Click here and insert name]

ADDRESS: 	[Click here and insert address]

I, [Click here and insert name of Declarant], of [Click here and insert name of entity]  do solemnly and sincerely declare that:

- I am a [insert role of Declarant] of [Click here and insert name of entity] and am authorized by [Click here and insert name of entity]  to make this declaration which relates to a tender (“the Tender”) submitted by [Click here and insert name of entity]  in response to an RFT dated titled [insert description of competition] published by [insert name of contracting authority] (“the Contracting Authority”).

- Neither [Click here and insert name of entity]  nor any person who is a member of the administrative, management or supervisory body of [Click here and insert name of entity]  nor any person who has powers of representation, decision or control in [Click here and insert name of entity]  has:

- ever been the subject of a conviction for participation in a criminal organisation, as defined in Article 2 of Council Framework Decision 2008/841/JHA.

- ever been the subject of a conviction for corruption, as defined in Article 3 of the Convention on the fight against corruption involving officials of the European Communities or officials of Member States of the European Union and Article 2(1) of Council Framework Decision 2003/568/JHA as well as corruption as defined in the national law of the Contracting Authority or the law of the state in [Click here and insert name of entity] is established.

- ever been the subject of a conviction for fraud within the meaning of Article 1 of the Convention on the protection of the European Communities’ financial interests.

- ever been the subject of a conviction for terrorist offences or offences linked to terrorist activities, as defined in Articles 1 and 3 of Council Framework Decision 2002/475/JHA respectively, or for inciting or aiding or abetting or attempting to commit an offence, as referred to in Article 4 of that Framework Decision.

- ever been the subject of a conviction for money laundering or terrorist financing, as defined in Article 1 of Directive 2005/60/EC of the European Parliament and of the Council.

- ever been the subject of a conviction for child labour and other forms of trafficking in human beings as defined in Article 2 of Directive 2011/36/EU of the European Parliament and of the Council.

- [Click here and insert name of entity]:

- is not in breach and has not breached its obligations relating to the payment of taxes or social security contributions.

- has carried out the preparation of the Tender independently.

- [Click here and insert name of entity]:

- has, in the performance of all public contracts, complied with applicable obligations in the field of environmental social and labour law that apply at the place where the works are carried out or the services provided, that have been established by EU law, national law, collective agreements or by international, environmental, social and labour law listed in Schedule 7 of the European Union (Award of Public Authority Contracts) Regulations 2016 (Statutory Instrument 284 of 2016).

- is not bankrupt or the subject of insolvency or winding-up proceedings, its assets are not being administered by a liquidator or by the court, it is not in an arrangement with creditors, its business activities are not suspended nor is it in any analogous situation arising from a similar procedure under national laws and regulations.

- is not guilty of grave professional misconduct.

- has not entered into agreements with other economic operators aimed at distorting competition.

- is not aware of any conflict of interest due to its participation in the Competition;

- has not had any prior involvement in the preparation of the Competition;

- has not shown significant or persistent deficiencies in the performance of a substantive requirement under a prior public contract, a prior contract with a contracting entity, or a prior concession contract, which led to early termination of that prior contract, damages or other comparable sanctions.

- is not guilty of serious misrepresentation in supplying the information required for the verification of the absence of grounds for exclusion or the fulfilment of the Selection Criteria for this Competition and did not withhold such information and did not fail or is not able to submit supporting documents in respect of this Competition as required under Regulation 59 of the European Union (Award of Public Authority Contracts) Regulations 2016 (Statutory Instrument 284 of 2016).

- has not undertaken to unduly influence the decision-making process of the Contracting Authority in respect of the Competition, or obtain confidential information that may confer upon it undue advantages in respect of the Competition; or negligently provided misleading information that may have a material influence on decisions concerning exclusion, selection or award.

- [Click here and insert name of entity] does not come within the category of prohibited economic operators identified in Regulation (EU) No 833/2014 of 31 July 2014 (as amended by EU Regulation 2022/576 or any subsequent amendments to same);

- The origin of goods connected to the Tender, if any, are not subject to the prohibitions set out in Regulation (EU) No 833/2014 (as amended by EU Regulation 2022/576 or any subsequent amendments to same);

- Any subcontractor, supplier or other entity on whose capacity [Click here and insert name of entity]  relies as part of the Tender does not come within the category of prohibited economic operators identified in Regulation (EU) No 833/2014 of 31 July 2014 (as amended by EU Regulation 2022/576 or any subsequent amendments to same).

I understand and acknowledge that the provision of inaccurate or misleading information in this declaration may lead to my business/firm/company/partnership being excluded from participation in this or future tenders, and I make this solemn declaration conscientiously believing the same to be true and by virtue of the Statutory Declarations Act, 1938. This declaration is made for the benefit of the Contracting Authority.

________________________

Signature of Declarant

Declared before me by ___________________________________ who is personally known to me

(or who is identified to me by ______________________________who is personally known to me) or*

at ____________________________  this ___________ day of _______________ 20__

________________________________

(signed)

Practicing Solicitor/Commissioner for Oaths

*Please include such other form of identification used to identify the Declarant as permitted by the Statutory Declarations Act, 1938 (as amended)

## A12: Appendix 5 Declaration and Warranty on the Use of Artificial Intelligence in the Preparation of the Tenderer’s Response

The Contracting Authority is aware that Large Language Models and Generative Artificial Intelligence (collectively “AI Tools”) are being used extensively in the preparation of responses to tenders in both the Public and Private Sectors. In responding to this Competition, Tenderers are required to declare if and where AI Tools have been used to draft, develop, inform or augment any part of their Responses.  While the use of AI Tools in responding to this Competition is not prohibited, it is necessary that Tenderers maintain control of and responsibility for their Response and that they confirm as much to the Contracting Authority.

It should be noted that the evaluation team will assess Responses to this Tender based on the information provided by the Tenderer regardless of how the Response was generated. Any declarations made here on the use, or non-use, of AI Tools will not influence the evaluation process or outcome.

However, should the Contracting Authority determine that any declaration made in this regard by the Tenderer is incomplete or has been incorrectly or inaccurately completed then the Tenderer’s Response may, subject to the provisions in Part A – Introduction to Tender, Instructions to Tenderers, be deemed to be non-compliant and will not be evaluated.

| Tenderers are required to provide responses to the following question: |  |
| --- | --- |
| - Have AI Tools been used in the preparation of any part of this Tender response? | Yes/No |
| If you answered 'yes' to the previous question, Tenderers are required to provide responses to the following additional questions: |  |
| - Pease provide details regarding the use of AI Tools (e.g., drafting text, data analysis, generating images, proofreading): |  |
| Response |  |
| - Describe the measures taken to ensure human oversight and validation of any AI-generated content in your Tender response. |  |
| Response |  |
| - If AI Tools were used for data analysis or generating content, identify the sources of data used and explain how you ensured the data's accuracy, reliability, and ethical use. |  |
| Response |  |
| - The Tenderer confirms that all elements of the Response generated or informed by AI Tools have been validated and the Tenderer warrants that, if it is successful in this Competition, it can deliver all elements of the Response to the standard and in the way described in the Response at the proposed price. | Yes/No |

| Tenders MUST detail below where and why AI was used to generate your response. Note: All AI generated information MUST be cross checked to ensure alignment with the delivery of goods and costs etc. |  |  |  |
| --- | --- | --- | --- |
| RFT Section | RFT Documentation AI Analysed | RFT Full / Partial AI Generated Response | Purpose of AI Use |
| RFT |  |  |  |
| Print Audit |  |  |  |
| Tender Response Document |  |  |  |
| Appendix 2 Pricing Schedule |  |  |  |

# Section B: Tender Response

## B1: Requirements & Specifications

Suppliers MUST provide all information relating to the instrument and the associated consumables in the response boxes below. Responses will be scored accordingly by the Technical Evaluation team as set out in the Award Criteria in the RFT document.

Please ensure all cost submissions are kept separate and included in the Excel file entitled: Appendix 2 – Pricing Schedule.

This is the minimum specification required.  All tenderers must meet or exceed this minimum specification to progress through the award criteria to cost evaluation.  Additional functionality should be identified in the comments section to facilitate scoring under the award criteria. The Contracting Authority may award a higher marking if the additional functionality is desirable or adds value.

Supplier Notes

- This specification document is designed to assist the contracting authority in its evaluation of your tender and to enable suppliers provide structured responses to key deliverables. It is structured to reflect the award criteria as set out in the RFT.

- Please note responses to individual questions will NOT be scored separately but collectively they will provide a basis for assessing the merits of the tenderer’s overall response to each award criterion and sub-criterion.

- Where specifications include reference to manufacturer or brand names this is a result of difficulties in specifying the product required. In these instances the phrase “or equivalent” is inferred.

- It is acknowledged that each Tender’s product line may differ from these specifications.  As such, each seller is free to propose variances from these specifications. Alternative proposals favourable to the Contracting Authority may be suggested. It is required, however, that, whenever a variance from these specifications occurs, the proposed item must meet or exceed the specified characteristics or level of performance.

- Where abnormally low tenders are received the Evaluation team shall require the tenderer to explain the price proposed in Appendix 2. (Per Article 69 of the EU Procurement Directives 2014/24). Where the tenderer cannot provide a sufficient explanation, the contracting authority is entitled to reject the tender.

- The award of any contract is subject to the successful tenderer meeting the required specification and standards as set out in this RFT and Appendix 1. Should the required specification and standards not be met, the EPS reserves the right, in consultation with the Contracting Authority, to award any contract to the next highest-ranked Tenderer.

- If the evaluation outcome results in a tie between two or more Tenders, then the Tender with the highest score for the highest weighted criterion shall be deemed the most economically advantageous tender. If the score remains in a tie, the remaining criterion will be assessed by the highest weighted until there is no longer a tie.  If the score remains a tie after all criterions are assessed, the tenderers will be asked to submit a Best and Final Offer (BAFO).

A failure to answer question(s) may leave the contracting authority with insufficient information on which to base a thorough evaluation and in turn may result in loss of marks in respect of the relevant Award Criterion. Therefore, please ensure you complete ALL sections of this questionnaire.

Only those Tenderers who have qualified in accordance with paragraph 3.2 of this RFT will proceed to be evaluated under this paragraph 3.3.1

## B2: Award Criteria

The Goods Contract will be awarded on the basis of the most economically advantageous tender(s) as identified in accordance with the following award criteria:

| Award Criteria LA3673C – Request-for-Tenders under Open (OJEU) Procedure for the Supply, Delivery, Installation, Commissioning & Calibration of a Fully Integrated Precision Ferroelectric and Piezoelectric Analysis Platform, for Tyndall National Institute, University College Cork (UCC) |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | Award Criteria: Please Note: All sub criteria for sections B, C, D & E carry equal weight/marks. For section A some sub criteria carry its own weighting. | Weighting | Total Maximum Score | Minimum Score (60%) Requirement |
| - Technical Merit |  |  |  |  |
| A | - General Functional Requirements - Performance Requirements - Hardware & Software Requirements | 30% 5% 15% 10% | 300 50 150 100 | 30 90 60 |
| - Quality, Health & Safety & Validation |  |  |  |  |
| B | - Quality Requirements - Health & Safety Requirements | 5% | 50 | 30 |
| - Warranty/Spare Parts, After-Sales and Training Support |  |  |  |  |
| C | - Warranty & Customer Support, - After-Sales, Repair & Breakdown Support - Replacement and Spare Parts & Equipment Consumables - Training & Technical Support Proposal | 10% | 100 | 60 |
| - Delivery Plan |  |  |  |  |
| D | - Lead Time & Delivery - Installation, Commissioning, Calibration Requirements - Documentation | 10% | 100 | 60 |
| - Energy, Environmental & Sustainable Considerations |  |  |  |  |
| E | - Environmental & Sustainable Considerations | 5% | 50 | 30 |
| - Ultimate Cost |  |  |  |  |
| F | Ultimate Cost Including: (this includes the Total Purchase Cost/Capital Cost (ex. VAT) and the After Sales Service & Warranty Costs (Service Cost), exc. VAT as per Appendix 2 – Pricing Schedule | 40% | 400 | No Minimum Score Applies to Ultimate Cost |
|  | - Total Maximum Score Available | 100% | 1000 |  |

Tenderers must achieve a minimum of 60% on all tender award criteria above, excluding Ultimate Cost, in order to avoid elimination from the competition. Only tenderers who have achieved the 60% will be evaluated against Ultimate Cost.

No commitment of any kind, contractual or otherwise shall exist unless and until a formal written contract has been executed by or on behalf of the Contracting Authority. Any award of notification of preferred bidder status by the Contracting Authority shall not give rise to any enforceable rights by the Tenderer. The Contracting Authority may cancel this public procurement competition at any time prior to a formal written contract being executed by or on behalf of the Contracting Authority. The Contracting Authority does not bind itself to accept the lowest priced or any tender.

The award of contract will be on the basis of the Most Economically Advantageous Tender received, in accordance with the detailed award criteria and weightings set out below.

The lowest Ultimate Cost tender, will receive the maximum score achievable under this criterion.

| Pricing Score = | Lowest Tendered Price x Max Points |
| --- | --- |

Price being evaluated.

Responses for non-cost/qualitative Award Criteria items will be scored using the following matrix:

| Scoring Methodology |  |
| --- | --- |
| Scoring Range | Meaning |
| 91% – 100% | Excellent response that fully meets and exceeds requirements, providing comprehensive, detailed, and convincing assurance that the Tenderer will deliver to an exceptional standard. |
| 80% – 90% | A very good response that demonstrates real understanding and fully meets the requirements and provides convincing assurance that the Tenderer will deliver to high standard. |
| 60% – 79% | A satisfactory response which demonstrates a reasonable understanding of requirements and provides reasonable confidence that the minimum requirements will be delivered. However, it does not provide sufficiently convincing assurance to award a higher mark. |
| 30% - 59% | A response where reservations exist due to a lack of convincing details and comprehensive credibility. There is a significant risk that the requirements will not be delivered to a satisfactory standard. |
| 1% – 29% | A response that raises significant reservations, due to insufficient detail, critical or notable deficiencies, that seriously lacks credibility, presenting a high risk of non-delivery. |
| 0% | No Response or partial Response that completely fails to address the criterion under consideration |

## B3: Appendix 1-Technical Specification

| Product Description | Tendering (Y/N) |
| --- | --- |
| Fully Integrated Precision Ferroelectric and Piezoelectric Analysis Platform |  |

Suppliers MUST provide all information relating to the instrument and the associated consumables in the blue response boxes below. Responses will be scored accordingly by the Evaluation team as set out in the Award Criteria in the RFT document.

Where an additional document is required to support your response, the document MUST be attached and the relevant section within the document MUST be clearly referenced in your answer.

Please note that website links are not sufficient as a supporting document.

Only referenced documents will form part of the evaluation. Brochures that have not been referenced will NOT be evaluated.

Please ensure all cost submissions are kept separate and included in the Excel file entitled:  Appendix 2 – Pricing Schedule

| LA3673C –Supply, Delivery, Installation, Commissioning and Calibration of a Fully Integrated Precision Ferroelectric and Piezoelectric Analysis Platform, for Tyndall National Institute, University College Cork (UCC) |  |
| --- | --- |
| Platform Sub Systems | State proposed model name, or Module description here: |
| Central electrical tester and control electronics for ferroelectric/piezoelectric/leakage tasks |  |
| Sub-picometre displacement sensor |  |
| Automated heating/cooling stages and controllers for bulk and thin films |  |
| High-voltage amplifiers (up to ±10 kV) and safety/protection interfaces |  |
| Automated wafer stages, bulk sample holders, and precision probe positioners |  |
| Pre-configured PC, high-resolution monitor(s), and full software suite |  |
| Vibration isolation (e.g., optical table) and environmental shielding (acoustic/light) |  |

- Technical Merit

General Functional Requirements

|  | The Tenderer MUST supply a NEW fully integrated Precision Ferroelectric and Piezoelectric Analysis Platform. The Precision Ferroelectric and Piezoelectric Analysis Platform MUST be supplied as a fully integrated turnkey solution comprising the following seven mandatory physical subsystems: - Precision Electrical Characterisation Core: A high-speed central unit for all ferroelectric, piezoelectric, leakage, and pulse-based measurements, providing integrated signal generation and high-speed data acquisition for both electrical and electromechanical characterisation. - Sub-Picometre Optical Metrology Subsystem: A non-contact optical displacement sensor (e.g., laser-based) capable of meeting the stated picometre-level resolution (<1 pm) and accuracy requirements. - Automated Thermal and Environmental Control Modules: Dedicated thermal stages and controllers for both bulk and thin-film samples, capable of reaching the specified low-temperature and high-temperature limits (e.g. -100°C to 600°C for bulk samples, and -100°C to 500°C for thin-film samples). - High-Voltage (HV) Expansion & Safety Subsystem: A safety-interlocked amplification system extending the operational range to at least ±10 kV. - Multi-Scale Sample Fixturing & Probing: A modular set of sample holders and automated stages to accommodate full wafers (at least 100 mm) and bulk ceramics (at least 25 mm diameter). This subsystem MUST include integrated electrical contacting hardware—such as precision micro-positioners and conductive chucks—to establish reliable electrical connections with user-fabricated electrodes. - Control & Analysis Workstation: A dedicated, pre-configured computer system with high-resolution monitor(s) and an integrated software environment for measurement execution, data acquisition, and data export. - Performance-Critical Infrastructure: All necessary vibration isolation (active or passive) and environmental shielding (e.g. acoustic/light) required to achieve the system's guaranteed performance in a standard laboratory setting. Please confirm and detail update the Appendix 2 pricing schedule with the relevant pricing. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |
| --- | --- |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | The platform MUST support dedicated and clearly defined measurement configurations for: - electrical ferroelectric characterisation of bulk samples, - electrical ferroelectric characterisation of thin-film samples, - piezoelectric characterisation of bulk samples, - piezoelectric characterisation of clamped thin-film structures on substrates and wafer-level samples by interferometric (or equivalent optical) displacement measurement. Please confirm and detail |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Platform Footprint Tenderers MUST meet the total footprint and access clearance requirements for the platform, which MUST NOT exceed 3.3 m (Width) x 2.0 m (Depth); including 0.8 m clearance. The total height of any system component MUST NOT exceed 2.0 m. Tenderers MUST explicitly state the dimensions a Width x Depth x Height (W x D x H) in meters format. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
| Performance Requirements |  |
|  | Bulk Sample Characterisation - temperature control The platform MUST support bulk ceramic and related samples for electrical ferroelectric and piezoelectric characterisation by means of a dedicated bulk sample holder with temperature control. Bulk samples are defined as macroscopic, discrete, and self-supporting ferroelectric or piezoelectric materials, including but not limited to: ceramic disks, single crystals, multilayer actuators, and large-area capacitors. (Specific sample dimensions and holder capacities are detailed in Ref A6). Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Bulk Sample Characterisation - high-voltage The platform MUST support bulk ceramic and related samples for electrical ferroelectric and piezoelectric characterisation by means of a dedicated bulk sample holder with high-voltage compatibility. Bulk samples are defined as macroscopic, discrete, and self-supporting ferroelectric or piezoelectric materials, including but not limited to: ceramic disks, single crystals, multilayer actuators (MLAs), and large-area capacitors. (Specific sample dimensions and holder capacities are detailed in Ref A6). Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Bulk Sample Characterisation – Sample Size The bulk measurement configuration MUST accommodate samples with a diameter of up to 25 mm and support thickness range spanning at least 0.2 mm to 3.0 mm. Tenderers MUST specify the following technical limits for the proposed hardware: - The maximum supported diameter (in mm). - The guaranteed thickness range (min/max in mm) that ensures full electrical contact and mechanical stability. - Any dependencies between sample diameter and allowable thickness. - Supported electrode configuration: Tenderers MUST confirm if the hardware supports both: a) Vertical/Sandwich structures (electrodes on top and bottom sides); b) In-plane/Planar structures (multiple electrodes on a single side, e.g., interdigitated electrodes) Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Bulk Sample Characterisation – Temperature Range The platform MUST support automated temperature-dependent piezoelectric and ferroelectric measurements on bulk samples over a range from -100°C to at least 600°C. Extended high temperature capability will be scored higher. Tenderers MUST state the full supported temperature range for the supplied bulk configuration. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Bulk Sample Characterisation – Voltage Range The integrated bulk sample configuration MUST support electrical characterisation using drive voltages of up to at least ±10 kV. The system MUST include all necessary; high-voltage interfaces, amplifiers, and safety-interlocked sample holders required to achieve this range. Tenderers MUST detail the current compliance and safety mechanisms (e.g., arc suppression) active at the 10 kV limit. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Thin Film Characterisation: General The platform MUST support electrical ferroelectric characterisation of thin-film and clamped thin-film samples by means of a dedicated temperature-controlled thin-film sample holder with microscope/camera-assisted probe positioning and high-voltage compatible contacting. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Thin Film Characterisation – Sample Size The thin-film electrical configuration MUST support discrete substrate-mounted thin-film samples. These are defined as individual, diced chips (e.g., 10 mm x 10 mm silicon or sapphire substrates) featuring ferroic thin films. Tenderers MUST state the exact supported sample formats and the minimum/maximum substrate dimensions compatible with the thin-film sample holders. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Thin Film Characterisation – Electrode Formats The system MUST accommodate both vertical (sandwich) and planar (interdigitated) electrode formats for the characterisation of ferroic thin films. This MUST be achieved using a minimum of two (2) high-precision probe positioners. Tenderers MUST state the exact supported sample formats and the minimum/maximum substrate dimensions compatible with the thin-film sample holders and probe station. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Thin Film Characterisation - Temperature Range: The platform MUST support thin-film electrical ferroelectric characterisation from at least -100°C to 500°C. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Clamped Thin-Film Piezoelectric Characterisation - General The platform MUST support piezoelectric characterisation (d33 measurements) of clamped thin-film structures on substrates by optical interferometric (or equivalent) displacement measurement with suppression of substrate bending artefacts. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Clamped Thin-Film Piezoelectric Characterisation For wafer-level characterisation, the platform MUST include a dedicated interferometric subsystem integrated with an automatic x-y stage, capable of performing automated, software driven measurements across a range of sample sizes, from small discrete substrate pieces (e.g., 10 mm x 10 mm) up to full wafers, with a diameter of at least 100 mm (4 inches). Proposals offering stages that support larger wafer sizes (e.g., 150 mm or 200 mm) will be scored higher. Tenderers MUST state the maximum travel range of the automatic stage and the maximum wafer diameter supported. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Clamped Thin-Film Piezoelectric Characterisation For clamped thin-film d33 measurements, the measurement principle MUST enable exclusion of substrate bending effects from the measured signal (e.g., via dual-side access, algorithmic correction, or equivalent). Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Clamped Thin-Film Piezoelectric Characterisation The tenderer MUST provide a comprehensive technical description of the proposed methodology for clamped thin-film characterisation. The response MUST include: - A detailed description of the physical measurement principle used to resolve sub-nanometre material displacements - Definition of the required sample architecture, including: - Electrode accessibility requirements for both vertical (sandwich) and planar (interdigitated) structures. - Optical requirements, such as the necessity for backside access or specific surface reflectivity. - The supported substrate thickness range compatible with the holder and optical path. - Confirmation and detail of the bending suppression method used to effectively isolate the intrinsic material strain from parasitic substrate curvature or environmental vibrations. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Clamped Thin-Film Piezoelectric Characterisation The tenderer MUST clearly distinguish their proposed hardware and methods for: - Electrical and ferroelectric thin-film measurements, - Local optical piezoelectric measurements on discrete samples or MEMS-like structures, - Wafer-level interferometric (or equivalent) measurements on clamped thin-film structures. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | ELECTRICAL FERROELECTRIC CHARACTERISATION: Dynamic Hysteresis Measurements: The platform MUST support dynamic hysteresis measurements for both bulk and thin-film samples. Hysteresis Frequency Requirements: - Mandatory Lower Limit: The system MUST be capable of operating at frequencies as low as 1 Hz (or lower) to support quasi-static material characterisation. - Mandatory Upper Limit: The system MUST support a dynamic frequency to an upper limit of at least 250 kHz. - Performance Bonus: Systems offering an expanded range (e.g., lower than 0.05 Hz or higher than 1 MHz) are desirable and will score higher. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | ELECTRICAL FERROELECTRIC CHARACTERISATION: Dynamic Hysteresis Measurements: The platform MUST meet the following minimum frequency and electrical performance requirements: - Internal amplifier frequency range: The system MUST be capable of a continuous frequency range from at least 1 Hz (lower) to at least 250 kHz (upper). - External high-voltage amplifier range: Tenderers MUST state the achievable frequency range when used with the required ±10 kV amplifier. - Operational Constraints: Tenderers MUST detail any operational limitations arising from sample capacitance, amplifier loading, or current compliance limits. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | ELECTRICAL FERROELECTRIC CHARACTERISATION: Static / Quasi-Static Hysteresis Measurements The platform MUST facilitate: - static or quasi-static hysteresis measurements suitable for minimising leakage. - relaxation artefacts and support dedicated static hysteresis measurement routines. - automated data analysis. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | ELECTRICAL FERROELECTRIC CHARACTERISATION: Leakage Current Measurements The platform MUST support highly sensitive leakage current measurements on both bulk and thin-film samples with a minimum current sensitivity of at least 1 pA or better. Tenderers MUST provide details of: - current range, - current resolution, - current measurement accuracy, - supported leakage measurement modes. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | ELECTRICAL FERROELECTRIC CHARACTERISATION: Pulse and PUND-Type Measurements The platform MUST support pulse-based ferroelectric measurements, including PUND-type pulse sequences. Shorter pulse widths will be scored higher. Tenderers MUST specify: - write pulse limits, - read pulse limits, - pulse amplitude capability, - timing parameter control, - whether acquisition is available during the shortest pulse widths. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | ELECTRICAL FERROELECTRIC CHARACTERISATION: Capacitive-Voltage Measurements: The platform MUST support capacitance-versus-voltage (C-V) measurements with adjustable DC bias and variable small-signal frequency (spanning at least 1 Hz to 1 MHz). Tenderers MUST state the exact frequency resolution and maximum DC bias supported for C-V measurements. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | ELECTRICAL FERROELECTRIC CHARACTERISATION: Capacitive-Voltage Measurements: The system MUST provide measurement outputs including capacitance and dielectric loss-related quantities such as tan(delta) and/or phase angle. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | ELECTRICAL FERROELECTRIC CHARACTERISATION: Capacitive-Voltage Measurements: Where a compatible piezoelectric displacement subsystem is supplied, simultaneous acquisition of small-signal displacement coefficient during C-V measurement SHOULD be available. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | ELECTRICAL FERROELECTRIC CHARACTERISATION: Fatigue Measurements: The platform MUST support high-frequency fatigue pulse cycling. Tenderers MUST state: - The absolute maximum fatigue pulse frequency supported by the system's internal generator. - The maximum output current (in Amps or mA) and typical voltage slew rate (e.g., V/µs) of the driving unit/amplifier at these high frequencies. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | ELECTRICAL FERROELECTRIC CHARACTERISATION: Retention Measurements: The tool MUST support retention measurements with a minimum applied voltage of ±10 V. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | ELECTRICAL FERROELECTRIC CHARACTERISATION: Retention Measurements: The system SHOULD accommodate an extended drive voltage input of up to ±10 kV and allow retention testing under two conditions: - shorted electrodes - continuous application of a retention bias. Please confirm and detail. Also state if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | ELECTRICAL FERROELECTRIC CHARACTERISATION: Imprint Measurements The tool MUST support automated imprint measurements and treatment sequences for Imprint measurements to evaluate the effect of applied voltage or temperature treatments on the shift of the hysteresis loop in thin-film and bulk ferroelectric systems. The vendor MUST state, with respect to thin film and bulk samples: - The minimum charge resolution that can be measured - The temperature range imprint measurements can be performed at - The maximum DC voltage that can be applied in Imprint measurements Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | ELECTRICAL FERROELECTRIC CHARACTERISATION: Imprint Measurements The system SHOULD possess a compensation method to differentiate parasitic switching from material switching currents during imprint measurements. Tenderers SHOULD confirm if this capability is included as a standard feature of the supplied software suite. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | ELECTRICAL FERROELECTRIC CHARACTERISATION: High Voltage Measurements: The platform MUST support high-voltage electrical ferroelectric characterisation up to at least ±10 kV and with at least 2mA current for suitable sample holders and sample geometries (scalable based on configuration). Tenderers MUST state for each amplifier and holder combination: - maximum voltage, - maximum current, - compatible measurement modes, - compatible sample environments, Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | BULK PIEZOELECTRIC CHARACTERISATION: Piezoelectric Characterisation of Bulk Samples - General The platform MUST support piezoelectric characterisation of bulk samples by simultaneous electrical excitation and optical displacement measurement and enable simultaneous acquisition of displacement and polarisation/current during compatible measurement sequences. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | BULK PIEZOELECTRIC CHARACTERISATION: Piezoelectric Characterisation of Bulk Samples – Longitudinal Response: The integrated platform MUST be capable of measuring the longitudinal piezoelectric coefficient (d33) on bulk samples across the following mandatory range: - For temperatures ranging from -100°C RT to at least 600 °C - with voltages up to ±10 kV Tenderers MUST state: - the measurement method, - displacement sensor type, - displacement resolution, reproducibility - voltage range, - temperature range, - the specific hardware configuration provided to reach the 10 kV / 600°C limit - achievable data outputs such as displacement versus voltage and effective large-signal d33. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | BULK PIEZOELECTRIC CHARACTERISATION: Pyroelectric Characterisation of Bulk Samples The platform SHOULD support pyroelectric current measurements on bulk samples under temperature ramp conditions (from -100°C to 600°C), where compatible sample holders and software are provided. Tenderers SHOULD provide the following information for evaluation: - Measurement Method: Detail the method used to extract pyroelectric coefficients (e.g., Byer-Roundy method or equivalent). - Configuration: Define the required sample configuration, including electrode type, minimum/maximum sample thickness, and diameter. - Sensitivity: State the minimum detectable pyroelectric current (must be 1 pA or better to align with standard leakage specs). Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | THIN-FILM PIEZOELECTRIC CHARACTERISATION ON CLAMPED STRUCTURES: Measurement Principle: For clamped thin-film structures on substrates, and wafer-level measurements, the platform MUST provide a dedicated optical interferometric (or equivalent) measurement method suitable for measuring electrically induced thickness displacement while suppressing substrate bending artefacts. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | THIN-FILM PIEZOELECTRIC CHARACTERISATION ON CLAMPED STRUCTURES: Effective d33 Measurement on Clamped Thin Films The platform MUST be capable of measuring the effective longitudinal piezoelectric coefficient of clamped thin-film capacitor structures on substrates and wafers. Tenderers MUST complete Appendix A. d33 Measurement on Clamped Thin Films. to define the performance and requirements of their proposed thin-film piezoelectric characterisation subsystem. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | THIN-FILM PIEZOELECTRIC CHARACTERISATION ON CLAMPED STRUCTURES: Bending Suppression Requirement: The thin-film piezoelectric measurement system MUST provide a methodology to ensure that reported displacement data represents the intrinsic electromechanical response of the material, effectively isolating or compensating for parasitic contributions such as mechanical substrate bending or ambient environmental vibrations. Tenderers MUST describe the technical approach used to achieve this and define the calculation method for the reported effective thin-film d33,f result. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | THIN-FILM PIEZOELECTRIC CHARACTERISATION ON CLAMPED STRUCTURES: Wafer-Level Mapping The platform MUST support automated wafer-level mapping for thin-film piezoelectric measurements, including motorised positioning and software-assisted mapping of parameters such as effective d33 and related electrical properties across the wafer. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | THIN-FILM PIEZOELECTRIC CHARACTERISATION ON CLAMPED STRUCTURES: Displacement Resolution and Repeatability The integrated optical metrology subsystem (utilising interferometric principles, such as laser doppler vibrometry, laser interferometry, or equivalent) MUST provide sub-nanometre displacement sensitivity. Tenderers MUST provide: - guaranteed displacement resolution better than 1 pm, - repeatability specification (e.g., ≤5% sigma on d33,av of quartz or equivalent reference sample (e.g. calibrated PZT thin films or calibrated AlN MEMS)), - reproducibility specification for measurements conducted across different sessions or sample loadings, - reference sample or test structure and test methodology used for acceptance testing. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | MEASUREMENT OF TRANSVERSE THIN-FILM PIEZOELECTRIC COEFFICIENTS: Piezoelectric Characterisation – Thin Film Transverse Measurements The platform MUST support the measurement of transverse piezoelectric coefficients (e.g. e31 or equivalent) of thin-film samples at both wafer-level (without dicing) and on individual discrete diced samples/cantilevers (e.g., via Kanno/Muralt inverse method, cantilever bending, or equivalent). Tenderers MUST clearly detail: - the operational constraints of their setup - Measurement Method: A detailed description of the physical measurement principle - Sample Size and Geometry: The minimum and maximum sample dimensions required or supported by the fixture - Electrode Configuration: The required electrode layouts, top/bottom accessibility, and minimum required electrode dimensions. - Temperature Range: The exact temperature range over which the measurement can be successfully executed using the proposed hardware. (If the fixture is limited to ambient room temperature, this must be explicitly stated). Higher temperature ranges will be scored higher. - Voltage range: The voltage range over which the measurement can be successfully executed using the proposed hardware. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | TEMPERATURE CAPABILITY MATRIX: The tenderer MUST complete Appendix B. Temperature Capability Matrix for every sample environment (e.g., cryo-stages, thermal chucks, or bulk holders) included in their proposal. This matrix will be used to evaluate the system’s performance range across the mandatory temperature and voltage specifications. Tenderers MUST ensure that the data provided in this matrix aligns with the detailed technical descriptions provided in their response to Appendix 1. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
| Hardware & Software Requirements |  |
|  | Software and Data Handling: Integrated Software Suite: The platform MUST include an integrated software environment for the control, acquisition, and analysis of all electrical, temperature-dependent, and optical piezoelectric measurements. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Software and Data Handling: Sequencing and Routines: The software MUST support: • Automated measurement sequencing: Linking multiple tasks into a single execution script. • User-defined test routines: Allowing researchers to customise pulse widths, delay times, and voltage steps. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | ELECTRICAL FERROELECTRIC CHARACTERISATION: Aging and Reliability Measurements (Fatigue, Retention, Imprint): The platform MUST include hardware architecture and automated software routines for fatigue, retention and imprint measurements on bulk and thin-film ferroelectric samples, including: - programmable stress sequences, - intermittent measurement routines, - long-term retention measurements, - automated imprint evaluation, - data logging and export. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Software and Data Handling: Visualisation and Formats: The software MUST provide: - Data visualisation: e.g. Real-time plotting of polarisation (P) vs. voltage (V) and displacement (D) vs. “butterfly” loops - Data export: Support for open formats (e.g. csv, txt., .xml, or .md) to ensure compatibility with external analysis tools e.g. Excel, Origin or equivalent. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Software and Data Handling: User Access & Offline Analysis: The system SHOULD support: • Multi-user operation: Support for a minimum of 5 unique user profiles (or unrestricted). • Offline analysis: Provision for at least 2 additional software licenses (or a site licence) for installation on secondary workstations/laptops to allow data processing and evaluation without occupying the measurement hardware. Tenderers SHOULD state any licensing limits or costs associated with additional user accounts or offline installations. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | The tenderer MUST provide all control and analysis software updates (including bug fixes and security patches) free of charge during the initial warranty period. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | All software updates to the system SHOULD be provided free of charge for a minimum period of five (5) years after the warranty expiry. Please state policy regarding the provision and cost of software updates beyond the warranty period. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Software Access Control: The software SHOULD allow for different levels of access including administrator, user, and maintenance, supporting multiple unique user profiles across these three distinct functional levels. Please confirm if these access levels are included as a standard feature within the proposed software suite. Please detail the functionality of each access level. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | The platform MUST include a dedicated PC/Workstation with the following minimum specifications: - Operating system: Windows 11 or equivalent (most current version, not end of life version) - Keyboard: One standard or ergonomic keyboard (required for defining measurement tasks, entering sample parameters, and software navigation) - Input: one high-resolution optical mouse (and any required proprietary positioning controllers for the stage or probes) - Monitor(s): At least one flat-panel colour monitor(s) (for live full-field video imaging, laser beam positioning and scan measurement set up) - Cabling: All necessary high-speed data links (USB 3.0/Ethernet), system integration cables (Analyzer to PC), and safety interlock cabling (to disable high voltage when the enclosure is open) - Memory (RAM): A minimum of 32 GB Random Access Memory (RAM). This is required to support high-speed data acquisition and the processing of complex software tasks like 2D mapping and remanent hysteresis construction - Storage (Hard Drive): At least 500 GB Solid State Drive (SSD). This ensures sufficient "suitably large" space for high-resolution data storage and fast read/write speeds for full-field video imaging Graphics: Dedicated graphics card with at least 8GB VRAM (required for real-time rendering of complex butterfly loops and high-speed video imaging). Tenderers MUST list the exact specification of the PC/Workstation(s) proposed. Please confirm and detail if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |
|  | Confirm (Yes / No) |
|  | Insert Comment |

- Quality, Health & Safety & Validation

Quality Requirements

|  | All goods supplied in accordance with the contract MUST be of good construction, sound material, robust, free of design defects and manufactured to be safe and without risk when properly used. Please confirm and detail. |
| --- | --- |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Goods and Services MUST comply with all necessary standards, i.e. Irish Standards Institution, European Union, CE, International Standards Organisation or other applicable standard the Tenderer MUST, on the request of the buyer, provide a copy of the standard free of charge. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | If there are manufacturing defects known to the tenderer, they MUST identify such defects and state their policy regarding the repair of known defects. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | The equipment operation MUST be FAT (Factory Acceptance Test) validated prior to shipment to Tyndall to ensure it meets performance specifications. Validation certificates MUST be provided prior to shipment to Tyndall. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Site Acceptance Test After functional acceptance at Validation, the Site Acceptance Test (SAT) of the equipment shall be done jointly by engineers assigned by Tyndall and by the tenderer. A final acceptance document shall be prepared with the following content: - The purchase order reference. - The procurement specification reference. - A list of outstanding issues, including an action time frame for resolving these issues and the division of responsibility for this work. - The signatures and functions of the accepting engineers on each side, the Tyndall, CMOS++ Head of Group and Tenderer’s appointed engineers. - The date of acceptance. - Serial number of the equipment. If the items in the list are satisfied, the equipment will be accepted. If one or more items are not accepted, the tenderer will be required to rectify the deficiency or, alternatively, a partial acceptance may be agreed combined with an extension of the warranty period. Terms shall be negotiated between Tyndall and tenderer based on the results of the acceptance tests. If there is a serious problem on the system, a mutually agreed plan of correction will be generated to correct the problem. Tyndall may waive any portion of the plan. The plan may include the replacement of the system at the discretion of the tenderer. Please confirm compliance. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
| Health & Safety Requirements |  |
|  | Tenderers MUST advise if there are any hazards associated with the operation or maintenance of this instrument? (E.g. manual handing, chemicals, gases etc.). Please detail and confirm. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | The Tenderer MUST, as part of the submission, include copies of any relevant testing or compliance certificates for the Goods. Please confirm and detail. Also state if any additional cost applies to options provided above. Any such additional costs must be detailed in Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | High Voltage Measurements: The high-voltage unit MUST include appropriate safety interlocks and high voltage protection. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | SAFETY & SYSTEM INTEGRATION: The complete platform MUST provide integrated safety handling for high-voltage operation and SHOULD support switching between connected sample holders without repeated manual recabling. If manual recabling is required, the tenderer MUST clearly describe the safety procedure and hardware used to execute the changeover. Tenderers MUST describe: - safety interlocks, - sample chamber interlocks, - amplifier switching concept, - holder switching concept, - electrical and operator safety features. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |

| - Warranty/Spare Parts, After-Sales and Training Support |  |
| --- | --- |
| Warranty & Customer Support |  |
|  | A minimum warranty of two (2) years MUST be supplied with the Integrated Precision Ferroelectric and Piezoelectric Analysis Platform from Validation & SAT and MUST be included as part of the quoted price, to include all hardware, software, spare parts, labour, call outs (24 hour), visit, travel and parts. Note: Longer Warranty period is preferable. Please state warranty period provided. Detail specific inclusions and exclusions of warranty. Preventative Maintenance MUST be provided for one year if necessary. Provision of additional warranty is desirable. Please confirm and detail and outline breakdown in Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | If UCC decide to extend the warranty period, Tenderers are asked to outline cost of any additional warranty. The Tenderer MUST state the details of any extended warranty and details of preventative maintenance and calibration and servicing check and any associated costs. Detail specific inclusions and exclusions of extended warranty with respect to: the Integrated Precision Ferroelectric and Piezoelectric Analysis Platform for Years 3, 4 &5 Please include price per annum, details of coverage and maximum term of coverage. Please include price per annum for additional years of coverage and maximum term of coverage. Please confirm, detail and also specify what is included in the warranty options outlined in the Appendix 2 pricing schedule, Optional to purchase. Detail specific inclusions and exclusions of extended warranty. Note: Pricing MUST only be included in the, Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Please provide separately a list and approximate cost of any consumable items/parts that are to be replaced during each service visit. Please update Appendix 2 Pricing Schedule – Consumables Parts tab with detail. Prices provided for consumable items/parts must be valid for a minimum of one (1) year from effective date of contract. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Please advise of any items/part of the instrument that require routine replacement during the normal lifetime of the instrument, the timeframe in which they require replacing and if such items will be supplied with the instrument or as part of the warranty. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | A list of checks carried out at Preventative Maintenance MUST be provided, along with any information about which (if any) parts are routinely replaced. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | The tenderer MUST provide a description of the arrangements and lead times for guaranteed responses times in the event of equipment breakdown under warranty both in relation to phone support, online/remote support, the provision of technical support, annual service/ maintenance, emergency call-out response, and spare parts. Please detail lead time for arrival on site by a fully trained service engineer during the warranty period (s) and post-warranty. This response time SHOULD be no more than 24 hours via phone call or remote service, and within seven (7) working days for onsite repair. Please confirm and detail information on remote diagnostic / Telecom-based service support. Please confirm detailing response times. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Outside of Warranty Period: The Tenderer MUST provide typical prices and lead time for servicing and repairs (including travel, subsistence and price per day) charged outside a warranty/ maintenance contract. Detail specific inclusions and exclusions outside of warranty period. Please confirm and detail. Please indicate the price in Appendix 2 – Pricing Schedule, Optional to Purchase. Note: Pricing MUST only be included in the, Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | If UCC decide to take out a Service Level Agreement, tenderers are asked to outline Options for subsequent years. Detail specific inclusions and exclusions of Service Level Agreement with respect to: Integrated Precision Ferroelectric and Piezoelectric Analysis Platform: Years 3, 4 & 5 Option to include: - On-going calibration to appropriate schedule - preventative maintenance - remote diagnostics of faults, if applicable - repair of the Instrument in case of breakdown) Tenderers are asked to outline full details of what is included and to outline cost of maintenance and calibration for each year option offered in Appendix 2, Pricing Schedule. Please confirm detail and specify what is included in the Service Level Agreement options outlined, as optional to purchase, in the Appendix 2, Pricing schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | The tenderer MUST have dedicated fully trained and qualified service engineering specialists at their disposal who can respond to all service calls and any form of breakdown call. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
| After-Sales, Repair & Breakdown Support |  |
|  | Please outline the company’s Customer Support Proposal and escalation policy. Tenderers MUST describe what process is in place for escalating a call-out, if a breakdown cannot be resolved on the first visit. Please confirm detail and specify what is included in Customer Support Proposal, as optional to purchase, in the Appendix 2 Pricing schedule, Optional to Purchase. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | The Tenderer MUST indicate the arrangements for and specify the level of experience of applications specialists and engineers who will be in charge of after sales service. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | The tender MUST specify the lead-time for Customer Support during the following scenarios. - warranty period(s) - post-warranty, - with service contract - without service contract Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | The tender MUST provide onsite response, when required, within 7 working days. Please specify the lead-time for Onsite Support during the following scenarios. - warranty period(s) - post-warranty, - with service contract - without service contract Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | The tender SHOULD specify the lead-time for repairs during the following scenarios. - warranty period(s) - post-warranty, - with service contract - without service contract Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Warranty commences post full installation and Final Tool Acceptance and SAT Validation by nominated UCC representative. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
| Replacement and Spare Parts & Equipment Consumables |  |
|  | Please indicate the lead-time for parts during the warranty period(s) and post-warranty, with and without service contract options. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | In event of a system breakdown, the tenderer MUST state the availability and guaranteed minimum response time for spare and replacement parts. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Spare Parts: Availability of all spare parts SHOULD be maintained ten (10) years after the warranty period. Please confirm and detail. Please indicate the availability and lead time of spare part stock centres, with maximum lead-time for any part. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Critical spares SHOULD be held at a location no more than 48 hours delivery time from Tyndall. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Recommended spares lists: If applicable, should be provided by Tenderer. E.g. typical consumables, high-turnover parts. The tenderer MUST supply, where information is available, to the buyer a schedule of spare parts showing in tabular form or, otherwise, the following information on OEM parts: - Names and address of all OEM suppliers - Part Name - Part Description - OEM Supplier Part Number. This SHOULD include typical consumables and high turnover parts. Prices provided for consumable items/parts MUST be valid for a minimum of one (1) year from effective date of contract. Please detail in Appendix 2, Pricing Schedule- Consumable Parts tab. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. This is optional for the Contracting Authority to purchase and will not be included for the purposes of cost evaluation. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Consumable Parts: Sufficient consumable parts MUST be supplied for the initial installation, commissioning and training. These consumable parts MUST be provided free of charge and at the point of commissioning. A full one year of consumable parts MUST also be supplied post-commissioning. Prices provided for consumable items/parts MUST be valid for a minimum of one (1) year from effective date of contract. Please confirm and detail outlining costs in the Appendix 2 – Pricing Schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | The tenderer MUST provide confirmation of the availability of service for at least five (5) years after commissioning. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
| Training and Technical Support Proposal |  |
|  | The tenderer MUST specifically outline the level and duration of application support offered with the instrument where “application support” refers to the expert technical support (on-line or on-site) available from the manufacturer for using the instrument in novel research applications. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Following the delivery of the Integrated Precision Ferroelectric and Piezoelectric Analysis Platform, the successful tenderer MUST provide on-site or online training for a minimum of two (2) Laboratory staff by fully qualified personnel immediately after installation and for at least 2 days. Training SHOULD be completed within three (3) Weeks, or earlier, of delivery. Both Validation and demonstration/training of equipment MUST conducted on same day. The training MUST include: - Equipment set up - Ongoing necessary maintenance - Daily Operation - Software proficiency - Optical metrology alignment - Sample handling and contacting - Cryogenic and thermal management - High voltage safety protocols - System calibration and verification Please specify if training is onsite or online. Please confirm and detail and update Appendix 2 Pricing schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | The training MUST incorporate: SAT and Validation tests (this can be delivered on-site or online) following the equipment installation at Tyndall. The (Site Acceptance Test) will ensure accuracy and overall performance of the system has not been affected by transport and shipping. These tests will also be used for on-site training purposes. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Tenderers are requested to provide a summary of scheduled maintenance and calibration required to keep instrument working to the full specification as per manufactures recommendations. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Relevant training and maintenance manuals in English MUST be provided in soft copy with the system. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |

| - Delivery Plan. Lead Time & Delivery. |
| --- |
|  | The Integrated Precision Ferroelectric and Piezoelectric Analysis Platform SHOULD be delivered within nine (9) months of receipt of Purchase Order. Please state the exact number of months for delivery. A range of months is not acceptable. (Note that shorter delivery time is highly desirable and will be scored favourably). - Tenderers must provide full details of lead times for the delivery of equipment including installation, commissioning & validation (if applicable), from receipt of the purchase order. - Tenderers must provide the facility requirements to UCC a minimum of four (4) weeks in advance of the equipment delivery. - Tenderers must provide advance notice, three (3) weeks of any requirements for tool unloading, number of containers and weight of each. Height, length and width of containers/ crates to be provided as well. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Delivery MUST include all shipping charges, insurance, customs clearance and all aspects as outlined in DAP (Delivery At Place) incoterms. Delivery must be to Tyndall National Institute-University College Cork: Goods In/Goods Out Dept Tyndall National Institute, University College Cork Lee Maltings Dyke Parade T12R5CP Cork, Ireland Advance notice MUST be given, two (2) weeks of any requirements for tool unloading, number of containers and weight of each. Height, length and width of containers/crates MUST also be provided. Please confirm and detail and update Appendix 2 Pricing schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1 document. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | The tenderer MUST provide details of the logistics/supply chain to the point of customer delivery, i.e. where the goods will come from, how the goods will be transported and what borders the goods will cross. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | The tenderer MUST provide details of actions that have been taken to ensure the future timely supply and delivery of goods and services, including ongoing supply of replacement/spare parts and consumables. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | The tenderer MUST provide details of their contingency plans for any changes in the delivery programme (alternative sources/routes of supply) for any delays in delivery should they arise. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
| Installation, Commissioning & Calibration Requirements |  |
|  | The supply of all electrical items MUST be in full accordance with, but not limited to, European standards / regulations and all relevant amendments. Please provide electrical requirements for operation and approximate power consumption. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | The proposed system MUST be suitable for use with a standard 230V, 50 Hz, single-phase electrical supply at the proposed installation site. If a transformer unit is required to achieve this, it MUST be supplied as part of the system. The system MUST utilise standard EU/CEE-7/VII grounded plugs (or equivalent 15A fused mains lead). Tenderers should also note that the following facilities will be available at the installation site: - Electrical supply: 230 V AC, 50 Hz single-phase - Environment: Temperature-controlled laboratory (19°C – 30°C) with humidity control (25% – 60%). - Nitrogen supply: Access to a nitrogen (N2) gas line and liquid nitrogen (LN2) for cryogenic modules - Connectivity: standard Ethernet port Please confirm and detail any additional requirements. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | Installation, commissioning and calibration of the equipment MUST be carried out on-site, by trained engineers provided by the tenderer. Tyndall will provide for any labour and materials as reasonably requested for the purpose of installation. Please confirm and detail and update Appendix 2 Pricing schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1. Document. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | The Tenderer MUST provide a qualified Engineer to perform the SAT, post installation and validation, to ensure accuracy and overall performance have not been affected by shipping. These tests will also be used for on-site training purposes. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | The tenderer MUST validate the system and verify full functionality at the Tyndall National Institute site (as an on-site or on-line process). Tenderer has to make a written request for this support at least four (4) weeks in advance of installation. Please confirm and detail and update Appendix 2 Pricing schedule. Note: Pricing MUST only be included in the Appendix 2 Pricing Schedule. No pricing is to be included in this Appendix 1. Document. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | A signed certificate for successful Site Acceptance Test (SAT) MUST be provided upon completion of training delivered by the tenderer on the operation and use of the equipment. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
| Documentation |  |
|  | A detailed User Operational Manual for the Integrated Precision Ferroelectric and Piezoelectric Analysis Platform MUST be provided. This should include the following. - Full electrical and design specifications. - Manuals MUST be supplied in English language. - Manuals MUST be provided in softcopy, and preferentially in hard copy too Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | A maintenance and validation manual for the system MUST be provided to include the following: - Cleaning requirements service if needed - Maintenance and Validation requirements. - Manuals MUST be supplied in English language. - Manuals MUST be provided in softcopy, and preferentially in hard copy too. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
|  | A Technical Specification for the Unit MUST be provided covering all major components of the system. The specification MUST be supplied in the English language and at least as soft copy. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |

- Energy, Environmental & Sustainable Considerations

Environmental & Sustainable Considerations

| E1 | Please outline your company’s commitment to the sustainable life cycle of the tendered equipment in a manner that minimises adverse environmental impacts. Please confirm and detail. |
| --- | --- |
|  | Confirm (Yes / No) |
|  | Insert Comment |
| E2 | Energy & Environmental Requirements: The Contracting Authority requires the equipment provided to be as energy and economically efficient as possible. The Integrated Precision Ferroelectric and Piezoelectric Analysis Platform equipment SHOULD have low running costs allowing the equipment to be as energy & economically efficient as possible. Please provide details on any “warm-up” times required for the equipment to operate to specification. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
| E3 | Tenderers SHOULD provide details of any takeback arrangements they may offer when the equipment reaches end of life and how that may be availed of. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
| E4 | Tenderers SHOULD indicate typical Energy Consumption during operational time while the equipment is in use. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
| E5 | Where possible, use of a standby mode/reduced power mode of operation when the system is not in use. The tenderer SHOULD specify the stand-by energy consumption. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
| E6 | It is DESIRABLE that all Tender requirement is delivered using environmentally friendly packaging materials; packaging, tape, labels and stretch wraps. Ideally, this packaging SHOULD avoid any single use plastic and be recovered or at a minimum reusable. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |
| E7 | Tenderers SHOULD provide details of expected lifetime of the equipment required. Tenderers must state how long the model of equipment has been on the market and indicate the expected working life of the proposed models and longevity of equipment, etc. Please confirm and detail. |
|  | Confirm (Yes / No) |
|  | Insert Comment |

| - Ultimate Cost |  |
| --- | --- |
| F1 | Please see Appendix 2 – Pricing Schedule for details. |

## B4: Appendix 2- Pricing Schedule

Please use the Appendix 2 – which is provided as a supplementary Excel – titled:

LA3673C Appendix 2 Pricing Schedule

The Appendix 2 Pricing Schedule provided in the Excel file MUST be completed per Lot/section(s) you are Tendering for. Please submit in excel format. The relevant Pricing Schedule must be completed in full, totalled, and returned as part of the Tender submission.

Tenderers are reminded that the prices/discounts quoted MUST be in compliance with the requirements set out in paragraph 2.10 of the RFT, i.e.: they MUST be all-inclusive, fixed, expressed in Euro, exclusive of VAT and remain fixed for the term stated in the RFT.

Any particulars and merits should be identified in the comments section to facilitate scoring under paragraph 3.3 Award Criteria of the RFT.

## B5: Checklist for Tender Submissions

| Please confirm that the information in TRD Section A. (A1-A9), relating to RFT Section 3. have been completed and returned as requested. | Confirm Yes/No |
| --- | --- |
| Please confirm that all the information requested in Appendices provided have been included in the tender response. | Confirm Yes/No |
| Please confirm that the Appendix 2 Pricing Schedule has been returned and included in your proposal. | Confirm Yes/No |
| Please confirm that Appendix 3 (on headed paper) has been completed and returned as requested. | Confirm Yes/No |
| Please confirm that Appendix 4 has been signed and completed and returned as requested. This document MUST also be signed, stamped and dated by a Practising Solicitor / Commissioner for Oaths. | Confirm Yes/No |
| Please confirm that Appendix 5 Declaration and Warranty on the Use of Artificial Intelligence has been completed. | Confirm Yes/No |
| Please confirm that the European Single Procurement Document (ESPD) as per 2.2.1 (a) of the RFT document has been completed. | Confirm Yes/No |
| Tenderers must address each of the issues and requirements and submit a detailed description in each case which demonstrates how these issues and requirements will be dealt with / met and their approach to the proposed delivery of the Goods. A mere affirmative statement by the Tenderer that it can / will do so or a reiteration of the tender requirements is NOT sufficient in this regard. | Confirm Yes/No |
| If multiple equipment types are available which suit the specifications required, please provide details of each option on a separate TRD and Appendix 2, clearly defining each option. | Confirm Yes/No |
| Please ensure that all supporting documentation is submitted in the English language. | Confirm Yes/No |
| Please ensure all documents are named accordingly and submitted within a zipped/compressed folder. | Confirm Yes/No |

Please ensure to sign confirmation form at end of this document.

This form must be completed and signed by an authorised officer of the tenderer’s organisation.

- This offer will remain open for acceptance by you for a period of 180 days from the date of deadline for submission of tenders.

- We acknowledge that you are not obliged to accept the lowest or any offer.

- I/We undertake to deliver the products/services in accordance with the terms and conditions of the tender specification.

- I/We undertake to maintain full confidentiality with regard to all aspects of this tender process

| Signed: |  | Position: |  |
| --- | --- | --- | --- |
| Print Name: |  | Phone No: |  |
| Company Name: |  | Date: |  |
| Address: |  |  |  |
| Email: |  |  |  |

Document end