## Instructions to Tenderers
Appendix 1 Tenderers Response Book for the Provision of Paediatric Cardiac Ultrasound for the New Children's Hospital, Dublin 8 CfT 8783106
Tenderers notes for completing Appendix 1: 1) This document contains an introduction and overview for this tender and the technical and functional specifications for a quantity of five paediatric ultrasound devices required for the cardiology department of the new children's hospital. The quality, aesthetics, functional characteristics, technical merit, environmental characteristics, regulatory standards and performance levels for the goods and services to be provided by each of these tendered solutions are listed in this document. Tenderers should respond to these stated requirements with solutions that meet or exceed the stated requirements and provide any necessary details to reassure that the stated requirement could be met by the solution tendered.
2) Tenderers may submit by completing each of the five numbered Tabs in this response book via etenders before the deadline expires
3) Each Tab may comprise a mix of mandatory and assessed requirements, as follows: (i) Mandatory : are the minimum requirements that a tendered solution must meet to remain in the competition (ii) Assessed: is how well the tendered solution is considered to meet the specific requirement. The maximum marks available are stated on each tab.
4) Tenderers responses should be concise, complete and relevant to the solution tendered. Evaluation will be based on Tendered responses, product sampling may be required to validate scores. Tenderers should complete all questions, only the response provided will be considered in evaluation. A blank response to an assessed requirement will not achieve a maximum score.
5) Where relevant, a response to a previous question may be repeated or clearly cross referenced.
6) Tenderers response areas on each tab of this document are highlighted in yellow.
7. Tenderers should check the followng Tab's are complete before submitting with any supplementary material requested: Tab # 1. Specification & Standards Tab # 2. Dicom EHR Tab # 3. After Sales Service Tab # 4. Reference Sites Tab # 5. Infection Control

## Introduction
Overview of Paediatric Echocardiograph requirements for the New Children's Hospital, Dublin 8
Introduction The outline requirements are for an ultra-high end, fully configurable cardiac ultrasound platform optimised for advanced paediatric echocardiography to meet the requirements for scanning Neonatal, Paediatric and Adolescent/Adult patients. The system must be capable of custom configuration and incorporate the latest advancements in 2D, 3D & 4D imaging, probe technologies, and AI enabled functionality. It should be built on a modern scalable platform that supports future upgrades and represents the most current technologies available from manufacturers. These systems are essential to ensure efficient, safe and minimally distressing ultrasound imaging for all our patients within the hospital environment. The system must support Neonatal, Paediatric and Adolescent/Adult patient studies on a single platform. Particular emphasis will be placed on its suitability for paediatric patients, including adolescents approaching adulthood, in line with National Children’s Hospital(NCHI) requirements. Tenderers should note the patient cohort is diverse, and although the hospital is a paediatric facility, a significant proportion of adolescent patients are, to all intents and purposes, young adults. This necessitates a system that also incorporates full adult imaging capabilities. NCHI serves as the national tertiary and quaternary referral centre for paediatric patients and also provides primary and secondary care for the greater Dublin area. As such, the latest state of the art equipment is required to meet the demands of a high-volume cardiology ultrasound imaging service. The system must be suitable for use across multiple clinical environments, including: Cardiac Outpatient Department Cardiac Intensive Care Unit Cardiac Catheterisation Theatres, Cardiac Hybrid Imaging Theatres Open Heart Surgery Theatres. Mobile bedside imaging should patients be unable to attend cardiac inpatient clinics. It is imperative that all equipment is highly reliable, supports efficient clinical workflows, and operates consistently to minimise risks to patients and staff The system must deliver high quality imaging that ensures heart structures are quickly and easily identified and provide precise quantitative data. It must also seamlessly integrate with PACS Systems ensuring that imaging data is efficiently transferred and readily accessible. This will support clinicians in determining optimal treatment pathways and monitor patient progress in a safe, efficient and child friendly environment. This document sets out the requirements for the system and invites the market to respond with a suitable configuration to meet these needs.
Project Overview
The current programme indicates the construction of the new children's hospital will be complete in August 2026. However, lessons learnt from other large hospital builds indicate these dates may vary slightly as the complexities of the building programme may result in timeline changing in response to factors that cannot be predicted. Therefore, once awarded, this timeline will be updated in accordance with the handover and commissioning plan.
Tender Scope
The initial requirement is for a quantity of five (5) Cardiac Ultrasound scanners for delivery before year end (2026). Children's Health Ireland may draw down additional (<5) scanners should replacement be necessary and funds available before they transfer to the new facility.
Clinical Overview
Echocardiography (cardiac ultrasound) is an essential imaging modality in the diagnosis and assessment of both congenital and acquired heart disease, from prenatal stages through to adulthood. Transthoracic (TTE) and transoesophageal echocardiography (TOE) are key imaging tools in infants and children with congenital heart disease. TTE is non-invasive, while TOE is minimally invasive; importantly, neither technique involves ionising radiation, and both provide detailed anatomical, haemodynamic, and functional information. Advances in technology, including AI, tissue doppler imaging, myocardial strain assessment, and 4D echocardiography, have expanded the role of echocardiography from a purely structural diagnostic tool to a comprehensive functional imaging modality. 4D imaging enhances pre-surgical assessment, particularly in complex intracardiac anomalies, septal defects, and valvular pathology. Tissue doppler and ventricular strain techniques support detailed evaluation of myocardial contractility and ventricular function. In addition, foetal echocardiography enables prenatal diagnosis of congenital heart disease, facilitating early detection, improved survival rates, and effective perinatal planning. The ultra-high specification paediatric cardiology ultrasound system shall provide an advanced platform for use within the main cardiac outpatient ultrasound department, the Cardiac Intensive Care Unit (CICU), and inpatient settings. Systems may also be required for advanced bedside and emergency imaging throughout the hospital, enabling rapid clinical assessment when needed. The system must represent a state-of-the-art solution from an established manufacturer and support the full range of ultrasound probes and clinical applications. Furthermore, it must accommodate all advanced imaging modalities and probes while ensuring a future-proof platform capable of supporting ongoing technological developments and enhancements. Paediatric imaging presents unique challenges, which must be reflected in the system’s design and functionality. The system must support advanced examinations in neonates, infants, and children of all ages and varying levels of cooperation, while promoting a positive patient experience—particularly important given the likelihood of repeated imaging over time. To meet the required specification, the system must not only comply with the functional criteria outlined in this document and also successfully pass the mandatory requirements stated to progress to assessment and clinical evaluation by end users within a paediatric hospital setting.
Technical Overview The ultra-high specification paediatric cardiology ultrasound scanner must comply with the minimum standards set out in this document.
Detailed Description The supplier is asked to quote for the supply, installation, commissioning and user plus technical training of an ultra-high specification paediatric cardiology ultrasound scanner and associated equipment to meet the needs set out in this document. Please note the ultra high premium-range system must:
1. Operate as a cart-based unit (i.e., Permanently mounted on suitable wheels/casters that incorporate two wheel lock for steering and four wheel lock for stability when in use)
2. Incorporate a large-screen video display on an articulated arm.
3. Include an independently adjustable user-interface control panel.
4. Include first-rate interoperability when integrating into CHIs Health IT infrastructure. All aspects including local storage for fast access to recent studies, PACS, NIMIS and integration of the Healthcare Enterprise (IHE) provision must be seamless.
5. Support a variety of transducers, including curved linear, flat linear, sector, micro convex, intraoperative, transoesophageal and endocavity in addition to newly emerging transducer technologies.
6. Support and offer innovative AI/Deep learning technologies that enhance the quality of care and processes.
7. It is desirable that the system supports the following advanced imaging modes or features: - 4D imaging - Atrial Function Analysis - Ventricular Function Analysis - Strain Elastography - CT Fusion Imaging - Advanced contrast-specific imaging capabilities designed to enhance myocardial border delineation and visualize fine anatomical structures, especially in technically difficult-to-scan patients - Innovative AI/Deep learning technologies - Advanced colour flow imaging - Custom imaging transducers for small body types and challenging anatomies regularly found when scanning paediatrics
CHI is keen to be at the forefront of paediatric imaging and any partnership or innovative solutions in research or data sharing opportunities proposed by Tenderers would be welcome. In addition to the above standard package requirements, Tenderers should provide details (as optional extras) on any additional features, functionality or innovations included or offered as part of this tender response that would maximise or enhance the operational, performance maintenance or extend the potential usable life of the scanner.

## 1. Specification
Appendix 1 Tenderers Response Book for the Provision of Paediatric Cardiac Ultrasound for the New Children's Hospital, Dublin 8
Technical and functional specifications listed below comprise Mandatory requirements, which are the minimum requirements to be met and Assessed requirements which will be scored based on how well the tendered response meets the stated requirement.
Tenderers must complete and respond to each line item on this tab
A maximum of 500 marks are available for scored requirements in this section
| Requirement | Mandatory or Assessed | Confirm (Select Yes or No) | Vendors Response Insert detail in this column where it is indicated
1 | System must support both paediatric transthoracic (TT) and Multiplane Transoesophageal (TOE) Imaging including all standard and advanced modalities as detailed in this document. | Mandatory
2 | Proposed system must include complete compliment of multi frequency transducers to image the patient cohort and support the studies detailed in this document. Specify what is provided in the bid and clearly indicate any options/extras on the Pricing Sheet (appendix 2) | Mandatory
3 | The system must be capable of future software/hardware expansion without replacement of the base platform where technically feasible. Please confirm | Mandatory
4 | The system must be enabled to support contrast injection imaging, including LVO assessment and other contrast-enhanced studies. Confirm | Mandatory
5 | The system must support fast boot-up times and incorporate a battery mode that enables operation between scanning locations and provides backup power in the event of a mains power failure. Confirm | Mandatory
6 | System must support automated image-alignment that uses AI or image-recognition algorithms to identify cardiac anatomy and automatically orient views. This should reduce the need for operator manipulation of datasets without compromising images for structural heart imaging and 4D transoesophageal echocardiography. Particularly desirable for views along complex anatomical structures such as valves, regurgitant jets. catheters. Specify how the solution meets this requirement. | Mandatory
7 | The system must include image processing filters that smooth borders and structures while preserving anatomical accuracy. Confirm | Mandatory
8 | The ultrasound system must incorporate real-time adaptive image optimisation technology that automatically provides uniform gain, brightness and contrast across the full imaging field while compensating for variations in tissue attenuation and imaging depth. The system must allow for selection of different optimisation profiles (e.g. soft or sharp) according to clinical application. Confirm and detail how this is achieved. | Mandatory
9 | The system must include automatic recognition of the most common Doppler spectra and, where supported, automatically initiate the corresponding Doppler measurements. If automatic measurement is not available, the system shall open the appropriate manual measurement table for user input. Detail how the solution meets this requirement. | Mandatory
10 | The system must incorporate automatic low power settings that are adjustable/reprogrammable for paediatric protocols. Confirm | Mandatory
11 | The system must be fully DICOM and IHE enabled without the use of proprietary offline software. It must support interoperability with third-party clinical analysis and reporting platforms through standard DICOM formats and, where available, vendor-supported native/raw data export. Tenderers must complete Dicom EHR sheet (Tab # 3) in this document. | Mandatory
12 | System must support real time 2D, 3D & 4D Imaging, Ventricular Function Imaging and assessment, M- Mode, Spectral Doppler, Colour Flow, Power Doppler, PW Doppler, Duplex and Triplex Scanning. Confirm a complete M&A package is included for all imaging modalities. | Mandatory
13 | System must include Z-scores for standard structural modalities to provide structured reporting and Paediatric Heart Network (PHN) data, Boston (Share), Detroit, and typical European models. Solutions should offer a seamless automated workflow, where entering patient's height and weight auto-calculates Z-scores across multi-dimensional 2D, M-mode, and Doppler acquisitions. In addition, the ability to quickly switch between or compare different regional datasets directly on the measurement touch screen is desirable. Detail what is included in the core bid solution to meet these requirements. | Mandatory
14 | System must include a three lead ECG lead and module for cardiac gating, real time monitoring and automatic measurements. Confirm what is included in the core bid solution. | Mandatory
15 | A range of user selectable automated measurements (i.e., PS, ED, TAMx, TAMn, PI, RI, S/D ratio etc.) must be available in real-time and frozen doppler modes. Please list what is available. | Mandatory
16 | The unit must be supplied with a minimum of eight (8) paediatric, age and/or weight-based programmes as standard plus additional user configurable programmes. Users must have free control of these parameters when selected. Confirm. | Mandatory
17 | System must include Harmonic Imaging as standard. Confirm | Mandatory
18 | The system must include high-definition zoom including full capability to perform offline or post-acquisition calculations and measurements on both frozen static images and cine loops. Confirm capability, range, parameters and any restrictions for the tendered solution. | Mandatory
19 | The proposed ultrasound system must provide real-time, automatic speed-of-sound optimisation to compensate for varying tissue densities and acoustic characteristics within the patient. Please confirm. | Mandatory
20 | Confirm that the system proposed incorporates a physical alphanumeric keyboard that is suitable for a healthcare environment. Provide make/model or image. | Mandatory
21 | System must provide simultaneous connection of four or more imaging transducers, with ease of switching and fast initialisation. Confirm | Mandatory
22 | Confirm that the system is capable of displaying user selectable mechanical and thermal indices. | Mandatory
23 | System must have suitable castors for mobile scanning that can be braked and two locking castors for ease of steering. It is desirable to have front and rear brake pedals for improved control. Specify what is included in the proposal. | Mandatory
24 | In the event that the device tendered is superseded by a next generation model version, or platform change prior to commissioning and clinical use. Tenderer must agree to substitute and supply, at no additional cost to the contracting authority, the latest solution with equivalent or enhanced functionality. | Mandatory
25 | List and cost any optional extras (i.e. not essential for core solution functionality) in the options section of Appendix 2 Pricing Schedule, for any peripherals (printers, external footswitch, ECG cables etc.) that may be integrated on system. | Mandatory
26 | To support confident diagnosis of paediatric routine and complex congenital heart disease, it is essential the system delivers exceptional image quality, spatial and contrast resolution, and tissue definition across the paediatric patient cohort. Please detail how the proposed system meets this requirement. | Assessment
27 | Clear visualisation of complex anatomy is essential for paediatrics, therefore a mode to enhance blood tissue boundaries would be beneficial. Please detail what modes are included in the solution and highlight any with particular benefits for paediatric patients. | Assessment
28 | Children with specific congenital heart diseases (CHDs) benefit from LAA procedures primarily to prevent thromboembolisms (strokes), eliminate sources of refractory arrhythmias, or treat localized tissue defects. The proposed system should be capable of automatically identifying the LAA in 2D or biplane view to enable targeted zooming of the LAA image, to facilitate fast acquisition of high-resolution 4D recordings of the LAA. Describe how the solution supports this. | Assessment
29 | In addition to # 28 it would be advantageous for the proposed solution to incorporate a semi automated boundary tracing algorithm to cater for our highly variable patient sizes. Detail what is provided in the core bid or if it is an option listed in Appendix 2 Pricing Schedule. | Assessment
30 | It is advantageous for the system to support real time 2D & 4D semi-automated TSP(Transseptal Puncture) guidance for safe navigation of the interarterial septum. Does the solution provide any assistance in procedural planning by calculating the distance from the operator identified mitral annulus to the catheter tip ? Detail what is available and included in the core bid or specify any optional costs in the Pricing Schedule (appendix 2) | Assessment
31 | System should support automatic assessment of left heart anatomy and function utilising volumetric echocardiographic datasets. Confirm how this functionality is validated on paediatric patients. | Assessment
32 | It is desirable that the system supports 4D volume rendering of boundaries between cavities and tissue structures (or devices) to visualize internal structures or to avoid overlying tissue degradation/blockage of colour flow acquisition. Confirm and detail what is provided. | Assessment
33 | The proposed system should ideally be capable of producing a three-dimensional visual representation on standard 2D colour flow images in order to improve visualization of the myocardium, valves and defects such as VSDs. Detail what is provided. | Assessment
34 | System architecture must be programmable and capable of supporting 2D, 3D and 4D Imaging without any deterioration in frame rates or spatial resolution. Acoustic acquisition and processing must ensure the ultrasound images offer the optimal in resolution. Detail how the solution meets this requirement. | Assessment
35 | System should provide optimal clinical information from each scan. Paediatric patients are more technically difficult so the ability to perform fast, consistent exams is highly desirable for patient ease and comfort. Describe how the solution proposed supports this. | Assessment
36 | The overall system design should provide consistent optimal clinical performance in a high utilisation environment. An Ergonomic solution is required in high usage areas such as paediatric cardiology outpatients, cardiac catheterization theatre, cardiothoracic theatres, CICU etc. where accuracy, reliability and user adaptability are paramount. Please provide image and dimensions plus weight of the solution tendered and how it meets this requirement. | Assessment
37 | It is desirable the system provides an easily accessed control to optimise speckle reduction and edge enhancement for clear visualization and differentiation between anatomical structures. This should be supported on all transducers. Provide details. | Assessment
38 | Proposed system should contribute to an efficient workflow when in use for department or mobile imaging studies. Detail software included in the bid that meets this requirement, identify any options costed separately in options on Pricing Schedule (appendix 2) | Assessment
39 | It is desirable the system can provide automated 4D left heart function metrics and volumes that may be included in routine echo examinations. Detail what is provided in the bid, indicate any options costed separately in the options section of the Pricing Schedule. | Assessment
40 | The proposed system should ideally be capable of providing a solution for the detection and classification of regurgitant or paravalvular jets, Specify if this is provided and how it is achieved?. | Assessment
41 | The user interface must be well designed and intuitive for ease of use. It should incorporate a customisable touchscreen and minimal switches/rotaries. These controls should be clustered in an efficient manner and automatically illuminate in use when available in specific applications Please provide detail of touchscreen size, functionality, controls and intuitive features with reference images where appropriate. | Assessment
42 | The user interface should be height adjustable to accommodate users standing or seated. Please provide details of adjustable capabilities and ranges. | Assessment
43 | System must include an independent, articulated Hi-resolution image display monitor, minimum twenty-two inches. Please provide details of monitor, resolution, grayscale shades, viewing angle, contrast ratio etc. | Assessment
44 | It is desirable that the system supports CT / MRI Fusion Imaging plus Strain Elastography. Please provide details indicating what is included as standard and list any options costed separately in options section of Pricing Schedule. | Assessment
45 | System should include a complete anatomical labelling library to include both diagrammatic and user programmable narrative. Specify what is provided as standard in the core bid. | Assessment
46 | System must incorporate factory range of pre-sets plus support user programming of custom pre-sets for each probe. Provide details plus maximum number of pre-sets that can be assigned to each probe. | Assessment
47 | Please list all factory grayscale plus tint maps and number of user programmable maps the proposed system supports. | Assessment
48 | It is desirable that the system is capable of compensating for or reducing motion artefact particularly with young patients. How is this provided by the tendered solution? | Assessment
49 | The proposed system should ideally be capable of maintaining high spatial and temporal resolution even when switching to 4D, ensuring that volumetric images are sharp and detailed. Please indicate how this can be achieved. | Assessment
50 | It is desirable that the System proposed has capacity to store studies should a transfer to PACS be delayed or unavailable. Detail storage capacity included in the bid as standard. | Assessment
51 | An optional wireless PACS connection is desirable, specify what is included as standard? | Assessment
52 | Please provide details on: Transducer holders that securely store cardiac probes when not in use. Cable Management: to keep transducer cables organized and off the floor. Indicate what is included in core bid or costed separately as options in Pricing Schedule. | Assessment
53 | Please provide details on the original launch date of the system, list any major hardware or software iterations released since launch, and the anticipated product lifecycle | Assessment
Imaging Transducers Tenderers to confirm compliance with all Mandatory Requirements
54 | Intended Use: Transducers must be designed for diagnostic medical ultrasound imaging. Transducers must support intended clinical applications as specified. Device must be suitable for continuous clinical use in hospitals. | Mandatory
55 | The transducers included in the core solution bid (Appendix 2 Pricing Schedule) must be suitable for the following paediatric Cardiac Applications: Neonatal echocardiography 2D & 4D Paediatric echocardiography 2D & 4D Adolescent/Adult echocardiography 2D & 4D Stress echo Foetal echocardiography | Mandatory
56 | All transducers and accessories supplied must be manufacturer approved and certified for use with the proposed system. System must be manufacturers latest/current production model. | Mandatory
57 | The following Imaging transducers must be included in tender bid: 2D Paediatric Probe 4.0 - 12.0 MHz 2D Neonatal Probe 2.0 - 8.0Mhz 2D Adolescent/large Paediatric patient probe 1.5 - 5.0MHz Please provide details stating model, frequency range, sector angle range, footprint size and maximum depth of field provided by each transducer. If the specified frequency range requires additional transducers (i.e. >3) they must be included in the submitted core bid price (Appendix 2 Pricing Schedule). | Mandatory
58 | Transducers will be decontaminated in a high-level disinfection system that uses sonically activated hydrogen peroxide mist to eliminate bacteria, fungi, viruses, and spores from ultrasound probes. The transducers proposed must be manufacturer approved for high level disinfection in the Nanosonics Trophon 3 System (device used by Children's Health Ireland). Please confirm. | Mandatory
59 | Transducers must provide: High-resolution B-mode imaging. Uniform image quality throughout field of view. High contrast resolution. Low image noise. Excellent near-field visualization. Excellent far-field penetration. Tissue harmonic imaging compatibility. Speckle reduction capability. Compound imaging compatibility. Confirm this is provided (subject to sampling verification) | Assessment
60 | Transducer cables must be: Lightweight and flexible. Incorporate reinforced strain relief. Sufficiently long to compliment ease of use. Resistant to repeated flexing and twisting. Please provide details on how your proposed solution meets these requirements. (subject to sampling verification) | Assessment
61 | Probe Housing must be: Durable & impact resistant. Smooth with a sealed surface. Easily cleaned and disinfected. Waterproof or fluid-resistant. (subject to sampling verification) | Assessment
62 | Transducer Ergonomics should be complimented through features that enhance operator comfort, manoeuvrability and ease of use during routine and complex paediatric examinations. Desirable features are: Lightweight construction. Ergonomic grip. Balanced handling. Reduced operator fatigue. (subject to sampling verification) | Assessment
63 | Tenderers should list and cost all additional probes available for the system under Options (Appendix 2 Pricing Schedule) and state any volume based discount available. Additional probes should include 4D Probes(both TT & MPTE), standard MPTE Probes, Foetal Echo Probes etc. Although the base system specification includes three probes, tenderers should identify any specific optional transducers within their portfolio that may provide additional clinical benefit for our paediatric patient cohort. The response should describe the key features of each recommended transducer, including its intended clinical applications, frequency range, footprint, and any unique technologies that enhance imaging performance across neonatal, infant, paediatric, and adolescent patients. Tenderers should note that the Procurement Evaluation Group (PEG) may request selected optional transducers to be made available for review and evaluation during the sampling stage of the tender assessment process.
A.I.
64 | Manufacturers are invited to provide a comprehensive description of all Artificial Intelligence (AI) programmes and functionalities available within the proposed system. Submissions should clearly outline how these AI-driven tools enhance clinical performance, workflow efficiency, and diagnostic accuracy, for all patients and particularly within a paediatric and adolescent patient cohort. Responses should include detailed information on the scope and maturity of AI applications embedded within the system, such as (but not limited to): • Automated image acquisition guidance and optimisation • Real-time image enhancement and artefact reduction • Automated measurements (e.g. chamber quantification, ejection fraction, strain analysis) • AI-assisted interpretation and decision support tools • Workflow automation, including exam protocol guidance and reporting assistance Manufacturers must clearly articulate the clinical advantages of each AI feature, supported where possible by published evidence, validation studies, or real-world clinical use cases. Particular emphasis should be placed on how these technologies: • Improve accuracy and reproducibility of measurements • Reduce operator dependency and inter-observer variability • Shorten examination and reporting times • Enhance patient experience, especially in paediatric settings where scan time and cooperation are critical • Support less experienced operators such as trainee sonographers while maintaining high diagnostic standards In addition, suppliers should describe: • The extent to which AI tools are configurable for both paediatric and adult populations • Integration with existing hospital systems (e.g. PACS, reporting platforms) • Requirements for training, updates, and ongoing support • Regulatory status (e.g. CE marking) and cybersecurity considerations All AI functionality must be clearly identified as standard or optional, and any associated licensing requirements should be disclosed. The objective of this request is to ensure that proposed systems leverage state-of-the-art AI technologies to support safe, efficient, and high-quality cardiac ultrasound imaging in a demanding, high-acuity paediatric healthcare environment | Assessment

## 2. Dicom EHR
DICOM and Electronic Health Record (EHR) Connectivity Requirements for the Provision of Paediatric Cardiac Ultrasound for the New Children's Hospital, Dublin 8
Tenderers must complete this tab
1 | The ultrasound system must be fully DICOM compliant and capable of seamless integration with existing hospital information systems, radiology information systems, cardiovascular information systems, picture archiving and communication systems, and electronic health record platforms. | Mandatory
2 |  | Mandatory
3 | The system must support, at a minimum, the following DICOM services: | Mandatory
4 | DICOM Storage (Send/Receive) | Mandatory
5 | DICOM Verification (C-ECHO) | Mandatory
6 | DICOM Modality Worklist (MWL) | Mandatory
7 | DICOM Modality Performed Procedure Step (MPPS) | Mandatory
8 | DICOM Query/Retrieve | Mandatory
9 | DICOM Structured Reporting (SR) | Mandatory
10 | DICOM Storage Commitment | Mandatory
11 | DICOM Print | Mandatory
12 | DICOM Media Export | Mandatory
13 | The system must automatically retrieve patient demographics from the hospital information system through DICOM Modality Worklist, minimizing manual data entry and reducing identification errors. | Mandatory
14 | The ultrasound platform must support secure transmission of images, measurements, waveforms, reports, and structured data to PACS and enterprise archives in real time. | Mandatory
15 | The system must be fully Electronic Health Record (EHR) enabled and support interoperability with major healthcare information systems utilizing industry-standard communication protocols, including HL7 and FHIR where applicable. | Mandatory
16 | The system must be capable of: 1) Bidirectional communication with Hospital Information Systems (HIS) 2) Integration with Electronic Medical Records (EMR) and Electronic Health Records (EHR) 3) Integration with Radiology Information Systems (RIS) 4) Integration with Cardiovascular Information Systems (CVIS) 5) Automatic patient demographic synchronization 6) Automated exam scheduling and worklist management 7) Electronic report generation and distribution 8) Secure clinical data exchange across enterprise networks | Mandatory
17 | The ultrasound system should support user authentication, role-based access control, audit trails, and data encryption to comply with applicable healthcare cybersecurity and patient privacy requirements. The vendor must provide all necessary software licenses, interfaces, configuration services, testing, validation, and implementation support required for full DICOM and EHR integration at the customer site. Compliance with current DICOM standards and interoperability testing documentation shall be provided at the time of bid submission. | Mandatory

## 3. After Sales & Service
After Sales & Service Requirements for the Provision of Paediatric Cardiac Ultrasound for the New Children's Hospital, Dublin 8
A maximum of 100 marks are available for scored requirements in this section
Tenderers must complete this tab
Question Number | Requirement to be met by the Tendered Solution |  | Indicate Compliance with Mandatory requirements (Yes or No) | Tenderers response to be inserted in this Column (Any supplementary material requested, should be listed in this column and the supplement cross referenced to the question number)
1 | Please provide details of warranty duration on goods supplied to include parts & labour for all components provided as part of the proposed solution. Tenderers must include warranty terms and conditions to meet these requirements Additional marks will be awarded for extended warranty provisions that demonstrate enhanced value and reduced lifecycle risks. | Assessed
2 | Frontline training in system operation, routine checks, and basic troubleshooting shall be provided to HTM (Healthcare Technology Management) staff upon installation. Please provide details and confirm it will not be a breach of warranty should one of our trained clinical engineers attend to the system during warranty period. Please provide details and number of staff that will be trained. | Assessed
3 | The vendor must also supply full manufacturer certified service and technical training for the tendered system in all aspects of device operation, preventative maintenance and service to two HTM staff. Please provide full details and associated costs. | Assessed
4 | Subject to HTM personnel receiving appropriate manufacturer-approved training and certification, all service-level passwords, access credentials, and engineering functions available to the vendor’s service engineers should also be made available to the authorised HTM personnel. This will enable HTM undertake maintenance, diagnostics, configuration, and servicing activities to the full extent permitted by the manufacturer. | Mandatory
5 | All technical / maintenance manuals for the proposed system and accessories must be supplied on delivery in both soft and/or hard copy. Provide full details. | Mandatory
6 | Replacement parts and consumables for specified equipment must be readily available for a period of not less than 10 years from the date of acceptance of the equipment. Please confirm agreement. | Mandatory
7 | An engineer digital signed report/transmission will be required after any service call (routine, ad-hoc or emergency). This report must list all tests, investigations and remedial works performed with evidence of test results and any further actions required. Details for call outs must include but are not limited to; - Logged problem or fault occurrence (who , what, when, where?). - Engineer to report on arrival to CHI service team. - Detection to identify root cause for fault or problem. - Any parts replaced or error logs cleared. - Solution implemented and any mitigations to avoid repeat. - Identify any follow-up or further investigation required. - A full electrical safety test is to be carried out after any repairs and copy submitted. A sample service call out report must be included with response. | Mandatory
8 | Please provide details of all Service Contracts on offer for this device and include details on response times, escalation and parts plus accessories covered. | Assessed
9 | Please state how many Irish based service engineers fully trained in service and repair are available in the company | Assessed
10 | Please confirm that all corrective, safety and manufacturer issued mandatory software and firmware updates will be supplied and installed free of charge for the lifetime of the system. HTM must be informed in a timely manner of these updates stating the defect/resolution, time for Engineer to carry out the work, changes to menus etc. and training required if necessary. | Mandatory
11 | The tenderer must state if they propose marketing a new or more technologically advanced product, other than that quoted for in the tender within 12 months from date of purchase. Should a new or more technologically advanced product become available, within 12 months, the supplier must replace the supplied or placed products with the new equipment without any charge to the purchaser. Please confirm agreement. | Mandatory
12 | A minimum of two years' notice will be given before a product or service is declared "Out Of Service" or has reached "End of Life". Please confirm. | Mandatory
13 | Please state the Risk Class of Medical Device (i.e. Class 1, Class 11a, Class 11b or Class 111) applicable to the equipment being offered as part of you tender submission. | Mandatory
14 | Please state the Electrical Safety Type of Medical Device (i.e. Type B or Type BF) | Mandatory
15 | Please detail any adverse incidents involving the equipment which have resulted in reports to the Medical Devices Agency (U.K.), the Food and Drugs Administration (USA) or Health Products Regulatory Authority(IRL). | Mandatory
16 | Tenderers must indicate if there have been any medical device safety notices issued for the system. If so, provide documentation outlining the issue and the corrective actions taken by the manufacturer. | Mandatory
17 | On goods delivery the supplier must unpack and commission the complete system, including full Electro safety assessment of each device and remove all packaging (in accordance with WEEE directive). Confirm compliance. | Mandatory

## 4. Reference Sites
Reference Sites for the Provision of Paediatric Cardiac Ultrasound for the New Children's Hospital, Dublin 8
For the purposes of verifying the information submitted by tenderers as part of this tender process, the Contracting Authority intends to refer to the reference sites provided in this Section and also to customers of the tenderer where the services offered by the tenderer in its tender response are currently being performed. Please note that this exercise is solely to allow Contracting Authority to assess the tenderer’s written tender response in the context of how proposals included in that tender response are being performed in a practical/live environment.
Each tenderer to nominate two reference sites per Lot, of similar size and complexity to the new children's hospital from which one or more site will be contacted / visited in order to evaluate the suitability of the solution tendered.
Tenderers must complete this tab
Tenderers Response | Reference Site # One | Reference Site # Two
Name of Hospital.
Hospital Address.
Make, Model and quantity of equipment provided.
Equipment age profile.
Year Date of installation of equipment in this site.
Clinician contact in this site.
CNM contact in this Reference Site.
Clinical Engineer contact in this Reference Site
Purchasing contact in this Reference Site.
Main clinical application(s) in this hospital.
Note to Tenderers: This tab must be completed

## 5. Infection Control
Infection Prevention and Control Questionnaire for the Provision of Medical Equipment for the New Children's Hospital, Dublin 8
Tenderers must complete this tab
Name of Tenderer
Product Description:
Brand Name/Product Code:
Non Invasive Medical Equipment & Furniture (please tick as appropriate):
Q# |  | Indicate YES or NO
1 | Suitable for cleaning with neutral detergent and water /detergent wipes?
2 | All Ultrasound Imaging Transducers must be suitable for high level disinfection using the Nanosonics Trophon 3 System
3 | In line with the New Children's Hospital green campus initiatives, chlorine dioxide will be used as a disinfectant. Please confirm if the equipment is compatible with chlorine dioxide
4 | Please list all products that are compatible for disinfection purposes.
5 | If chlorine dioxide is not compatible, please state preferred disinfection method
6 | Do cleaning instructions restrict the choice of disinfectants/hypochlorite brands that can be used?
7 | State to what level the complete solution supplied can be cleaned e.g., steam cleaned. Please provide relevant IPX rating including the IPX rating of the hand control to support your submission.
8 | Written instructions on the care, cleaning and maintenance checks for the solution must be submitted with tender. Confirm these are included.
9 | The item should have durable smooth surfaces with low dirt retention and minimal grooves to facilitate easy cleaning?
10 | The item should be made of a material which prevents fluid ingress and able to withstand repeated applications of detergent, water and disinfectant?
11 | Does the tendered solution has easy access to all parts for manual cleaning? If 'Yes' skip Q11
12 | List any parts that cannot be accessed for manual cleaning.
13 | Is there operational training for users provided?
14 | Provide details of operational training provided:
Manufacturers Decontamination Guidelines.
14 | Is the equipment supplied with a user manual stating clear detailed instructions on the methods of decontamination? (Electronic copy of same to be included with tender submission)
15 | Instructions for the care and maintenance of the product from the company are provided?
16 | A life expectancy for the item should be indicated to estimate a time frame for a replacement policy.
17 | Please state device life expectancy
| Education/Training Service | Indicate Yes or No
18 | Is there an educational programme provided
19 | Provide details
Declaration
This document covers relevant infection prevention and control aspects of Non Invasive Medical Equipment cleaning and disinfection. This must be signed and returned together with other relevant information.
Company:
Signature:
Position in company:
Date:
Note to Tenderers: This tab must be completed