APPENDIX 1 : Technical Specification & Supplier Response Document

Tender Reference: NCTR2639

Request for Tenders for the Supply, Delivery, Installation and Commissioning of a Single Quadrupole Gas-Chromatography Mass Spectrometry (GC-MS) System, with coupled Pyrolser Unit for TU Dublin

Open (National) Procedure

Location

Health, Engineering and Material Sciences Building

| Issue Date: | Tuesday 18th August 2026 |
| --- | --- |
| Contact Details for Queries: | E-tenders Messaging facility |
| Closing Date for Receipt of Queries: | 14:00pm on Wednesday 2nd September 2026 |
| Deadline for Receipt of Tenders:. | 14:00pm on Wednesday 9th September 2026 |

Contents

Introduction	3

Summary of Requirements	3

Funding Acknowledgment	4

Section A : Selection Criteria – Mandatory Requirements	5

Section B  - Requirements & Specification	10

Technical Specification	10

Programme Requirements	11

Award Criteria	12

Scoring Methodology	13

Specification Required & Supplier Response	16

Declaration	30

# Introduction

TU Dublin, Ireland’s first Technological University, is where career-focused students, dedicated staff and academic excellence in science, the arts, business, engineering and technology converge to create the leaders of tomorrow. We offer an inclusive and open learning experience with pathways to graduation, from Apprenticeship to PhD. Our 28,500 students learn in a practice-based environment informed by the latest research and enabled by technological advances.

TU Dublin hosts a thriving research community engaged in applying innovation and technology to solve the world’s most pressing challenges. We are deeply committed to collaborating with our national and international academic partners and our many networks in industry and civic society to create new learning experiences and develop impactful research.

TU Dublin is an exciting new milestone in Irish Higher Education.  With campuses in Dublin City, Tallaght and Blanchardstown, it spans the largest population centres of Ireland’s capital city.  Building on the rich heritage of its founding organisations - DIT, IT Blanchardstown and IT Tallaght – TU Dublin is inclusive and adaptable, creating educational opportunities for students at all stages.  TU Dublin students are socially responsible, open-minded global thinkers who are ambitious to change the world for the better. As graduates, they will be enterprising and daring in all their endeavours, ready to play their part in transforming the future.

# Summary of Requirements

Tenders are sought for the supply, installation, commissioning, and training for a Single Quadrupole Gas-Chromatography Mass spectrometry (GC-MS) System, with a coupled Pyrolyser Unit.

The proposed GC-MS system shall be a fully programmable, high-performance instrument suitable for a broad range of analytical applications. The system must incorporate a dual column oven configuration, a single quadrupole mass spectrometry detector, and an autosampler. Additional evaluation marks may be awarded where the proposed autosampler is a connected multifunctional autosampler system capable of performing liquid injection, headspace analysis, and solid phase microextraction (SPME), with automated switching between sampling modes, as assessed in accordance with the published evaluation criteria.

The GC-MS system shall be fully compatible with a the identified pyrolyser unit and the whole system be supplied with all necessary hardware, interfaces, and connections to enable seamless integration and operation. The system must be able to accommodate multiple carrier gases, including nitrogen, helium, and hydrogen. Where necessary for use, a high-purity hydrogen generator, complete with all associated tubing and fittings, shall be included.

The system must incorporate integrated gas leak detection, automatic gas switching, and gas shut-off functionality. A split injector system is preferred. The complete solution must also include a selection of GC columns appropriate for the intended range of analyses and, where required to support the multifunctional autosampler. A nitrogen generator should be supplied to ensure full operational capability upon installation, if required.

The system must be supplied with all relevant computers, software and licenses, tubing, connectors and relevant components/accessories necessary for its function as a single quadrupole GC-MS system with coupled pyrolyser unit.

# Funding Acknowledgment

This procurement is being funded under the Higher Education Authority (HEA) Research and Innovation Equipment Grant, supported by public funding administered by the Higher Education Authority.

## Section A : Selection Criteria – Mandatory Requirements

| REF | PASS/FAIL CRITERIA | PASS REQUIREMENT |
| --- | --- | --- |
| A1 | Applicant Summary | Tenderers must complete this section. If Tenderer is a grouping, then a separate questionnaire must be completed for each group member. |
| A2 | Financial | Tenderers are required to self-declare compliance with the requirements. Tenderers should note that if deemed the most economically advantageous tender, they will be required to provide the evidence self-declared prior to contract award. |
| A3 | Tax Compliance |  |
| A4 | Insurance |  |
| A5 | Previous Experience | Tenderers must supply contact details and information on services provided for two referees to whom you have provided similar services to the requirements of this RFT. |
| A6 | Health & Safety | Tenderers must demonstrate that they comply with the relevant legislation with reference to Environmental, Health and Safety. |
| A7 | Data Protection Compliance | The Tenderer must ensure the security of Personal Data at all times is in line with the Data Protection Acts 1998 and 2003 and guidance issued by the Data Protection Commissioner of Ireland. Tenderers are required to be compliant under the EU General Data Protection Regulation (GDPR) which came into effect on 25/05/2018. Please confirm compliance. |
|  |  | Fully completed Appendix 1 – Supplier Response Document and Appendix 2 – Pricing Schedule must be submitted. |

SECTION A: PASS/FAIL CRITERIA

| A1. APPLICANT SUMMARY Weighting: Pass/Fail only Pass requirement: Tenderers must complete this section |  |
| --- | --- |
| Organisation Name | Insert Response |
| Contact Name | Insert Response |
| Address | Insert Response |
| Telephone | Insert Response |
| Email | Insert Response |
| Legal Status (if any) (e.g., Company, Partnership, Sole Trader, etc.) | Insert Response |

A2/A3/A4. FINANCIAL/ TAX COMPLIANCE / INSURANCE Weighting: Pass/Fail only Pass requirement: Tenderers must confirm that the meet the following minimum rules

| A2. Financial Summary |  |
| --- | --- |
| Requirement | Yes / No |
| The company is solvent and of good economic standing and can if requested supply an accountant’s/ auditor’s statement for the last three financial years. Availability of a signed statement from accountant/auditor confirming an annual company turnover of €220,000 for each of the last 3 years. | Choose an item. |

| A3. Tax Compliance |  |
| --- | --- |
| Requirement | Yes / No |
| Tenderers hold a current and valid Tax Clearance Certificate(s) from the Revenue Commissioners or in the case of a Candidate outside of the Republic of Ireland hold or can provide an equivalent statement from the relevant taxation authority. | Choose an item. |

| A4. Insurances |  |
| --- | --- |
| Requirement | Yes / No |
| That Tenderers holds, or will hold should they be successful, the following insurances. TU Dublin reserves the right to amend the requirements before a Purchase Order is issued. |  |
| Employer’s Liability - €13 million limit for any one claim or series of claims arising out of a single occurrence. | Choose an item. |
| Public Liability - €6.5 million limit for any one claim or series of claims arising out of a single occurrence. | Choose an item. |
| Product Liability - €6.5 million limit for any one claim or series of claims arising out of a single occurrence. | Choose an item. |

Documentary evidence of economic and financial standing is not required at this stage of the Competition. Instead, Candidates are asked to declare that they meet the minimum standards required by the Contracting Authority and will be in a position to produce the necessary documentation (e.g., bank statements, audited accounts, proof of insurance, etc.) if prior to any award of contract.

| The following are a requirement of each Applicant |  |
| --- | --- |
| Requirement | Yes / No |
| Tenderers must submit a signed Tenderers statement.as per Appendix 3 of the RFT | Choose an item. |
| Tenderers must submit a Signed Declaration as to Personal Circumstances. as per Article 45 of Directive 2004/18/EC & Regulation 53 of SI329); Appendix 4 | Choose an item. |

A5: PREVIOUS EXPERIENCE/REFERENCES

The tenderer MUST have a track record of successful supply and installation of similar systems to those being tendered for.

The Tenderer MUST provide the relevant reference sites of where two (2) similar systems related to that being tendered for, have been successfully installed.

Please complete tables below for each comparable example.

| PREVIOUS EXPERIENCE (CONTRACT ONE) |  |  |  |
| --- | --- | --- | --- |
| Title of Project |  |  |  |
| Client Name |  |  |  |
| Client Contact Details | Contact Name |  |  |
|  | Contact E-mail Address |  |  |
|  | Contact Telephone Number |  |  |
| Annual Value |  |  |  |
| Start Date |  | End Date |  |
| Please provide a summary of the service provided and indicate how it is similar to the goods being sought by TU Dublin. |  |  |  |
| Please insert details here |  |  |  |
| PREVIOUS EXPERIENCE (CONTRACT TWO) |  |  |  |
| Title of Project |  |  |  |
| Client Name |  |  |  |
| Client Contact Details | Contact Name |  |  |
|  | Contact E-mail Address |  |  |
|  | Contact Telephone Number |  |  |
| Annual Value |  |  |  |
| Start Date |  | End Date |  |
| Please provide a summary of the service provided and indicate how it is similar to the goods being sought by TU Dublin. |  |  |  |
| Please insert details here |  |  |  |

| A6. Health & Safety |  |
| --- | --- |
| Requirement | Yes / No |
| Candidates must demonstrate that they comply with the relevant legislation with reference to Environmental, Health and Safety. | Choose an item. |

A7: DATA PROTECTION COMPLIANCE

Compliant Yes/No:. Signature:. Tenderers are required to be compliant under the EU General Data Protection Regulation (GDPR) replacing the Data Protection Directive 95/46/EC which came into effect on 25/05/2018. Please confirm compliance.

PLEASE NOTE

In the first instance, all submissions will be checked to ensure they meet the Mandatory Pass/Fail requirements in Section A in order to determine compliancy and responsiveness.

All submissions that meet the mandatory pass/fail criteria will be deemed to be compliant and responsive and will be assessed under the weighted criteria set out in Section B.

Any submission that does not meet the mandatory pass/fail requirements will be deemed to be non-compliant and non-responsive and will be eliminated.

In relation to minimum score required - Tenderers should note that they must achieve a minimum rating set out for each of the individual qualitative criteria to avoid elimination from the competition.

## Section B  - Requirements & Specification

# Technical Specification

This is the minimum specification required.  All tenderers must meet or exceed this minimum specification to progress through the award criteria to cost evaluation.

Suppliers MUST provide all information relating to the instrument and the associated consumables in the response boxes below. Responses will be scored accordingly by the Technical Evaluation team as set out in the Award Criteria in the RFT document.

Tenderers must use “Appendix 1: Requirements and Specifications to structure their response on how they meet the requirements.

Please ensure all cost submissions are kept separate and included with your submission in the Excel file entitled: Appendix 2 – Pricing Schedule.

Notes

- Where specifications include reference to manufacturer or brand names this is a result of difficulties in specifying the product required. In these instances the phrase “or equivalent” is inferred.

- Compliance with all the standards as regards health and safety, quality assurance and analytical protocols is expected.

- It is acknowledged that each Tender’s product line may differ from these specifications.  As such, each seller is free to propose variances from these specifications. Alternative proposals favourable to the Contracting Authority may be suggested. It is required, however, that, whenever a variance from these specifications occurs, the proposed item must meet or exceed the specified characteristics or level of performance.

- Additional functionality should be identified in the comments section to facilitate scoring under the award criteria.

Where the bidder has more than one system matching the specifications, details for each system should be provided separately and is limited to 2 per tenderer.

Ex-Demo Goods/Instrument Models

Proposals for the provision of ex-demo model Goods/Instrument(s) are not acceptable to TU Dublin in this instance.

# Programme Requirements

- Tenderers MUST note that contractors do not have exclusive access and will be required to work around and accommodate other suppliers and contractors who will be active in the building and laboratories simultaneously.

- TU Dublin working hours are Monday to Friday from 9.00am to 5.00 pm (Business Hours). Access to the site outside the Business Hours shall at all times be arranged through the appointed representative in TU Dublin.

- Note: Working outside Business Hours may be required to comply with work programmes. Where this becomes necessary as a result of changes to the start dates set out above, all reasonable efforts will be made to accommodate same.

- Contractors are responsible for security of their own equipment, materials, etc.

- Prior to any deliveries, the successful supplier(s) (and/or  relevant sub-contractors, etc.) will be required to meet the representatives from TU Dublin to clarify the requirements for Health & Safety, site access, services etc. to ensure compliance with standard operating procedures, and other regulatory/ statutory requirements.

# Award Criteria

The Goods Contract will be awarded on the basis of the most economically advantageous tender(s) as identified in accordance with the following award criteria:

| Award Criteria |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | Award Criteria: Tenderers must meet the minimum score of 60% for Sections B - E | Overall Weight | Total Maximum Score | Minimum Score (60%) Requirement |
| B | Technical Merit - Technical Requirements of the System including Software Requirements | 35% | 350 | 210 |
| C | After Sales and Technical Assistance - Customer Support, Repair, Maintenance & Breakdown Requirements - Warranty & Preventative Maintenance - Training Proposal | 15% | 150 | 90 |
| D | Delivery Date or Date of Completion - Lead Time & Delivery/Completion Plan - Installation, Commissioning, Calibration & Validation. | 10% | 100 | 60 |
| E | Environmental & Sustainability - Energy Consumption Emissions - Life Cycle & Waste | 10% | 100 | 60 |
| F | Ultimate Cost (Total Cost of Ownership) Including - Capital Cost and all Delivery, Installation, Training, Warranty, Maintenance & Preventative Maintenance Costs | 30% | 300 | n/a |
|  | Total Maximum Score Available 100% |  | 1000 |  |

The tender award criteria will be applied to each tender who has successfully passed the Selection Criteria. Tenderers must achieve a minimum of 60% on all qualitative tender award criteria above, excluding Ultimate Cost, in order to avoid elimination from the competition. Only tenderers who have achieved the 60% will be evaluated against Ultimate Cost.

No commitment of any kind, contractual or otherwise shall exist unless and until a formal written contract has been executed by or on behalf of the Contracting Authority. Any award of notification of preferred bidder status by the Contracting Authority shall not give rise to any enforceable rights by the Tenderer. The Contracting Authority may cancel this public procurement competition at any time prior to a formal written contract being executed by or on behalf of the Contracting Authority. The Contracting Authority does not bind itself to accept the lowest priced or any Tender.

A failure to answer question(s) may leave the contracting authority with insufficient information on which to base a thorough evaluation and in turn may result in loss of marks in respect of the relevant Award Criterion. Therefore, please ensure you complete ALL sections of this questionnaire.

TU Dublin reserves the right to purchase up to an additional one (1) item  should such a need arise within eighteen (18) months commencing from the Award Notification date.

# Scoring Methodology

The lowest Ultimate Cost tender, will receive the maximum score achievable under this criterion.

| Pricing Score = | Lowest Tendered Price x Max Points |  |
| --- | --- | --- |
|  | Price of bid being evaluated |  |

Responses for Non-cost/qualitative Award Criteria items will be scored using the following matrix:

| - Band | Scoring Range | Meaning |
| --- | --- | --- |
| Excellent | 91% – 100% | Excellent response that fully meets and exceeds requirements, providing comprehensive, detailed, and convincing assurance that the Tenderer will deliver to an exceptional standard. |
| Very Good | 80% – 90% | A very good response that demonstrates real understanding and fully meets the requirements and provides convincing assurance that the Tenderer will deliver to high standard. |
| Good | 60% – 79% | A satisfactory response which demonstrates a reasonable understanding of requirements and provides reasonable confidence that the minimum requirements will be delivered. However, it does not provide sufficiently convincing assurance to award a higher mark. |
| Fair | 30% - 59% | A response where reservations exist due to a lack of convincing details and comprehensive credibility. There is a significant risk that the requirements will not be delivered to a satisfactory standard. |
| Poor | 1% – 29% | A response that raises significant reservations, due to insufficient detail, critical or notable deficiencies, that seriously lacks credibility, presenting a high risk of non-delivery. |
| 0 | 0% | No Response or partial Response that completely fails to address the criterion under consideration. |

Suppliers please note:

- This specification document is designed to assist the contracting authority in its evaluation of your tender and to enable suppliers provide structured responses to key deliverables. It is structured to reflect the award criteria as set out in the SRFT.

- Please note responses to individual questions will NOT be scored separately but collectively they will provide a basis for assessing the merits of the tenderer’s overall response to each award criterion and sub-criterion.

- Compliance with all the standards as regards health and safety, quality assurance and analytical protocols is expected.

- Where abnormally low tenders are received the Evaluation team shall require the tenderer to explain the price proposed in Appendix 2 (per Article 69 of the EU Procurement Directives 2014/24). Where the tenderer cannot provide a sufficient explanation, the contracting authority is entitled to reject the tender.

- Next Ranked Supplier - The award of any contract is subject to the successful tenderer meeting the required specification and standards as set out in this RFT and Appendix 1. Should the required specification and standards not be met, the Contracting Authority reserves the right to award any contract to the next highest-ranked Tenderer.

- Tie Break - If the evaluation outcome results in a tie between two or more Tenders, then the Tender with the highest score for the highest weighted criterion shall be deemed the most economically advantageous tender. If the score remains in a tie, the remaining criterion will be assessed by the highest weighted until there is no longer a tie. If the score remains a tie after all criterions are assessed, the tenderers will be asked to submit a Best and Final Offer (BAFO).

# Specification Required & Supplier Response

Suppliers MUST provide all information relating to the instrument and the associated consumables in the blue response boxes below. Responses will be scored accordingly by the Evaluation team as set out in the Award Criteria in the RFT document. Yes/No answers will not suffice.

| NCTR2639 - Request for Tenders for the Supply, Delivery, Installation and Commissioning of a Single Quadrupole Gas-Chromatography Mass Spectrometry (GC-MS) System, with coupled Pyrolser Unit for TU Dublin |  |  |  |  |
| --- | --- | --- | --- | --- |
| (Suppliers to state model type here) |  |  |  |  |
| REF. |  | REQUIREMENT | CONFIRM (YES/NO) | COMMENT |
| B. Technical Merit |  |  |  |  |
| Functional Requirements |  |  |  |  |
| B1 | A gas chromatography instrument coupled to a single quadrupole mass spectrometry detector is required. Please confirm and provide details of the proposed instrument. |  | Compliant (Yes / No) | Insert Response |
| B2 | The proposed instrument must have an autosampler, with the preferred option being a multifunctional autosampler to include, but not specific to, liquid, headspace (vapour) and solid phase microextraction (SPME) methods. A gas generator for the multifunctional autosampler must be included, if required for operation. Rough vacuum pump and attachments must be included. Please confirm and detail, indicating the cost, if any, in Appendix 2 – Pricing Schedule. |  | Compliant (Yes / No) | Insert Response |
| B3 | The proposed instrument must be supplied with a fully compatible Pyrolysis Unit with all necessary connections for fully functional linking of the pyrolyser system to the GC-MS system. Versatile operating modes for full material characterisation is required, including evolved gas analysis, thermal desorption, flash pyrolysis, heart-cutting and double-shot modes. Please confirm and detail, indicating the cost, if any, in Appendix 2 – Pricing Schedule. |  | Compliant (Yes / No) | Insert Response |
| B4 | The proposed instrument must have fully programmable dual column oven. Please confirm and details. |  | Compliant (Yes / No) | Insert Response |
| B5 | The proposed system must be supplied with a selection (Minimum 3) of generic GC-columns for use such as, a 5% Phenyl / 95% Dimethyl Polysiloxane column, a Polar and Nonpolar column, that are at least 30 meters long, ID 0.25mm with ferrules and fittings to attach to the instrument. Please confirm and detail, indicating the cost, if any, in Appendix 2 – Pricing Schedule. |  | Compliant (Yes / No) | Insert Response |
| B6 | The proposed system must be capable of taking different carrier gases such as nitrogen, helium and hydrogen. A high purity hydrogen generator, with all necessary high-quality tubing, must be included as required for the operation of the system. Similarly for nitrogen gas, if required. Please confirm and detail, indicating the cost, if any, in Appendix 2 – Pricing Schedule. |  | Compliant (Yes / No) | Insert Response |
| B7 | Gas leak testing capability, gas switching and gas cut-off must be integrated in the system. Please confirm and detail. |  | Compliant (Yes / No) | Insert Response |
| B8 | The proposed system must be supplied with a split/splitless injection capability and full functionality and compatibility for liquid sample injection (in addition to pyrolysis functionality). Please confirm and detail. |  | Compliant (Yes / No) | Insert Response |
| B9 | The proposed system must be supplied with a mass analyser with capability for operation in full scan mode (untargeted analysis) and selected ion monitoring (SIM) mode (targeted analysis). Please confirm and detail. |  | Compliant (Yes / No) | Insert Response |
| B10 | The proposed system must be supplied with Electron Ionisation (EI) as standard, with optional configuration for other ionization modes (e.g. Chemical Ionisation) which can be upgraded. Please confirm and detail. |  | Compliant (Yes / No) | Insert Response |
| B11 | The proposed system must be capable of a mass range suitable for routine organic compound analysis (typically up to/ > m/z 600 or higher) and a sensitivity capable of trace-level detection. Please confirm and detail. |  | Compliant (Yes / No) | Insert Response |
| B12 | The tenderer must provide comprehensive details of all detection capabilities of the proposed instrument, including, but not limited to, limits and types of analyte detection. The response must include a complete list of operational modes, together with their corresponding performance specifications, including accuracy, precision, and any applicable reference standards or validation methods. |  | Compliant (Yes / No) | Insert Response |
| B13 | The proposed System must meet CE certification and EU safety compliance. Please confirm and provide details. |  | Compliant (Yes / No) | Insert Response |
| B14 | The proposed system must be supplied with a computer for seamless control of all components. The tenderer must specify the details of the PC requirements (hardware and software) required for the proposed instrument to operate in its entirety. Please confirm and detail. |  | Compliant (Yes / No) | Insert Response |
| B15 | The proposed system should be flexible and allow for purchase of additional functionality in the future, to keep up to date with developments in the field. Please confirm and provide details. |  | Compliant (Yes / No) | Insert Response |
| B16 | Please provide details of any additional technical features or specifications that may offer a performance advantage. Such features shall be clearly described, including an explanation of how they enhance the overall capability, efficiency, or quality of the instrument. Please provide details indicating the cost associated, if applicable, under additional purchases in Appendix 2 – Pricing Schedule. |  | Compliant (Yes / No) | Insert Response |
| System Software Requirements |  |  |  |  |
| B17 | The system must be provided with all necessary hardware and software licences required for full system control, along with running and processing of the acquired data. The system must provide a perpetual software to enable complete operation of the instrument and subsequent analysis of the data. Please confirm and provide details. |  | Compliant (Yes / No) | Insert Response |
| B18 | The system must include a minimum of one (1) software licence for data analysis on a remote computer. All/any third-party software required to operate the system must be supplied. Please confirm and detail, indicating the cost, if any, in Appendix 2 – Pricing Schedule. |  | Compliant (Yes / No) | Insert Response |
| B19 | Latest software updates available at the time of system commissioning must be included/installed. All software updates over the lifetime of the equipment must be made available within the package price provided to TU Dublin. Continued development of the software should be demonstrated by the supplier, with facilities for updating systems in use. Please confirm and provide details. If not, please clearly detail any exclusions, limitations, or additional charges, indicating the cost, if any, in Appendix 2 – Pricing Schedule. |  | Compliant (Yes / No) | Insert Response |
| B20 | The system package must be easy to use/user-friendly, intuitive and accessible while the system is running. Please confirm and detail. |  | Compliant (Yes / No) | Insert Response |
| B21 | The tenderer must provide details on how data can be exported from the proposed instrument/software. The exported data must be in standard formats such as CSV, Excel, PDF, etc. Please confirm and detail. |  | Compliant (Yes / No) | Insert Response |
| B22 | IT networking: The system vendor must specify any IT infrastructure required to support the instrument package, including but not limited to, any wired and WiFi network specifications. Please confirm and provide details. |  | Complaint (Yes/No) | Insert Response |

| - After Sales and Technical Assistance |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Customer Support, Repair, Maintenance & Breakdown Requirements |  |  |  |  |  |
| C1 | The tenderer must have dedicated specialists at their disposal and fully qualified service personnel who can respond to service calls. Please confirm and provide details of access to technical support (remote and/or on-site) available to TU Dublin. | Compliant (Yes / No) |  | Insert Response |  |
| C2 | The tenderer must provide details of guaranteed response times in event of instrument issues or breakdown, in relation to phone support, online/remote support and arrival on site by a fully trained service engineer during the warranty period (s) and post-warranty. | Compliant (Yes / No) |  | Insert Response |  |
| C3 | The tender must specify the lead time for repairs during the warranty period(s) and post-warranty, with and without service contract options. | Compliant (Yes / No) |  | Insert Response |  |
| C4 | Please outline the company customer service plan and your escalation policy and process in place for escalating a call-out, if a breakdown cannot be resolved on the first visit. | Compliant (Yes / No) |  | Insert Response |  |
| C5 | The tenderer must provide confirmation of the availability of service and replacement parts for at least seven years (7) years after commissioning. Please provide details. | Complaint (Yes/No) |  | Insert Response |  |
| C6 | The tenderer must provide full details of service and support infrastructure. Customer service contact details and availability must also be provided. Information including, but not limited to, the names, telephone numbers and qualifications of contact individuals in respect of scientific, sales, technical and IT personnel available to service and support the tendered instrument will be requested from the successful tenderer in advance of any award. Please confirm and detail. | Complaint (Yes/No) |  | Insert Response |  |
| Warranty & Preventative Maintenance |  |  |  |  |  |
| C7 | One (1) year’s warranty must be included as part of the tendered price, to include all hardware, software, call outs, travel, and parts. Please detail specific inclusions and exclusions of warranty. Please confirm and provide details indicating additional pricing, if any, in Appendix 2 -Pricing Schedule. |  | Compliant (Yes / No) |  | Insert Response |
| C8 | The tenderer should state the details of yearly servicing contracts offered as part of any extended Warranty. Detail specific inclusions and exclusions of extended warranty for up to three (3) years optional to purchase extended warranties on a yearly basis. Please include price per annum, details of coverage and maximum term of coverage, and indicating the cost in Appendix 2 – Pricing Schedule. |  | Compliant (Yes / No) |  | Insert Response |
| C9 | The warranty must include one (1) preventative maintenance visit per year. The tenderer must state details of annual service, calibration and preventative maintenance schedule included along with a recommended 4-year preventative maintenance schedule. Detail specific inclusions and exclusions must be clearly outlined. Please confirm annual price in Appendix 2 – Pricing Schedule including the option to purchase additional preventative maintenance for up to three (3) years, on a yearly basis. |  | Compliant (Yes / No) |  | Insert Response |
| C10 | The Tenderer must provide typical daily rate price for servicing and repairs (including travel, subsistence, and price per day) charged outside a maintenance contract. Please confirm pricing in Appendix 2 – Pricing Schedule. |  | Compliant (Yes / No) |  | Insert Response |
| C11 | The tender must specify the lead time for repairs during the warranty period(s) and post-warranty, with and without service contract options. |  | Compliant (Yes / No) |  | Insert Response |
| C12 | If, during the warranty period, the model does not meet the criteria specified, TU Dublin must have the right to return the instrument for a full refund of the cost of the instrument. Please confirm your agreement. |  | Compliant (Yes / No) |  | Insert Response |

| Training Proposal |  |  |  |
| --- | --- | --- | --- |
| C13 | Training for at least 4 personnel must be included after commissioning to cover set up / system familiarization, operation of the tendered instrument, troubleshooting, maintenance, and associated software once the system is in place. A detailed description of the training and support to be provided for staff and technical personnel is required. (Date(s) and timing of training to be confirmed with TU Dublin and the winning tenderer). Please confirm and detail, indicating the cost, if any, in Appendix 2 – Pricing Schedule. | Compliant (Yes / No) | Insert Response |
| C14 | The tenderer should provide details of any optional refresher training offered. Remote and on-site options should be detailed where available. Please indicate the cost, if any, under additional purchases in Appendix 2 – Pricing Schedule. | Compliant (Yes / No) | Insert Response |
| C15 | All relevant training and maintenance manuals in English must also be provided with the instrument. Please confirm and detail. | Compliant (Yes / No) | Insert Response |
| C16 | Training materials must include online access to comprehensive training materials (videos, digital resources). Please confirm and provide details. | Compliant (Yes / No) | Insert Response |
| - Delivery Date or Date of Completion |  |  |  |
| Lead Time & Delivery/Completion Plan |  |  |  |
| D1 | Delivery must include all Shipping Charges, insurance, customs duties and all aspects as outlined in DDP (Delivery Duty Paid) incoterms. Instrument must be delivered to: Health, Engineering and Material Sciences Building, Lab 1.03 Tallaght Campus Belgard Road Tallaght Dublin D24 FKT9 Please confirm cost, if any, in Appendix 2 - Pricing Schedule. | Compliant (Yes / No) | Insert Response |
| D2 | Instrument must be fully installed and operational no later than 30th November 2026. Please confirm. Tenderers must provide an appropriate lead time and delivery plan including installation and commissioning, sign off and acceptance, from PO placement, in weeks, for the required instrument. Details of timeline for the lead-time to installation, commissioning and training must be provided. Shorter lead times are highly desirable. Please confirm and detail. | Compliant (Yes / No) | Insert Response |
| D3 | The tenderer must provide details of the logistics/supply chain to the point of customer delivery, i.e. where the goods will come from, how the goods will be transported and what borders the goods will cross. Please confirm and detail. | Compliant (Yes / No) | Insert Response |
| D4 | The tenderer must provide details of their contingency plans (alternative sources/routes of supply) for any delays in delivery. Please confirm and detail. | Compliant (Yes / No) | Insert Response |
| D5 | The tenderer must provide details of any actions that have been taken to ensure the future timely supply and delivery of goods and services, including ongoing supply of replacement/ spare parts and consumables. Please confirm and detail. | Compliant (Yes / No) | Insert Response |
| D6 | The tenderer must provide details of the effects of any regulation change, if there are any, to the purchase of these goods. Please confirm and detail. | Compliant (Yes / No) | Insert Response |
| Installation, Commissioning & Validation |  |  |  |
| D7 | The instrument must be installed (if required) and commissioned in a location defined by the Contracting Authority, by an appropriately qualified engineer on an agreed date. Instrument must be fully installed and operational no later than 30th November 2026. Please confirm. Please confirm, providing details of experience/qualifications of installation team and indicating the cost, if any, in Appendix 2 – Pricing Schedule. | Compliant (Yes / No) | Insert Response |
| D8 | Tenderers must provide an appropriate installation plan (if required) for the instrument. Details of timeline for the lead-time to installation, commissioning and training (if applicable) must be provide. Please confirm and detail. | Compliant (Yes / No) | Insert Response |
| D9 | Should responsibilities performed by installation teams, training facilitators or service engineers be affected during a global crisis, please outline alternative solutions that are in place from your standard operating procedures. Please confirm and detail. | Compliant (Yes / No) | Insert Response |
| D10 | The instrument must be ready for use once delivered and installed. Following the installation of all instrument in the relevant area, the successful tenderer will be required to fully demonstrate the correct operation of all of the equipment to the satisfaction of a nominated employee. Please confirm and detail. | Compliant (Yes / No) | Insert Response |
| D11 | The installation of all electrical items must be in accordance with but not limited to European standards/ regulations and all relevant amendments. Installation and commissioning of the equipment must be carried out by the fully qualified and trained engineers. Please confirm and provide details. | Compliant (Yes / No) | Insert Response |
| D12 | All instrument documentation relating to service and maintenance, parts list, and operation must be supplied in English language. Please confirm and provide details. (Documentation is not required as part of your tender submission). | Compliant (Yes / No) | Insert Response |
| - Environmental & Sustainability |  |  |  |

Energy Consumption & Emissions

| E1 | Tenderers must state the energy rating of the proposed system. Your response should include details of the heat level generated, if the instrument requires additional cooling/ventilation, the range of environmental conditions in which it will operate and the energy consumption per hour. The system should ideally have low running costs allowing the instrument to be as energy and economically efficient as possible. Please confirm and detail. | Compliant (Yes / No) | Insert Response |
| --- | --- | --- | --- |
| E2 | Tenderers should demonstrate efforts to minimise adverse environmental impact in the methods of services delivery and in materials used. Please confirm and detail stating actions have you taken to reduce your impact on the environment? | Compliant (Yes / No) | Insert Response |
| E3 | Please confirm the expected lifespan of your equipment and for how long software updates will be available. | Compliant (Yes / No) | Insert Response |
| E4 | The tendered must provide details of any standby or low-power operational modes available on the proposed equipment that may be utilised during periods of non-use or reduced activity. The tenderer must also provide the energy consumption and/or energy efficiency ratings of the equipment under both fully operational and standby conditions, including any relevant power-saving features. | Compliant (Yes / No) | Insert Response |
| Life Cycle and Waste |  |  |  |
| E5 | Tenderers must dispose of all Packaging & Waste, using environmentally friendly sources, after installation and commission has been completed. The instrument should comply with Directive 2012/19/EU (WEEE) relating to waste electrical and electronic equipment. Tenderers should be a registered producer of electrical and electronic equipment or shall supply the equipment via a registered producer compliant with Irish WEEE requirements. Please confirm and detail. | Compliant (Yes / No) | Insert Response |

| - Ultimate Cost |  |
| --- | --- |
| F1 | Please see Appendix 2 – Pricing Schedule for details. |

# Declaration

This form must be completed and signed by an authorised officer of the tenderer’s organisation.

- This offer will remain open for acceptance by you for a period of 12 months from the date of deadline for submission of tenders.

- I/We acknowledge that you are not obliged to accept the lowest or any offer.

- I/We undertake to deliver the products/services in accordance with the terms and conditions of the tender specification.

- I/We undertake to maintain full confidentiality with regard to all aspects of this tender process

| Signed: |  | Position: |  |
| --- | --- | --- | --- |
| Print Name: |  | Phone No: |  |  |
| Company Name: |  | Date: |  |  |
| Address: |  |  |  |  |
| Email: |  |  |  |  |

End of Document