| TENDER RESPONSE DOCUMENT |  |
| --- | --- |
| Open Procedure |  |
| Award of a Contract for | Supply, Delivery, Installation and Commissioning of an Environmental Chamber for DCU |
| Name of Tenderer |  |
| Contact Name |  |
| Email address |  |
| Phone Number |  |

CONTENTS

INSTRUCTIONS FOR COMPLETION	3

1.	RESPONSE TO SELECTION CRITERIA	4

2.	RESPONSE TO THE AWARD CRITERIA	16

# INSTRUCTIONS FOR COMPLETION

Tenderers must ensure they complete this document in full.

Tenderers should ensure they have read the Invitation to Tender before they attempt to complete this Tender Response Document.

If you consider that the Tender Response Document is missing any sections which would prevent you from preparing a comprehensive response, please contact the Contracting Authority as soon as possible.

Where there is a discrepancy between the contents / instructions in this document and the Invitation to Tender, the Invitation to Tender will take precedence. Please ensure you contact the Contracting Authority, and they may decide to issue an updated Tender Response Document.

Tenderers must follow the instructions contained in this document in relation to attachments, word limits, formats for submissions etc.

This document covers the selection and award criteria as relating to the open procedure.

Tenderers must complete all sections. However, tenderers should note they must meet and pass all of the rules contained in the selection criteria, in order to be considered eligible for assessment against the award criteria.

Where tenderers are relying on other parties to meet the selection criteria, those parties must be available to deliver elements of the contract and be clearly identified throughout the response.

# RESPONSE TO SELECTION CRITERIA

## GENERAL INFORMATION RELATING TO THE TENDERER

| Weighting: Pass/Fail Pass Requirement: Tenderers must complete this Section |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Organisation Name: |  |  |  |  |  |  |  |  |  |
| Contact Name: |  |  |  |  |  |  |  |  |  |
| Position: |  |  |  |  |  |  |  |  |  |
| Address: |  |  |  |  |  |  |  |  |  |
| Telephone (Office): |  |  |  |  |  |  |  |  |  |
| Telephone (Mobile): |  |  |  |  |  |  |  |  |  |
| Email: |  |  |  |  |  |  |  |  |  |
| Website: |  |  |  |  |  |  |  |  |  |
| VAT Registration Number: |  |  |  |  |  |  |  |  |  |
| Date of Establishment: |  |  |  |  |  |  |  |  |  |
| Legal Status (Company, Sole Trader, etc.): |  |  |  |  |  |  |  |  |  |
| Is the Tenderer a Small and Medium Enterprise (SME)? | Yes |  |  |  | No |  |  |  |  |
| SMEs are enterprises which employ fewer than 250 persons, and which have an annual turnover not exceeding €50 million, and/or an annual balance sheet total not exceeding €43 million |  |  |  |  |  |  |  |  |  |
| Is the tenderer a consortium / group of Economic Operators? | Yes |  |  |  | No |  |  |  |  |
| If yes, please provide the following information |  |  |  |  |  |  |  |  |  |
| Company / Party Name |  | Address |  | Element being relied upon - Selection criteria such as turnover or previous experience - Elements of contract delivery |  |  |  |  |  |
| Confirmation that relevant information is provided for all 3rd parties. |  |  |  | Yes |  |  | No |  |  |

## CONTRACT TERMS

| I confirm that I have read and accept DCU’s Terms and Conditions outlined in Appendix 1 of the Invitation to Tender, and agree to be bound by them, if successful. | Yes |  | No |  |
| --- | --- | --- | --- | --- |
| Name of Economic Operator |  |  |  |  |
| Authorised Signatory |  |  |  |  |
| Title |  |  |  |  |
| Mobile No |  |  |  |  |
| Email |  |  |  |  |
| Signature |  |  |  |  |

## FINANCIAL INFORMATION

| Weighting: Pass/Fail Pass Requirement: Tenderers must complete this Section |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Tax Compliance |  |  |  |  |  |  |  |  |
| Option 1: I confirm and declare being tax compliant. The Contracting Authority can verify your tax clearance status through Revenue’s online facility at http://www.revenue.ie/en/online/tax-clearance.html To this end, please confirm: |  |  | Yes |  |  |  |  |  |
|  |  |  | No |  |  |  |  |  |
| Tenderer Name: |  |  |  |  |  |  |  |  |
| Tax Reference Number (PPSN): |  |  |  |  |  |  |  |  |
| Access Number: |  |  |  |  |  |  |  |  |
| Option 2: I confirm that I hold a current valid Tax Clearance Certificate (generally relates to Non-Residents) To this end, please confirm: |  |  | Yes |  |  |  |  |  |
|  |  |  | No |  |  |  |  |  |
| Registration Number: |  |  |  |  |  |  |  |  |
| Certificate Number: |  |  |  |  |  |  |  |  |
| Option 3: I confirm that I have applied for Tax Clearance status or a Tax Clearance Certificate which will be made available on request |  |  | Yes |  |  |  |  |  |
|  |  |  | No |  |  |  |  |  |
| Financial Capacity |  |  |  |  |  |  |  |  |
| I confirm that our turnover exceeded € 500,000 per annum in one of the last three financial years. If turnover figures for 2025 are unavailable, please use 2024, 2023 and 2022. |  |  |  |  |  |  |  |  |
| Financial Year |  | 2025 | 2024 |  |  | 2023 |  |  |
| Turnover |  | € | € |  |  | € |  |  |
| Month of financial year end |  |  |  |  |  |  |  |  |
| I confirm that I will provide the following promptly on request at any time prior to the tender list being finalised: (Evidence of turnover for the past three financial years) |  |  |  |  | Yes |  | No |  |
| I confirm that I will supply evidence of financial standing, by way of auditor’s letter, ensuring the tenderer party has the financial capacity to pay its debts identified on the current statement of assets and liabilities as being the debts as they fall due. |  |  |  |  | Yes |  | No |  |
| Insurances |  |  |  |  |  |  |  |  |
| I confirm that we have the following insurances in place: |  |  |  |  |  |  |  |  |
| Insurance Type |  | Level in place |  | Details of Excess |  | Expiry Date |  |  |
| Employer Liability |  | € |  |  |  |  |  |  |
| Public Liability |  | € |  |  |  |  |  |  |
| Product Liability |  | € |  |  |  |  |  |  |
| I confirm that I will provide the following promptly on request at any time prior to the award decision being made: Evidence of insurances in place or Letter from Insurance Broker confirming that the required levels could be put in place if successful |  |  |  | Yes |  |  |  |  |
|  |  |  |  | No |  |  |  |  |
| I confirm that if successful, where the levels required under the contract or framework are higher than those currently in our possession, I will be in a position to put the required forms and levels of insurances required in place promptly. |  |  |  | Yes |  |  |  |  |
|  |  |  |  | No |  |  |  |  |
| Please note that the Contracting Authority will seek to verify self-declarations regarding financial capacity prior to next stage of the competition. |  |  |  |  |  |  |  |  |
| Signed |  |  |  |  |  |  |  |  |
| Position |  |  |  |  |  |  |  |  |
| Date |  |  |  |  |  |  |  |  |

## DECLARATION OF BONA FIDES & STATUTORY OBLIGATIONS

| Declaration of Bona Fides & Statutory Obligations (Under Public Sector Directive 2014/24/EU - Article 57) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Weighting: Pass/Fail Pass requirement: Tenderers must complete, sign and date this Declaration. The Contracting Authority reserves the right at its discretion to exclude a non-compliant tenderer under each heading. This must be completed by each group member. |  |  |  |  |  |
| Economic Operators will be excluded from the procurement process if, within the past five (5) years, there is evidence of a conviction relating to a specific criminal offence listed below (see 1.1) or if they have been the subject of a binding legal decision which found a breach of legal obligations to pay tax or social security contributions (see 1.2) (except where this is disproportionate e.g. where only minor amounts are involved). |  |  |  |  |  |
| 1.1 | Has the Economic Operator or a member of their proposed consortium, (if applicable), Director, or Partner or any other person who has powers of representation, decision, or control, been convicted of any of the following offences? |  | Yes | No |  |
|  |  |  | Please indicate your answer by marking ‘X’ in the relevant box |  |  |
| 1.1a | participation in a criminal organisation, as defined in Article 2 of Council Framework decision 2008/841/JHA; |  |  |  |  |
| 1.1b | corruption, as defined in Article 3 of the Convention on the fight against corruption involving officials of the European Communities or officials of Member States of the European Union and Article 2(1) of Council Framework Decision 2003/568/JHA as well as corruption as defined in Irish Law or the jurisdiction in which the Economic Operator is established; |  |  |  |  |
| 1.1c | fraud within the meaning of Article 1 of the Convention on the protection of the European Communities’ financial interests; |  |  |  |  |
| 1.1d | the subject of a conviction for terrorist offences or offences linked to terrorist activities or for inciting, aiding, abetting, or attempting to commit an offence; |  |  |  |  |
| 1.1e | the subject of a conviction for money laundering or terrorist financing; |  |  |  |  |
| 1.1f | the subject of a conviction of child labour and other forms of trafficking in human beings; |  |  |  |  |
| 1.2 | Non-payment of taxes or social security obligations Has it been established by a judicial or administrative decision having final and binding effect in accordance with Irish law or the legal provisions of the country in which the Economic Operator is established (if outside Ireland), that the Economic Operator is in breach of obligations related to the payment of tax and social security contributions? If the response to 1.2 above is in the affirmative, please provide further information on the decision and the amounts involved. |  |  |  |  |
| An Economic Operator who answers ‘Yes’ in any of the situations set out in paragraphs 2.1.a to 2.1.i will be excluded. |  |  |  |  |  |
| 2.1 | Please indicate if any of the following situations have applied, within the past three (3) years, or currently apply to your organisation. The Economic Operator: |  | Yes | No |  |
|  |  |  | Please indicate your answer by marking ‘X’ in the relevant box |  |  |
| 2.1a | has, in the performance of any public contract, failed to comply with applicable obligations in the field of environmental, social and labour law applying at the place where the works were carried out or the services provided, as established by EU law, national law, collective agreements or by international, environmental, social and labour law listed in Annex X of Directive 2014/24/EU; |  |  |  |  |
| 2.1b | is bankrupt or the subject of insolvency or winding-up proceedings, its assets are being administered by a liquidator or by the court, or has entered into an arrangement with creditors, suspended its business activities or is in any analogous situation arising from a similar procedure under national laws and regulations; |  |  |  |  |
| 2.1c | is guilty of grave professional misconduct which renders its integrity questionable; |  |  |  |  |
| 2.1d | has entered into agreements with other Economic Operators aimed at distorting competition; |  |  |  |  |
| 2.1e | has a conflict of interest within the meaning of Article 24 of 2014/24/EU that cannot be effectively remedied by other, less intrusive, measures; |  |  |  |  |
| 2.1f | confirms that it has had prior involvement in the preparation of the procurement procedure which has resulted in a distortion of competition, as referred to in Article 41 of 2014/24/EU, that cannot be remedied by other, less intrusive, measures; |  |  |  |  |
| 2.1g | has shown significant or persistent deficiencies in the performance of a substantive requirement under a prior public contract, a prior contract with a contracting entity, or a prior concession contract, which led to early termination of that prior contract, damages, or other comparable sanctions. |  |  |  |  |
| 2.1h | - is guilty of serious misrepresentation in supplying the information required for the verification of the absence of grounds for exclusion or the fulfilment of the selection criteria; or - has withheld such information or is not able to submit supporting documents required under Article 59 of Directive 2014/24/EU; or |  |  |  |  |
| 2.1i | has undertaken to: - unduly influence the decision-making process of the contracting entity, or - obtain confidential information that may confer upon the tenderer undue advantages in the procurement procedure; or - negligently provide misleading information that may have a material influence on decisions concerning exclusion, selection, or award. |  |  |  |  |
| Declaration re Statutory Obligations |  |  |  |  |  |
| We confirm that we are fully compliant with the following legislation, or equivalent legislation in our country of establishment / operation: |  |  | Yes |  | No |
| (i) | Employment Equality Acts 1998-2011 |  |  |  |  |
| (ii) | Equal Status Acts 2000-2011 |  |  |  |  |
| (iii) | National Minimum Wage Act 2000 as amended |  |  |  |  |
| (iv) | Organisation of Working Time Act 1997 as amended |  |  |  |  |
| (v) | Safety, Health and Welfare at Work Act 2005 and Safety, Health, and Welfare at Work (General Application) Regulations 2007 |  |  |  |  |
| (vi) | Disability Act 2005 |  |  |  |  |
| (vii) | General Data Protection Regulation 2016/679 |  |  |  |  |
| (viii) | We have procedures in place to ensure that our subcontractors, if any are used for this contract, apply the same standards. |  |  |  |  |
| Article 5K Declaration EU Regulation 2022/576 on restrictive measures in the Context of Russian Actions in the Ukraine |  |  |  |  |  |
| A. | In the light of Russian actions in the Ukraine, the European Council adopted a new regulation – EU Regulation 2022/576, aimed at restricting participation in economic activity by economic operators from the Russian Federation. Article 5K of the Regulation prohibits the award or continuation of contracts falling within the scope of the Public Procurement Directives to persons or undertakings related to the Russian Federation. To comply with EU Regulation 2022/576, the Contracting Authority is seeking a declaration from economic operators to this effect. |  |  |  |  |
| B. | I declare that none of the exclusions specified in EU Regulation 2022/576 apply to any party associated with this tender. |  | Yes |  | No |
| This Declaration is made for the benefit of the Contracting Authority I certify that the information provided in this declaration is accurate and complete to the best of my knowledge and belief. I understand that the provision of inaccurate or misleading information in this Declaration will lead to my organisation being excluded from participation in this and future tenders and I am signing on behalf of: |  |  |  |  |  |
| Name of Economic Operator: |  |  |  |  |  |
| Authorised Signatory: |  |  |  |  |  |
| Name in Block Capitals: |  |  |  |  |  |
| Position: |  |  |  |  |  |

## QUALITY ASSURANCE MANAGEMENT SYSTEM

| Weighting: Pass/Fail Minimum requirement to remain eligible in the competition: Tenderers should provide information on the measures in place to ensure the delivery of a quality service, for example by way of an externally certified system, or an equivalent in-house quality control process or system. Note: where multiple quality systems are operated, please repeat relevant tables regarding certification information. |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Do you operate a Quality Management System? |  | Yes |  |  |  | No |  |  |
| Manager responsible for Quality |  |  |  |  |  |  |  |  |
| Quality management |  | Externally Certified |  |  |  | In-house |  |  |
| If externally certified, please provide the following information: |  |  |  |  |  |  |  |  |
| Name of Certification Body |  |  |  |  |  |  |  |  |
| Date of most recent certification |  |  |  |  |  |  |  |  |
| Scope of Certification: |  |  |  |  |  |  |  |  |
| I confirm that evidence of compliance will be provided promptly on request. | Yes |  |  |  | No |  |  |  |

## HEALTH AND SAFETY MANAGEMENT SYSTEM

| Weighting: Pass/Fail Minimum requirement to remain eligible in the competition: Tenderers must demonstrate compliance with all the relevant Health & Safety legislation. The Contracting Authority reserves the right to seek further information to verify compliance. |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Do you comply with the Safety Health & Welfare at Work Act 2005? |  | Yes |  |  |  | No |  |  |
| Manager responsible for Health & Safety: |  |  |  |  |  |  |  |  |
| Health & Safety management |  | Externally Certified |  |  |  | In-house |  |  |
| If externally certified, please provide the following information: |  |  |  |  |  |  |  |  |
| Name of Certification Body: |  |  |  |  |  |  |  |  |
| Date of most recent certification: |  |  |  |  |  |  |  |  |
| Scope of Certification: |  |  |  |  |  |  |  |  |
| I confirm that evidence of compliance will be provided promptly on request. | Yes |  |  |  | No |  |  |  |

| Declaration re Health & Safety |  |
| --- | --- |
| We confirm that each member of staff is aware of his/her responsibilities under the Safety, Health and Welfare at Work Act 2005 and the Safety Health & Welfare at Work (Construction) Regulations 2006. In particular we are familiar with the general duties of Contractors as outlined in Part 3 Regulations 24-29 of the Safety Health & Welfare at Work (Construction) Regulations 2006; also, with the specific duties enumerated in Parts 4-14, Regulations 30-105 of those Regulations and Schedules 1-6 of those Regulations. We are aware of and will take into account the general principles of prevention as enumerated below when carrying out construction work associated with the project. |  |
| GENERAL PRINCIPLES OF PREVENTION APPLICABLE TO CONTRACTOR |  |
| The purpose of the General Principles of Prevention is to provide a framework within which account is taken when identifying hazards in the risk assessment required under section 19 of the Safety, Health, and Welfare at Work Act 2005. (i) The avoidance of risks. (ii) The evaluation of unavoidable risks. (iii) The combating of risks at source. (iv) The adaptation of work to the individual, especially as regards the design of places of work, the choice of work equipment and systems of work, with a view to alleviating monotonous work and work at a predetermined rate and to reduce their effect on health. (v) The adaptation of the workplace to technical progress. (vi) The replacement of dangerous articles, substances, or systems of work by non-dangerous articles, substances, or systems of work. (vii) The giving to collective protective measures priority over individual protective measure. (viii) The development of an adequate prevention policy in relation to safety, health, and welfare at work, which takes account of technology, organization of work, working conditions, social factors and the influence of factors related to the working environment. (ix) The provision of appropriate training and instruction to employees. |  |
| OBLIGATIONS UNDER SECTION 17(3) |  |
| We are aware as Contractors of our obligations under Section 17 (3) of the Safety Health & Welfare at Work Act 2005 to ensure so far as is reasonably practicable that the project ‘is constructed to be safe and without risk to health and that it complies in all respects, as appropriate, with the relevant statutory provisions. |  |
| COMPLIANCE BY STAFF |  |
| We confirm that all staff have received, read and will apply the Safety, Health and Welfare at Work (General Application) Regulations 2007 and the Guide to the Safety, Health, and Welfare at Work (General Application) Regulations 2007. Where the role required is Contractor risk assessments will be carried out and maintained on the job file. We confirm that in our opinion this declaration is deemed to satisfy our obligations in relation to the following areas given the scope and nature of the proposed works: (i) Health and Safety Policy and Organization; (ii) Arrangements. (iii) Competent Advice. (iv) Training and Information. (v) Individual Qualifications and Experience. (vi) Monitoring, Audit and Review. (vii) Workforce Involvement. (viii) Accident/Incident Reporting, Review. (ix) Sub-consulting Procedures. (x) Hazard Management and Risk Assessment. (xi) Health and Welfare. |  |
| ENFORCEMENT ACTIONS, LEGAL PROCEEDINGS, ACCIDENTS, FATALITIES, OR INCIDENTS |  |
| In relation to enforcement actions, legal proceedings, accidents, fatalities, or incidents associated with the discharge of our duties as Contractor or PSCS (whether the relevant discipline is the subject of this declaration or not) over the last three years: (Please attach any details or, alternatively, confirm that no enforcement actions, legal proceedings, accidents, fatalities, or incidents have arisen). |  |
| COMPETENT AND RESOURCED |  |
| We confirm that in our opinion our organization is competent and adequately resourced to fulfil its obligations under the Safety, Health, and Welfare at Work Act 2005 and that our organization has adequate resources to fulfil the role of Contractor as stated above. |  |
| Name of Economic Operator |  |
| Authorised Signatory |  |
| Signature |  |

## ENVIRONMENTAL MANAGEMENT SYSTEMS

| Weighting: Pass/Fail Minimum requirement to remain eligible in the competition: Tenderers must demonstrate compliance with all the relevant Environmental legislation and guidelines. The Contracting Authority reserves the right to seek further information to verify compliance. |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Do you comply with all relevant Environmental legislation? |  | Yes |  |  |  | No |  |  |
| Manager responsible for Environmental Systems: |  |  |  |  |  |  |  |  |
| Environmental management |  | Externally Certified |  |  |  | In-house |  |  |
| If externally certified, please provide the following information: |  |  |  |  |  |  |  |  |
| Name of Certification Body: |  |  |  |  |  |  |  |  |
| Date of most recent certification: |  |  |  |  |  |  |  |  |
| Scope of Certification: |  |  |  |  |  |  |  |  |
| I confirm that evidence of compliance will be provided promptly on request. | Yes |  |  |  | No |  |  |  |

# RESPONSE TO THE AWARD CRITERIA

## COST CRITERION - FORM OF TENDER

THIS FORM OF TENDER MUST BE COMPLETED AND RETURNED BY ALL TENDERERS.

Failure to sign this Form of Tender will invalidate the offer.

| To: |  |
| --- | --- |
| From: |  |
| Re: |  |

I/We have examined the tender documentation and hereby offer to provide the services in accordance with the details contained within the Invitation to Tender Document and the attached Detailed Breakdown of Charges and Expenses.

Environmental Chamber

| Requirement | Unit Price (€) Excluding VAT | Quantity | Total Cost (€) Excluding VAT |
| --- | --- | --- | --- |
| Equipment (please list all individual pieces of equipment separately) |  |  |  |
| Sub-Total A |  |  |  |
| Delivery (DDP) |  |  |  |
| Installation |  |  |  |
| Training |  |  |  |
| Warranty |  |  |  |
| Other Costs (Please Specify) |  |  |  |
| Sub-Total B |  |  |  |
| Discounts: - Discounts off list price - Other benefits (extended warranty, free issues, upgrades etc.) |  |  |  |
| Total Cost for Evaluation (A + B – Discounts) |  |  |  |

Additional Costs (Optional - not included in evaluation of cost)

| Requirement | Unit Price (€) Excluding VAT | Quantity | Total Cost (€) Excluding VAT |
| --- | --- | --- | --- |
| Warranty (Year 2) |  |  |  |
| Warranty (Year 3) |  |  |  |
| Warranty (Year 4) |  |  |  |
| Warranty (Year 5) |  |  |  |
| Service/Maintenance Contract (Year 2) |  |  |  |
| Service/Maintenance Contract (Year 3) |  |  |  |
| Service/Maintenance Contract (Year 4) |  |  |  |
| Service/Maintenance Contract (Year 5) |  |  |  |

I/We confirm that I/we:

- Have priced our tender on the basis of all rates being inclusive of all out of pocket expenses (i.e., mileage, subsistence, phone, postage, etc.) and account / contract management related costs,

- Will keep this offer for the contract / framework open for acceptance by you for a period of 12 months from the date of deadline for submission of tenders,

- Agree that you are not bound to accept the most economically advantageous or any tender you may receive,

- Agree that the rates stated are maximum prices for the duration of the contract,

- Have read and thoroughly examined the ITT,

- Fully understand the ITT and the Client’s requirements,

- Undertake to treat the details of the ITT, its tender, and any subsequent agreements as private and confidential,

- Acknowledge that acceptance by the Contracting Authority of this tender will not constitute a binding and enforceable agreement and that a legally enforceable agreement will not exist until and unless the contract is awarded between the Contracting Authority and the tenderer,

- Have availed of all offers for additional information or have otherwise satisfied myself/ourselves as to conditions that may in any manner affect the performance of the contract,

- Have included all elements necessary for the performance of the specified requirements, which are either expressly stated in the RFT or contained in any supplementary information, or which could reasonably be inferred therefrom,

- Have found no errors, omissions, conflicts, or ambiguities in the RFT except those which I/We have brought to the attention of the Contracting Authority before the latest date for submitting queries,

- Have included for compliance with all statutory requirements applicable in Ireland and those applicable in any country where parts of the contract may be performed that are in force 7 days prior to the deadline for receipt of tenders,

- Will not, if awarded a contract employ labour in a manner that is discriminatory in relation to gender, civil/family status, sexual orientation, religion, age, disability, race, or membership of the traveller community.

- Agree that as a condition of award, it shall be our sole responsibility to fulfil the obligations under the Contract, notwithstanding any changes in circulars, laws, regulations, taxation, duties, or other factors which might arise following the withdrawal of the United Kingdom from membership of the EU.

| Signed |  |
| --- | --- |
| Name (in Capital Letters) |  |
| On behalf of |  |
| Address |  |
| Telephone (Office) |  |
| Telephone (Mobile) |  |
| Email |  |
| Date |  |

## 3.2	RESPONSE TO MANDATORY REQUIREMENTS

| Ref | Requirement |
| --- | --- |
| M1 | Tenderer MUST ensure that the environmental chamber is supplied, installed, and fully commissioned by December 2026 in a pre-defined location within the Polaris Building. The successful Tenderer will be required to work collaboratively with the Contracting Authority and any other appointed contractors throughout the planning, installation, and commissioning phases. Tenderer MUST ascertain the existing site layout and levels, M&E connection points, facilities of access thereto, the supply of labour and materials the district will afford and the general convenience of working. They shall be taken to have acquainted themselves with the nature of the site and all other things insofar as they may have any connection with or affect the Works. Tenderer MUST consider the convenience of all persons who may be affected by the Works, including the methods to be used, appropriate delivery times, and any other relevant factors impacting site users or the surrounding environment. Tenderer MUST take these requirements into account when preparing their Tender. No additional claims for labour, materials, or other costs arising from a failure to do so will be accepted. Tenderer MUST confirm that, if awarded the contract, a site visit will be undertaken following contract award and prior to commencement of installation. |

Tenderers, please use this free text area for your response…

## 3.3	RESPONSE TO QUALITATIVE AWARD CRITERIA

| Criterion B | Weighting | Maximum Marks | Minimum Marks Required (50%) |
| --- | --- | --- | --- |
|  | 40% | 400 | 200 |
| Title | Fitness for Purpose of the Proposed Solution |  |  |
| Description | Technical and functional merit of the proposed solution to meet the specification requirements as described in sections 4.1.1, 4.1.2, and 4.1.3 of the ITT. Tenderers are required to provide responses to each of the mandatory, core, and optional requirements listed in sections 4.1.1, 4.1.2, and 4.1.3 of the ITT. Tenderers must submit a detailed description in response to each of the requirements M1, T1 to T22 inclusive and O1 to O6 inclusive which demonstrates in a clear and comprehensive manner how they will meet the requirement. The provision of affirmative statements by the Tenderer that it can or will meet a requirement, or a reiteration of the requirement will not be considered sufficient. Where no response or the response does not sufficiently address the subject matter of the requirement the tenderer may be eliminated from the competition. While Tenderers may upload supporting documents as separate documents with their response, they must be clearly identified (for example, naming the relevant document, referring to the specific Technical Requirement to which it applies) relevant to demonstrating that the Tenderer can meet the specific requirement, and noting the page limits outlined below. This criterion will be assessed on the totality of the response. |  |  |
|  | Tenderers’ responses under this Criterion are limited to a maximum of 2 A4 pages per Technical Requirement (that is 2 A4 pages for T1, 2 A4 pages for T2 and so on) including the use of additional supporting documents where required. |  |  |

| Ref | Requirement |
| --- | --- |
| T1 | Tenderer MUST supply, deliver, install and commission a high-performance, controlled environmental chamber specifically designed to support human and athlete performance testing, physiological research, and environmental conditioning during exercise. Precise and simultaneous control of temperature, humidity and simulated altitude is required, together with the ability to monitor carbon dioxide concentration and respond appropriately to maintain user safety. The environmental chamber and associated components MUST be installed before the year end (December 2026), in a pre-determined space as shown in Appendix A – Environmental Chamber Floor Plan. Please confirm and in detail the make, model and specifications of the proposed unit and all components. Note: Pricing MUST be clearly outlined in the Form of Tender (TRD Section 2.1 – Cost Criterion) |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| T2 | Tenderer MUST ensure that the environmental chamber makes optimal use of the available footprint and volume to maximise usable space. Tenderer are required to provide a detailed description of how the proposed equipment will fit within the available footprint and volume, as shown in “Appendix A – Environmental Chamber Floor Plan”. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| T3 | Tenderer MUST ensure that the environmental chamber and associated anteroom include doors which: - are accessible from Exercise Physiology Lab, FT112. - perform sufficiently to maintain the desired environmental conditions, as outlined in T4 – T6/O4 – O6 below. - are of a size, construction and location appropriate for the movement of typical exercise physiology equipment (e.g. cycle ergometers, metabolic analysers) between the environmental chamber and laboratory FT112. - can be opened from the outside and from the inside at any time, including in the case of failure of the environmental chamber control system or failure of the building power supply. - meet the requirements of local accessibility regulations. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| T4 | Tenderer MUST ensure that the environmental chamber includes windows. - which perform sufficiently to maintain the desired environmental conditions. - are of sufficient size to allow several observers to see into the chamber. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| T5 | Tenderer MUST ensure the environmental chamber perform to achieve the following temperature specifications: - Range: +10 to +50 ⁰C |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| T6 | Tenderer MUST ensure the environmental chamber perform to achieve the following simulated altitude conditions - Range: 0 to 3,500 m simulated altitude (~20.9% O2 to ~13.5% O2) |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| T7 | Tenderer MUST ensure environmental chamber perform to achieve the following relative humidity specifications - Range: 30 to 80% RH |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| T8 | Tenderer MUST ensure environmental chamber control system and associated components create an environment which does not allow airflow rates likely to induce windchill-like conditions and associated physiological responses in exercising users. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| T9 | Tenderer MUST ensure the environmental chamber control system and associated components ensure a stable and uniform environment (temperature, relative humidity, % oxygen, % carbon dioxide) across all areas of the room when occupied by the recommended number of exercising users. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| T10 | Tenderer MUST ensure the environmental chamber control system be able to detect the carbon dioxide level at multiple locations within the chamber and respond accordingly (e.g. alarms, automatic shutdown) when safety thresholds are exceeded. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| T11 | Tenderer MUST ensure the environmental chamber control system include easily accessible, clearly marked emergency stop buttons located both within and outside the chamber which will immediately halt operation of the chamber when activated. This user-operated shut-down function MUST be operational in the event of a building power outage. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| T12 | Tenderer MUST ensure the environmental chamber control system include audible and visible alarms automatically triggered by environmental conditions (oxygen, carbon dioxide, temperature) exceeding user-set safety thresholds. - The sensors MUST BE capable of sampling from more than one appropriate location within the environmental chamber. - The alarms MUST be audible and visible both within and outside the chamber. - The alarms MUST operate even in the event of a building power outage |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| T13 | Tenderer MUST ensure the environmental chamber control system include an automatic shut-down mechanism automatically triggered by environmental conditions (oxygen, carbon dioxide, temperature) exceeding user-set safety thresholds. - The sensors MUST sample from more than one appropriate location within the chamber. - This shut down function MUST operate even in the event of a power outage. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| T14 | Tenderer MUST ensure the environmental chamber and anteroom include lighting which will automatically function in the event of a building power outage to enable safe exit of chamber users. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| T15 | Tenderer MUST ensure the environmental chamber have fire resistant materials, smoke detection sensors, and integration with the building fire safety systems. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| T16 | Tenderer MUST ensure the environmental chamber have a level, non-slip floor with load bearing capacity appropriate for the intended use (treadmill etc.). |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| T17 | Tenderer MUST ensure the environmental chamber ventilation system include air filtration mechanisms to meet recommended safety standards in compliance with the Safety, Health and Welfare at Work Regulations, and HSA Code of Practice for Indoor Air Quality. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| T18 | The environmental chamber control system MUST provide an intuitive and user-friendly interface suitable for use by a range of users across educational, research, and sport performance environments. Tenderers MUST provide details of the control system interface, including ease of operation, user controls, and available features. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| T19 | Tenderer MUST ensure the environmental chamber control system allow users to schedule the system to automatically start and commence the ramping process at a user-defined time. Tenderers MUST provide details of the scheduling functionality available. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| T20 | Tenderer MUST ensure the user-accessible control system for the environmental chamber be located outside the chamber (in room FT112 shown on “Appendix A - Environmental Chamber Floor Plan”). |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| T21 | Tenderer MUST ensure the environmental chamber control system be designed so that access to adjust chamber settings is restricted (e.g. to password holders), but that this access restriction can be over-ridden to ensure that the chamber operation can be shut off, if necessary, in an emergency. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| T22 | Tenderer MUST ensure the environmental chamber provide high-frequency data logging of chamber conditions, including temperature, relative humidity, simulated altitude, oxygen concentration, and carbon dioxide concentration. Logged data MUST be accessible and exportable in a commonly used file format compatible with data management platforms (e.g., CSV format). |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| O1 | Tenderer SHOULD make sure the environmental chamber wall includes a small ‘hatch’ of up to 40cm x 40cm in size - for the transfer of small items (e.g. blood samples, documents) out of the chamber without opening the main door - of a construction suitable to maintain environmental conditions within the chamber the inclusion of a hatch would be considered highly desirable |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| O2 | Tenderers SHOULD provide cable pass-throughs within the control room/environmental chamber wall, as indicated in the reference image below. The final number, size, and location of cable pass-throughs will be agreed between the Tenderer and Contracting Authority during the design phase |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| O3 | Tenderer SHOULD ensure the internal and external surfaces of the environmental chamber and anteroom walls be of a structure and material suitable for mounting relevant equipment (e.g. display screens). |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| O4 | The environmental chamber should be capable of achieving the following temperature performance specifications: - Range: - +10 to +50⁰C - -20 to +50⁰C (highly desirable) - Accuracy: Tenderer to specify - Uniformity: Tenderer to specify - Ramp time: Tenderer to specify Note: Where Tenderers are in a position to meet and exceed the minimum temperature performance specifications, DCU requests that Tenderers provide these options in their Response. Tenderer should provide details of each option separately. All associated costs for each option MUST be clearly identified and outlined in the Form of Tender (TRD Section 2.1 – Cost Criterion). |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| O5 | The environmental chamber should be capable of achieving the following simulated altitude specifications: - Range: - 0 to 3500m simulated altitude (~20.9% O2 to ~13.5% O2) [minimum] - 0 to 6000m (~13.5% O2) [highly desirable] - Accuracy: Tenderer to specify - Uniformity: Tenderer to specify - Ramp time: Tenderer to specify |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| O6 | It is desired that the environmental chamber is capable of achieving following humidity specifications: - Range: 20 to 90% RH would be considered highly desirable. - Accuracy: Tenderer to specify - Uniformity: Tenderer to specify - Ramp time: Tenderer to specify |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the response to this requirement is a maximum of two (2) A4 pages as set out above.

| Criterion C | Weighting | Maximum Marks | Minimum Marks Required (50%) |
| --- | --- | --- | --- |
|  | 10% | 100 | 50 |
| Title | Service and Support |  |  |
| Description | Technical and functional merits of the proposed solution to meet the Service and Support requirements as described in sections 4.1.4 of the ITT. Tenderers are required to provide responses to each of the mandatory Service and Support requirements listed in sections 4.1.4 of the ITT. Tenderers must submit a detailed description in response to each of the requirements P1 to P10 inclusive which demonstrates in a clear and comprehensive manner how they will meet the requirement. The provision of affirmative statements by the Tenderer that it can or will meet a requirement, or a reiteration of the requirement will not be considered sufficient. While Tenderers may upload supporting documents as separate documents with their response, they must be clearly identified (for example, naming the relevant document, referring to the specific requirement to which it applies) and relevant to demonstrating that the Tenderer can meet the specific requirement. This criterion will be assessed on the totality of the response. |  |  |

| Ref | Requirement |
| --- | --- |
| P1 | Tenderers MUST provide a minimum two (2) year warranty included within the quoted price. The warranty must cover defects arising under normal operating conditions, including callouts, labour, travel, replacement parts, control systems, sensors, electrical systems, safety locks, and other critical components. Tenderers MUST clearly define all warranty inclusions, exclusions, limitations, and conditions. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| P2 | TENDERERS MUST provide a clear scope of warranty coverage, including all major components such as control systems, sensors, electrical systems, and safety mechanisms. All exclusions, limitations, and conditions MUST be clearly defined, including the distinction between normal operation, wear and tear, misuse, and improper maintenance. Tenderers MUST also outline warranty duration, support arrangements, and repair/replacement processes. Note: Tenderers MUST provide the associated costs for any extended warranty options in the TRD Section 2.1 – Cost Criterion. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| P3 | Tenderer MUST ensure that the warranty includes remote diagnostics within 48 hours of notification and, where required, on-site intervention, subject to site access and location constraints. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| P4 | Tenderer MUST ensure that all replacement parts supplied under warranty are new or manufacturer-approved equivalents, with appropriate traceability and certification where applicable. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| P5 | Tenderer MUST ensure that warranty support includes all software updates and firmware patches necessary to maintain system stability, performance, and security throughout the warranty period. Tenderers MUST confirm and provide full details of the warranty offering, including associated costs, in Form of Tender (TRD Section 2.1 – Cost Criterion). |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| P6 | Tenderers MUST provide an option to extend the warranty period through a Service Level Agreement (SLA), including preventative maintenance packages where applicable. Note: Tenderers MUST provide details and associated pricing for warranty extensions and/or maintenance packages for Years 3 and 4 in the Form of Tender (TRD Section 2.1 – Cost Criterion). |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| P7 | Tenderer MUST include comprehensive on-site training for five (5) staff for one (1) day following installation. The training MUST cover the full scope of operation, including extended use at end-range environmental conditions, safe operation, programming, emergency procedures, and troubleshooting. The environmental chamber specialist contractor MUST explain and demonstrate the purpose, function, and operation of the installation. Training dates and timings will be agreed between DCU and the successful Tenderer. Please confirm and detail, outlining costs in the Form of Tender (TRD Section 2.1 – Cost Criterion). |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| P8 | Tenderer MUST include a training package comprising dedicated training on routine maintenance procedures, calibration checks and verification, fault diagnostics and system resets, replacement of basic components, and other basic troubleshooting. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| P9 | Tenderer MUST ensure that all manuals, quick-reference guides, and digital resources, including any templates or systems for tracking maintenance and performance over time, are provided in English. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| P10 | Tenderer MUST include access to specialised support personnel with defined support hours and clear escalation pathways, including availability of trained personnel for on-site support, diagnostics, and repair. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Criterion D | Weighting | Maximum Marks | Minimum Marks Required (50%) |
| --- | --- | --- | --- |
|  | 15% | 150 | 75 |
| Title | Delivery, Installation and Commissioning |  |  |
| Description | Technical and functional merit of the proposed solution to meet the specification requirements as described in sections of the 4.1.5 of the ITT. Tenderers are required to provide responses to each of the Delivery, Installation and Commissioning Requirements listed in sections 4.1.5 of the ITT. Tenderers must submit a detailed description in response to each of the requirements D1 to D10 inclusive which demonstrates in a clear and comprehensive manner how they will meet the requirement. The provision of affirmative statements by the Tenderer that it can or will meet a requirement, or a reiteration of the requirement will not be considered sufficient. Where no response or the response does not sufficiently address the subject matter of the requirement the tenderer may be eliminated from the competition. While Tenderers may upload supporting documents as separate documents with their response, they must be clearly identified (for example, naming the relevant document, referring to the specific requirement to which it applies) and relevant to demonstrating that the Tenderer can meet the specific requirement. This criterion will be assessed on the totality of the response. |  |  |

| Ref | Requirement |
| --- | --- |
| D1 | Tenderer MUST provide a detailed lead time and delivery schedule from contract award through to final installation, commissioning, training and handover. - Delivery must include all shipping charges, insurance, customs duties and all aspects as outlined in DDP (Delivery Duty Paid) incoterms. - The environmental chamber and all its parts must be delivered to Dublin City University, Polaris Building, DCU Glasnevin, D09 YH9P, Ireland. - Installation MUST be complete by December 2026. Please confirm cost (if any) in Form of Tender (TRD Section 2.1 – Cost Criterion). |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| D2 | Tenderer MUST ensure that the environmental chamber is supplied, installed, and fully commissioned by December 2026. Tenderers are required to state the proposed lead time for delivery. Note: Shorter lead times are desirable. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| D3 | Tenderer MUST clearly identify all key milestones and logistics within the proposed delivery schedule of the environmental. Information provided MUST include: - Schedule for regular progress updates throughout the project. - Plan for identifying delivery risked with mitigation strategies included. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| D4 | Tenderer MUST provide all facility requirements (if any) in advance of the delivery of the environmental chamber equipment. This MUST include, as a minimum: - number of containers and weight of each - any clearance requirements (if necessary). - requirements for Health & Safety - requirements for site access |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| D5 | Tenderer MUST provide details of your contingency plans (alternative sources/routes of supply) for any delays in delivery. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref |  | Requirement |
| --- | --- | --- |
| D6 | Tenderers MUST install, commission and calibrate the equipment. Please outline your installation plan in terms of how you would aim to deliver this project and work within the boundaries of a live construction site. |  |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| D7 | Tenderer MUST provide a detailed handover plan and clearly define the handover acceptance criteria. This MUST include, as a minimum: - Completion and submission of all required documentation (e.g. user manuals, technical documentation, maintenance guides, warranties, and as-built documentation, where applicable). - Completion of all agreed training for designated University personnel. - Successful commissioning, testing, and verification of the system/equipment, where applicable. - Resolution of any outstanding defects or snag items identified during installation and testing. - Formal operational acceptance and sign-off by the University confirming that the system/equipment has been delivered, installed, tested, and is ready for operational use. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| D8 | Tenderer MUST outline the measures they will implement to prevent damage and describe how any damage caused during the works will be repaired or reinstated at the Tenderer's own cost, restoring the affected area to its original condition or better, to the satisfaction of the University. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| D9 | Tenderer MUST comply with all the legal acts and regulations which apply to this construction site. Particular attention is drawn to the requirements of: - Safety, Health and Welfare at Work Acts, 1989 and 2005 - Factories Act, 1955 - Safety in Industry Act, 1980 - Conditions of Employment Act - Safety, Health and Welfare (General Applications) Regulations, 1993 - Safety, Health and Welfare at Work (Construction) Regulations, 1995 - Safety, Health and Welfare at Work (Construction) Regulations, 2001 - Safety, Health and Welfare at Work Regulations, 2003 - Safety, Health and Welfare of Work (Construction) Regulations, 2006 - Safety, Health and Welfare of Work (Construction) Regulations, 2013 - Safety, Health and Welfare of Work (Construction) Regulations, 2019 - Safety, Health and Welfare of Work (Construction) Regulations, 2020 The environmental chamber specialist contractor shall provide for complying with these Acts, and relevant Regulations including Notifications and placards for display. Tenderer MUST confirm compliance and outline how they would deliver this project adhering to all safety, health and welfare guidelines. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| D10 | Tenderer is advised that personal protective equipment (PPE) MUST be worn at all times while on site. The minimum requirements for PPE on this site are - Safety Helmet - Safety type footwear (waterproof, if appropriate) - High visibility vest or jacket (MUST only bear the relevant Contractor’s name) - Gloves - Eye protection where necessary Additional PPE MUST be worn whilst carrying out activities such as grinding, welding, etc. Tenderer MUST confirm compliance and outline how they would deliver this project while adhering to these guidelines. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| D11 | The Tenderer MUST include a preventative maintenance schedule for the technical staff, including recommended calibrations, service intervals, inspection checklists and consumable replacement guidance. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| D12 | The Tenderer MUST provide details on proper classification, handling and disposal of any hazardous materials e.g. refrigerants, oils, electronic components, with Safety Data Sheets (SDS) provided |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| D13 | The Tenderer MUST provide detailed instructions for safe decommissioning, including identification of components requiring specialist handling. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| D14 | Tenderer MUST clearly set out the allocation of responsibility for decommissioning and disposal, including any proposed supplier take-back arrangements or recycling schemes. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| D15 | Tenderer MUST ensure that, on completion of the works the site is cleared of all plant, equipment, tools, unused materials, wiring, metal, broken blocks, stones, debris, and all other loose objects or waste arising from the works. All waste materials MUST be removed from site and disposed of at a licensed waste disposal facility at the Tenderer’s own cost. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Criterion E | Weighting | Maximum Marks | Minimum Marks Required (50%) |
| --- | --- | --- | --- |
|  | 5% | 50 | 25 |
| Title | Sustainability |  |  |
| Description | Technical and functional merit of the proposed solution to meet the specification requirements as described in section 3.2 of the ITT. Tenderers are required to provide responses to each of the Sustainability Requirements listed in section 3.2 of the ITT. Tenderers must submit a detailed description in response to each of the requirements S1 to S4 inclusive which demonstrates in a clear and comprehensive manner how they will meet the requirement. The provision of affirmative statements by the Tenderer that it can or will meet a requirement, or a reiteration of the requirement will not be considered sufficient. Where no response or the response does not sufficiently address the subject matter of the requirement the tenderer may be eliminated from the competition. While Tenderers may upload supporting documents as separate documents with their response, they must be clearly identified (for example, naming the relevant document, referring to the specific requirement to which it applies) and relevant to demonstrating that the Tenderer can meet the specific requirement. This criterion will be assessed on the totality of the response. |  |  |

| Ref | Requirement |
| --- | --- |
| S1 | Tenderers MUST provide details of any recyclable or low-impact materials used in the construction and composition of the environmental chamber. Tenderers should identify the materials used, their recyclability, and any environmental benefits associated with their use |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| S2 | Tenderers MUST provide details of the measures implemented to minimise waste during the manufacture, installation, and operation of the environmental chamber. Tenderers should outline any waste reduction practices, reuse or recycling initiatives, and processes adopted to minimise environmental impact throughout the lifecycle of the system. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| S3 | Tenderer MUST ensure that the chamber and associated components are designed to minimise energy consumption, thereby reducing operational costs and environmental impact during use. The Tenderer MUST outline the expected rate of energy consumption while establishing and maintaining a variety of typical environments (i.e., likely combinations of temperature, humidity, and simulated altitude conditions). |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.

| Ref | Requirement |
| --- | --- |
| S4 | The chamber and associated components SHOULD be designed to ensure minimum noise transmission to nearby rooms and areas while operating through acoustic insulation, low-noise fans and compressors, vibration isolation mounts, etc. Typical operational noise levels SHOULD comply with relevant standards and project requirements. |

Tenderers, please use this free text area for your response (that is, do not respond in the box above) and note the overall response to this Criterion is a maximum of five (5) A4 pages as set out above.