## 1 - Technical Merit
APPENDIX 1 - SPECIFICATION OF REQUIREMENTS
THE SUPPLY, SUPPORT AND MAINTENANCE OF FLUID WARMERS
SECTION 1 - CLINICAL SUITABILITY AND TECHNICAL MERIT OF PROPOSED EQUIPMENT AND CONSUMABLES
REF | REQUIREMENT | COMPLIANT (Yes / No) | DETAILED NARRATIVE COMMENT
1.1 | MANDATORY: The proposed equipment must be CE-marked. Tenderers must confirm that this is the case and must provide all relevant information. | Compliant (Yes / No) | Insert Comment
1.2 | MANDATORY: The proposed equipment and all components of the proposed equipment must meet the minimum requirements of IEC 60601-1, 60601-1-2 and 60601-2-41. Tenderers must confirm that this is the case and must provide all relevant information in this regard. Tenderers must also list all standards and regulations not listed above that apply to the tendered equipment and must provide evidence of compliance with each standard and regulation. | Compliant (Yes / No) | Insert Comment
1.3 | MANDATORY: The tenderer must prove compliance with EU MDR (EU) 2017/745. For legacy devices leveraging Regulation (EU) 2023/607 extensions (to 2027 or 2028), the tenderer must provide a valid MDD certificate, proof of a formal MDR application lodged with a Notified Body, and a signed transfer agreement. Additionally, the tenderer must certify that no significant design changes have occurred and demonstrate active adherence to MDR post-market surveillance, vigilance, and EUDAMED requirements. | Compliant (Yes / No) | Insert Comment
1.4 | MANDATORY: If the proposed devices/products have ever been the subject of a field safety notice or product recall Tenderers must provide all relevant information. | Compliant (Yes / No) | Insert Comment
1.5 | MANDATORY: Commissioning and electrical safety tests must be carried out on site on delivery of new devices by the company engineer.Please confirm that this is the case and state the electrical safety rating of the proposed device including IEC 62353 Class and Type. | Compliant (Yes / No) | Insert Comment
1.6 | MANDATORY: The device must be suitable for heating both blood and fluids as standard. The proposed devices must perform at high flow rates (>4 L/h) with temperature range for keeping blood and fluid warm between 37-42 degrees. The device must give audible/visual warning if it is outside of the tolerance range of 37-42 degrees. Please state flow rate (ml per minute). Tenderers must provide all relevant information. | Compliant (Yes / No) | Insert Comment
1.7 | MANDATORY: The device must be suitable for the delivery of blood and intravenous fluid at normothermic temperature under gravity flow conditions. | Compliant (Yes / No) | Insert Comment
1.8 | MANDATORY: The proposed device must have a clear, concise Graphical User Interface (GUI) which allows for ease of use. Tenderers must provide all relevant information and where applicable images of the GUI. | Compliant (Yes / No) | Insert Comment
1.9 | MANDATORY: Tenderers must state if the unit requires any disposable components and consumables that are brand-specific. Tenderers must provide all relevant information in this regard inclusive of a full list of components and consumables required. | Compliant (Yes / No) | Insert Comment
1.10 | MANDATORY: Tenderers must state if the unit can operate with commonly available third party disposable components and consumables. | Compliant (Yes / No) | Insert Comment
1.11 | MANDATORY: Tenderers must list all manufacturers of third party disposable components and consumables that are compatible with the monitor. | Compliant (Yes / No) | Insert Comment
1.12 | MANDATORY: Tenderers must state the weight and dimensions of the equipment. | Compliant (Yes / No) | Insert Comment
1.13 | MANDATORY: The administration set must be Latex free and free of Toxins. | Compliant (Yes / No) | Insert Comment
1.14 | DESIRABLE: Please state the compliance with needle-free and needleless systems. | Compliant (Yes / No) | Insert Comment
1.15 | MANDATORY: Tenderers must confirm that the administration set has a straight fluid pathway with no obstructions to allow for high flow rates and reduce haemolysis. The administration set must have a clamp on the line to prevent free flow of fluid or blood. | Compliant (Yes / No) | Insert Comment
1.16 | MANDATORY: Tenderers must provide details of cleaning instructions, including approved cleaning agents, for the proposed unit. | Compliant (Yes / No) | Insert Comment
1.17 | DESIRABLE: Please state the temperature of fluids on reaching IV access to patient without an extension set. | Compliant (Yes / No) | Insert Comment
1.18 | MANDATORY: Tenderers must state the time taken for the device to be ready for use from power up, once the adminstration set is in place. | Compliant (Yes / No) | Insert Comment
1.19 | MANDATORY: Tenderers must provide Infection Prevention and Control Data for the administration set. | Compliant (Yes / No) | Insert Comment
1.20 | DESIRABLE: Please state if the administration set has the Luer Lock compatibility option and whether the administration set is with protective cap and state the administration set's compatibility with MRI | Compliant (Yes / No) | Insert Comment
1.21 | MANDATORY: The administration set must be single use. State the durability of a set in hours once it is opened for patient use and the priming volume for all set types. | Compliant (Yes / No) | Insert Comment

## 2 - Support_Maintenance
SECTION 2- Proposed Support and Maintenance Methodology
REF | REQUIREMENT | COMPLIANT (Yes / No) | DETAILED NARRATIVE COMMENT
2.1 | MANDATORY: For the avoidance of doubt, the Contracting Authority’s requirements regarding support and maintenance are clearly outlined hereunder. Tenderers are advised that support and maintenance proposals must respond accurately to all subsequent points. In this regard, it may be necessary for tenderers to amend their standard service and support offerings (e.g., any pre-defined service packages) for the purposes of participating in this tender competition. Tenderers must explicitly confirm their understanding of this fact. | Compliant (Yes / No) | Insert Comment
2.2 | MANDATORY: While the successful tenderer may wish to implement a service and support offering that corresponds to a pre-defined service package (e.g., a ‘gold,’ ‘silver’ or ‘bronze’ service package), it is emphasised that such classification will be permitted for the sake of the successful tenderer’s convenience only, and that the support and maintenance specification outlined hereunder shall take precedence over any such classification. Any terms in any pre-defined service package that seek to vary, amend, qualify or mitigate the provisions of this specification of requirements are hereby expressly rejected. Tenderers must explicitly confirm their understanding of this fact. | Compliant (Yes / No) | Insert Comment
2.3 | MANDATORY: This support and maintenance specification shall form the basis of the service contract ultimately entered into with the successful tenderer. On this basis, any standard support, service or maintenance terms imposed by the successful tenderer shall be considered secondary to the Contracting Authority’s service contract. In the event of any discrepancy or contradiction between this specification, as codified by the Contracting Authority’s service contract, and any support, service or maintenance terms that the successful tenderer seeks to impose, this specification, as codified by the Contracting Authority’s service contract, shall take precedence. Tenderers must explicitly confirm their understanding of this fact. | Compliant (Yes / No) | Insert Comment
2.4 | MANDATORY: Tenderers’ attention is drawn to the fact that service contract prices are fixed in the associated Pricing Schedule (see Appendix 3). For the avoidance of doubt, these fixed prices shall not be subject to any increases or variations arising out of the consumer price index, any force majeure events, or any amendments to the successful tenderer’s operating costs. Having been afforded the opportunity to tender a separate price for each year of support and maintenance potentially required by the Contracting Authority, any subsequent price increase requested by the successful tenderer shall be automatically rejected. Tenderers must explicitly confirm their understanding of this fact. | Compliant (Yes / No) | Insert Comment
2.5 | MANDATORY: This support and maintenance specification applies equally to all aspects of the tendered solution. In the event that the successful tenderer is responsible for providing equipment from a range of manufacturers or third parties in order to meet the Contracting Authority’s requirements, the successful tenderer must assume full responsibility for co-ordinating the support and maintenance of the entire solution, including all third-party elements. In this regard, the Contracting Authority will not tolerate engaging with multiple separate companies or points of contact in order to ensure the resolution of a fault, unless such an approach would result in the resolution of an issue within a shorter timeframe. Tenderers must confirm their understanding of this fact and must outline all relevant procedures to be followed in this regard. | Compliant (Yes / No) | Insert Comment
2.6 | MANDATORY: The life expectancy of all equipment provided by the successful tenderer must be at least ten years from the date of commissioning. In order to demonstrate compliance with this requirement, tenderers are required to provide documented evidence demonstrating the legitimate life expectancy of the equipment. | Compliant (Yes / No) | Insert Comment
2.7 | MANDATORY: The tendered equipment must remain supported by the original equipment manufacturer (to include the provision of all parts, all technical support and all clinical support potentially required) for a period of at least ten years from the date of commissioning. In order to demonstrate compliance with this requirement, tenderers must provide appropriate written assurance from the original equipment manufacturer attesting to this fact. Tenderers are advised that, as a binding contract will be entered into on foot of this specification, should any situation arise in which the equipment becomes unsupported before the expiry of this ten year period, the Contracting Authority shall avail of every course of action necessary to protect its position, up to and including legal action. | Compliant (Yes / No) | Insert Comment
2.8 | MANDATORY: In the event that a tenderer is not the manufacturer of the equipment in question, the tenderer in question must provide objectively verifiable documentary evidence to the effect that the tenderer or any sub-contractor responsible for discharging the support and maintenance obligations detailed in this specification is authorised and approved by the equipment manufacturer to undertake all relevant support and maintenance activities. This documentary evidence must be included with tender submissions. | Compliant (Yes / No) | Insert Comment
2.9 | MANDATORY: Tenderers must state the Risk Class of Medical Device (i.e. Class 1, Class 11a, Class 11b or Class 111) applicable to the tendered equipment, and must provide all relevant information in this regard. | Compliant (Yes / No) | Insert Comment
2.1 | MANDATORY: In the event that the successful tenderer is made aware of any end-of-life notice or any analogous information regarding the equipment in question at any point during the warranty period or subsequent service contract, such information must be immediately forwarded to the Contracting Authority, with all necessary supporting documentation appended. Tenderers must confirm their understanding of this fact. | Compliant (Yes / No) | Insert Comment
2.11 | MANDATORY: In the event that any field safety notices or any analogous documents are raised regarding the equipment in question at any point during the warranty period or subsequent service contract, such information must be immediately forwarded to the Contracting Authority, with all necessary supporting documentation appended. Tenderers must confirm their understanding of this fact. | Compliant (Yes / No) | Insert Comment
2.12 | MANDATORY: A comprehensive warranty of 12 months must be provided by the successful tenderer, commencing upon the date of commissioning in question. The warranty must include the conduct of planned preventative maintenance visits in accordance with manufacturers’ recommendations, and the replacement of all parts and components of the tendered solution, whether minor ‘wear and tear’ items, routine spare parts or major spare parts. The warranty must also involve the immediate resolution of all faults, breakdowns and errors of every description. In order to demonstrate compliance with this requirement, tenderers must include in their tender submissions their proposed warranty terms for consideration by the Contracting Authority, ensuring that these terms do not vary or otherwise amend the requirements of this support and maintenance specification. | Compliant (Yes / No) | Insert Comment
2.13 | MANDATORY: Tenderers are advised that the failure of any component of a tendered solution during the warranty period will give rise to serious fitness for purpose concerns on the part of the Contracting Authority. Consequently, time shall be of the essence regarding the resolution of any fault during the warranty period. Tenderers are required to confirm their understanding of this fact and to outline all mechanisms to be relied upon in immediately rectifying any issues arising during the warranty period. | Compliant (Yes / No) | Insert Comment
2.14 | MANDATORY: Immediately prior to the expiration of the warranty period referred to in Point 2.13 above, the successful tenderer shall perform a final comprehensive service of the equipment in question. Tenderers must explicitly confirm their understanding of this fact. | Compliant (Yes / No) | Insert Comment
2.15 | MANDATORY: Tenderers are advised that, should a component fail during the warranty period, the warranty period shall recommence from the date upon which the fault has been resolved, to the satisfaction of the Contracting Authority, by the successful tenderer. Tenderers must confirm their understanding of this fact. | Compliant (Yes / No) | Insert Comment
2.16 | MANDATORY: In the event of component failure or non-performance resulting in the replacement of a component, the Contracting Authority must have the ability to reject options, based on objectively verifiable evidence, and must be permitted to specify possible alternatives. Tenderers must confirm their understanding of this fact and must detail all mechanisms that will be implemented to afford the Contracting Authority this level of discretion. | Compliant (Yes / No) | Insert Comment
2.17 | MANDATORY: Promptly following the award of this contract, the successful tenderer must provide a comprehensive set of technical and maintenance manuals (including diagrams, parts lists, schematics, etc). Such technical and maintenance manuals must be provided in the English language, in hard and soft copies, and must include patient safety recommendations, safety test procedures, disassembly and reassembly procedures and troubleshooting recommendations / solutions. Tenderers must explicitly confirm their understanding of this fact. | Compliant (Yes / No) | Insert Comment
2.18 | MANDATORY: As outlined below, it is the Contracting Authority's intention to enter into a comprehensive contract, involving breakdown repair and the provision of parts during core working hours. However, tenderers must commit to making a planned preventative maintenance support offering available, costed on a pro-rata basis, should the need arise. Tenderers must confirm their understanding of this fact and must detail all mechanisms that will be implemented to afford the Contracting Authority this level of discretion. | Compliant (Yes / No) | Insert Comment
2.19 | MANDATORY: Both during the warranty period and thereafter, the support and maintenance offering must involve preventative maintenance visits inclusive of electrical safety testing at a frequency consistent with equipment manufacturer recommendations and a comprehensive maintenance service involving the provision of spare parts and breakdown repair during core working hours (i.e., Monday to Saturday, 08:00 to 20:00). | Compliant (Yes / No) | Insert Comment
2.2 | MANDATORY: Tenderers must confirm that, as a component of the support and maintenance offering to be provided by the successful tenderer, any required electrical safety tests for the tendered equipment shall be performed in accordance with IEC60601 and 62353 standards. Tenderers must provide all relevant information in this regard. | Compliant (Yes / No) | Insert Comment
2.21 | MANDATORY: The work undertaken during each service visit (whether a planned preventative maintenance visit or a breakdown repair visit) must correspond in full to any manufacturer recommendations. Tenderers must confirm their understanding of this fact and must outline any procedures and processes to be adhered to in this regard. | Compliant (Yes / No) | Insert Comment
2.22 | MANDATORY: The planned preventative maintenance visits must include electrical safety testing, equipment performance, safety, quality, functional tests, calibration and the replacement of normal wear and tear items. Tenderers must confirm their understanding of this fact, and must outline the proposed frequency of planned preventative maintenance visits for the equipment in question, confirming that the frequency corresponds to the manufacturer’s requirements. | Compliant (Yes / No) | Insert Comment
2.23 | MANDATORY: The successful tenderer will be required to utilise appropriate servicing and testing instruments for the purposes of carrying out all maintenance obligations under this contract. Tenderers must comprehensively describe all instrumentation to be employed in this regard. Additionally, tenderers must confirm that all such instrumentation is appropriately calibrated and, where relevant, certified. | Compliant (Yes / No) | Insert Comment
2.24 | MANDATORY: The timing and duration of planned preventative maintenance visits must not jeopardise uptime requirements. Tenderers must provide full information in this regard, outlining the proposed duration of each visit and the manner in which the visits will be scheduled to take into consideration the day-to-day utilisation of the equipment. | Compliant (Yes / No) | Insert Comment
2.25 | MANDATORY: The support and maintenance offering must include all labour and any related ancillary expenses associated with undertaking both preventative maintenance visits and breakdown repair visits (e.g., travel, subsistence and accommodation costs). Tenderers must confirm their understanding of this fact. | Compliant (Yes / No) | Insert Comment
2.26 | MANDATORY: Following the expiry of the warranty period, the support and maintenance offering will include the replacement of ‘wear and tear’ parts and/or the provision of service kits (with such parts being replaced during the scheduled planned preventative maintenance visits) and the provision of all components and routine spare parts. Tenderers must confirm their understanding of this fact, and must clearly define all ‘wear and tear’ parts, service kits, components and spare parts that will be made available, without additional charge, under the proposed support and maintenance offering. While tenderers are permitted to exclude from their support and maintenance offering certain replacement parts (e.g., items that, in the opinion of the tenderer, do not constitute ‘wear and tear’ parts or items that are not considered routine spare parts), any such exclusions must be clearly listed in response to this requirement. In the event that an exclusion is not clearly listed in response to this point, the successful tenderer shall be contractually bound to provide the part in question to the Contracting Authority at no additional cost. Additionally, tenderers are advised that the exclusion of a high number of parts from the proposed support and maintenance offering, or the proposal of any unreasonable exclusions (i.e., the exclusion of a part that is genuinely a ‘wear and tear’ part or a routine spare part), may pejoratively impact upon the score awarded under this award criterion and may, depending upon the severity of the exclusion(s) in question, result in a failure to comply with this requirement and the consequent elimination of the tender submission as a whole. | Compliant (Yes / No) | Insert Comment
2.27 | MANDATORY: In carrying out all service visits (whether planned preventative maintenance visits or breakdown repair visits), the successful tenderer must comply with all local policies and instructions of the Contracting Authority regarding site access and the timing of visits. The successful tenderer shall visit the site only once a prior arrangement has been made with a representative of the Contracting Authority. Tenderers must confirm their understanding of this fact and must outline any relevant procedures to be adhered to in this regard. | Compliant (Yes / No) | Insert Comment
2.28 | MANDATORY: All service visits required under the support and maintenance offering (whether planned preventative maintenance visits or breakdown repair visits) must be carried out by appropriately trained, qualified and experienced service engineers. In order to demonstrate compliance with this requirement, tenderers are required to submit the training profile of their service engineers (to include typical qualifications, post-qualification training and all certifications or accreditations) for consideration by the Contracting Authority. In particular, where the manufacturer of the equipment requires that any service engineers are trained to a particular level or standard, verifiable documentary evidence that such training has been received by a tenderer’s service engineers must be included with tender submissions. | Compliant (Yes / No) | Insert Comment
2.29 | MANDATORY: The successful tenderer must make available a signed and sufficiently comprehensive service report to a nominated representative of the Contracting Authority upon completion of any service visit (whether a planned preventative maintenance visit or a breakdown repair visit). Additionally, the successful tenderer must forward soft copies of the report to designated individuals following any reasonable request from the Contracting Authority. Tenderers must confirm their understanding of this fact. | Compliant (Yes / No) | Insert Comment
2.3 | MANDATORY: The service report referred to in Point 2(A).29 above must clearly identify the serial number of the equipment in question, the reason for the intervention, a full description of any components changed, any effects (i.e., the need for additional calibration, validation, etc), any uncertainty of measurement (where applicable), all relevant safety checks and certification of equipment upon completion. In order to demonstrate compliance with this requirement, tenderers are required to include in their tender submissions sample copies of proposed service reports. | Compliant (Yes / No) | Insert Comment | #
2.31 | MANDATORY: Should the successful tenderer, over the course of any service visit (whether a planned preventative maintenance visit or a breakdown repair visit), become aware of any issue, defect or potential issue that would render the equipment unsafe for use in a hospital environment to any extent, the Contracting Authority must be immediately notified of this fact. Tenderers must confirm their understanding of this fact and must outline any relevant procedures that will be followed in this regard. | Compliant (Yes / No) | Insert Comment
2.32 | MANDATORY: Following the reporting of a fault during core hours (i.e., 08:00 to 18:00, Monday to Friday), and in the event that the matter requires the attendance of an engineer, the maximum time for the attendance of a competent engineer at the Contracting Authority’s premises or delivery of a loan device shall be within 24 hours. Tenderers must include in their tender submissions full information regarding the mechanisms to be implemented in ensuring adherence to this maximum response time. | Compliant (Yes / No) | Insert Comment
2.33 | MANDATORY: In the event that a fault can not be resolved within 12 hours of attendance of an engineer, the successful tenderer shall provide a replacement device as a 'loaner unit' to the Contracting Authority until the issue has been resolved. There shall be no additional cost to the Contracting Authority for the use of this device. Tenderers must confirm their understanding of this fact and must outline any relevant procedures that will be followed in this regard. | Compliant (Yes / No) | Insert Comment
2.34 | MANDATORY: On the basis that a comprehensive support and maintenance offering is required, the successful tenderer must have in place a stock contingency protocol to ensure that the availability of parts is not impacted by any unforeseen events, including force majeure events or changes in distribution agreements, etc. In order to demonstrate compliance with this requirement, any such stock contingency policies and/or protocols must be provided for consideration by the Contracting Authority. | Compliant (Yes / No) | Insert Comment
2.35 | MANDATORY: Tenderers must state the period of time that they undertake to guarantee availability of a complete range of spare parts in the maintenance manual. | Compliant (Yes / No) | Insert Comment
2.36 | MANDATORY: Tenderers must state the launch date of the proposed equipment and any upgrade paths that have occurred. | Compliant (Yes / No) | Insert Comment
2.37 | MANDATORY: Tenderers must state if there has been a medical device alert with this model of device. If this is the case, tenderers must provide comprehensive documentation regarding the alert and actions taken by the manufacturer. | Compliant (Yes / No) | Insert Comment
2.38 | MANDATORY: Tenderers must state if there have been any adverse incidents involving the equipment which have resulted in reports to the Medical Devices Agency (U.K.), the Food and Drugs Administration (USA) or Health Products Regulatory Authority (Ireland). If this is the case, tenderers must provide comprehensive details of all adverse incidents. | Compliant (Yes / No) | Insert Comment
2.39 | DESIREABLE: The hardware and software required for the proposed equipment should be upgraded after 4 years. Tenderers must confirm the extent to which their proposal can comply with this requirement and provide all relevant information in this regard. | Compliant (Yes / No) | Insert Comment

## 3 - Training
SECTION 3 - PROPOSED TRAINING MECHANISMS
REF | REQUIREMENT | COMPLIANT (Yes / No) | DETAILED NARRATIVE COMMENT
3.1 | MANDATORY: Certified onsite/Offsite training of Clinical Engineering team and clinical end-users of the equipment must be provided promptly following the installation of devices purchased under the contract. Such training must be sufficient to provide the Clinical Engineering team with sufficient technical knowledge to deal with day-to-day issues potentially arising in respect of the equipment. Tenderers must confirm their understanding of this requirement and, in order to demonstrate compliance, must provide comprehensive information regarding the format and duration of proposed training. Samples of all training material to be provided must be included with tender submissions. | Compliant (Yes / No) | Insert Comment
3.2 | MANDATORY: The successful tenderer's initial training programme must involve the provision of relevant training material in both hard and soft copy formats. Tenderers must confirm their understanding of this fact. | Compliant (Yes / No) | Insert Comment
3.3 | MANDATORY: The successful tenderer must ensure that the interval between the delivery of any new equipment and the completion of required training is minimal. Tenderers are required to confirm their understanding of this fact and must outline all procedures that will be adhered to in this regard. | Compliant (Yes / No) | Insert Comment
3.4 | MANDATORY: In terms of 'refresher' training, as a component of the support and maintenance offering, the successful tenderer will be required to provide two onsite training session per annum to Contracting Authority's clinical end-users and clinical engineering of the equipment. Tenderers must confirm their understanding of this requirement and, in order to demonstrate compliance, must provide comprehensive information regarding the format and duration of proposed training. Samples of all training material to be provided must be included with tender submissions. | Compliant (Yes / No) | Insert Comment
3.5 | MANDATORY: The successful tenderer must amend the 'refresher' training programme, or provide additional training, in response to instructions from the Contracting Authority (NOTE: The Contracting Authority accepts that such instructions would normally be in response to corrective and/or preventive actions arising out of non-conformance management, changes in clinical practice or clinical issues). In order to demonstrate compliance with this requirement, tenderers must provide details of the mechanism to be employed to adapt training measures on an as-needed basis throughout the duration of the contract. | Compliant (Yes / No) | Insert Comment
3.6 | MANDATORY: Tenderers must provide the Contracting Authority with details of any E-learning resources or online platforms that users may use to improve their understanding and utilisation of the proposed IONM system. Tenderers must allow user access to these resources for a trial period prior to tender completion. | Compliant (Yes / No) | Insert Comment

## 4 - Reliability Continuity
SECTION 4 - RELIABILITY AND CONTINUITY OF PROPOSED EQUIPMENT
REF | REQUIREMENT |  | DETAILED NARRATIVE COMMENT
4.1 | Tenderers must demonstrate the reliability and continuity of their proposed products and associated services by nominating two reference sites, which may be contacted by the Contracting Authority independently and without further reference to the tenderer. In order to be considered a legitimate reference site, the following conditions must be complied with: • Each reference site must have been supplied with similar fluid warming devices to the devices proposed in response to this Invitation to Tender document, or near equivalents; • Each reference site must comprise a hospital of comparable size to Beaumont Hospital. Scores under this criterion will be apportioned on an overall basis, taking into account the similarity of previous projects, in terms of nature, scale and complexity, to the Contracting Authority's advertised requirements and the extent to which information provided by reference sites enables the Contracting Authority to place due confidence in a tenderer's ability to provide the products and related services in question in a sufficiently reliable and continuous manner. |  | N/A | Insert Detailed Narrative Response
REFERENCE PROJECT #1 | (i) | Name of Reference Site. | Insert
| (ii) | Detailed description of Reference Site. | Insert
| (iii) | Name of client contact person / referee. | Insert
| (iv) | Position of client contact person / referee. | Insert
| (v) | E-mail address of client contact person / referee. | Insert
| (vi) | Telephone number of client contact person / referee. | Insert
| (vii) | Commencement date of this contract | Insert
| (viii) | Termination date of this contract | Insert
| (ix) | Annual value of this contract | Insert
| (x) | Detailed description of all products and equipment provided under this contract. | Insert
REFERENCE PROJECT #2 | (i) | Name of Reference Site. | Insert
| (ii) | Detailed description of Reference Site. | Insert
| (iii) | Name of client contact person / referee. | Insert
| (iv) | Position of client contact person / referee. | Insert
| (v) | E-mail address of client contact person / referee. | Insert
| (vi) | Telephone number of client contact person / referee. | Insert
| (vii) | Commencement date of this contract | Insert
| (viii) | Termination date of this contract | Insert
| (ix) | Annual value of this contract | Insert
| (x) | Detailed description of all products and equipment provided under this contract. | Insert