Appendix 7

Sample Analysis Information for State Laboratory NIR & Calibration Set Tender

SL202602

1            Equipment

The instrument and calibration set proposed for the tender must be used for this sample analysis.

2            Samples

A series of animal feeding stuff samples are provided to quantify (a) % Crude Fibre, and,

(b) % Crude Oil levels in the measurement ranges described in Appendix 1 Section B2, viz. 1% - 25% for Crude Oil and 1.5% - 40% for Crude Fibre.

If some sample levels are outside the above ranges please report results as found.

A total of 14 samples for the determination of % Fibre and of % Oil B are provided as follows:

| SL Ref. No. | Category | Matrix | Description |
| --- | --- | --- | --- |
| SL-Sample-01 | FEED_MAT | DAIRY | Lucerne pellets |
| SL-Sample-02 | CFS | SHEEP | Lamb Grow + Finish 18% Cube |
| SL-Sample-03 | FEED_MAT | ALL | Soya Bean Hulls |
| SL-Sample-04 | CFS | CALF | Calf 18% Crunch |
| SL-Sample-05 | CFS | SHEEP | Sheep Nuts |
| SL-Sample-06 | CFS | PIG | Gilt Feed Pellets |
| SL-Sample-07 | CFS | PIG | FV Silver Medal Pig Fattener |
| SL-Sample-08 | CFS | POULTRY | WBCC Free Broiler Finisher |
| SL-Sample-09 | FEED_MAT | ALL | Palm Kernel Expeller |
| SL-Sample-10 | FEED_MAT | ALL | Soya Bean Hulls |
| SL-Sample-11 | CFS | SHEEP | GV Ewe & Lamb 19% Coarse |
| SL-Sample-12 | CFS | HORSE | Horse & Pony Nut |
| SL-Sample-13 | CFS | HORSE | Horse Ration |
| SL-Sample-14 | CFS | BEEF | Prime Beef 12% Cubes |

CFS: Compound Feedstuff

FEED-MAT: Feed Material

3	Analysis Instructions:

All samples are ground to 1mm and must be analysed as received by NIR spectroscopy.

Run 1:	Please analyse samples in duplicate in the same run as follows:

Mix the sample thoroughly before filling the sample cell ensuring homogeneity and take the reading (Replicate A).

A second reading (Replicate B) is done on a freshly filled cell (i.e. on a fresh aliquot of sample)

Review the results. Where the repeatability criteria, as outlined in Section 4 below, have not been met for a sample analysed in duplicate, the sample should be re-analysed at least once more, i.e. Replicate C, and if necessary, a fourth replicate, Replicate D should be run.

Note: All replicates are done with freshly filled cells.

Run 2:	Repeat the above analysis, on another day.

All results must be reported to 2 decimal places using Sheet 1 (Fibre) and Sheet 2 (Oil) of the supplied Excel spreadsheet, Appendix 8 –Samples Results Worksheet SL NIR Tender SL202602.

4	Evaluation of Test Results:

The results will be evaluated as follows:

The difference between results for within run duplicate analyses should not exceed:

-	0.2% in absolute value for Crude Oil results

-	0.3% in absolute value for Crude Fibre results less than 10%

-	0.4% in absolute value for Crude Fibre results of 10% to 20%

-	0.5% in absolute value for Crude Fibre results greater than 20%

The Precision will be determined by assessment of within day and between day variation Accuracy will be assessed by comparison with results obtained by Official Methods as documented in Commission Directive 152/2009 for Crude Fibre & for Crude Oil (Procedure B). Results obtained for Sample Analysis should be within 18% of results obtained by the State Laboratory using the Official Methods for Fibre and Oil B determination.